## Laboratory Blanks
Appendix B. Results of laboratory blanks in water samples, Watershed Sources Project, Sinclair Inlet, Washington, January 2008 through October 2009.
[Abbreviations: NWQL, National Water Quality Laboratory; WMRL, Wisconsin Mercury Research Laboratory; CV-AFS, cold vapor atomic fluorescence spectrometry; ICP-OES, inductively coupled plasma optical emission spectrometry; UV, ultra violet; IR, infrared spectrometry; IC, ion chromatography; ng/L, nanogram per liter; mg/L, milligram per liter; µg/L; microgram per liter]
Analyte | Laboratory | Method | Type of Blank | Units | Inclusive Dates |  | Number of samples | Median | Mean | Standard Deviation | Minimum | Maximum
Total Mercury, filtered | NWQL | CV-AFS | Instrument Blank | ng/L | 2008-01-16 | 2008-07-29 | 171 | -0.017 | -0.041 | 0.98 | -2.474 | 2.4658
|  |  | Blind Blank |  | 2008-01-16 | 2008-07-29 | 37 | 3.124 | 3.63 | 3.18 | -1.47 | 8.896
Total Mercury, particulate | WMRL | CV-AFS | Instrument Blank | ng/filter | 2008-04-14 | 2008-08-14 | 3 | 0.03 | 0.072316 | 0.073294 | 0.03 | 0.15695
Methyl-Mercury, filtered | WMRL | CV-AFS | Instrument Blank | ng/L | 2008-04-22 | 08/14/08a | 5 | 0.0525 | 0.0534 | 0.022675 | 0.03 | 0.081
Total Iron | NWQL | ICP-OES | Instrument Blank | µg/L | 2008-05-07 | 2009-08-25 | 2983 | 0.322 | 0.309 | 0.9989 | -4.891 | 3.753
|  |  | Blind Blank |  | 2008-05-07 | 2009-08-25 | 65 | 0.079 | 0.2282 | 1.2726 | -2.682 | 4.954
Total Manganese | NWQL | ICP-OES | Instrument Blank | µg/L | 2008-05-07 | 2009-08-27 | 3016 | 0.031 | 0.036861 | 0.32498 | -2.049 | 15.01
|  |  | Blind Blank |  | 2008-05-07 | 2009-08-25 | 66 | -0.004 | -0.018409 | 0.064252 | -0.259 | 0.116
DOC, dissolved organic carbon | NWQL | UV promoted persulfate oxidation and IR | Instrument Blank | mg/L | 2008-04-01 | 2009-10-27 | 876 | 0.0147 | 0.023275 | 0.048014 | -0.0914 | 0.2107
|  |  | Blind Blank |  | 2008-05-01 | 2009-08-26 | 68 | 0.0635 | 0.071406 | 0.07529 | -0.023 | 0.3886
Ammonia | NWQL | Kone Discrete Analyzer | Instrument Blank | mg/L | 2008-05-20 | 09/17/08a | 41 | 2e-05 | 5.0244e-05 | 0.0023852 | -0.00602 | 0.00581
|  |  | Blind Blank |  | 2008-05-20 | 09/23/08a | 22 | 0.00072 | 0.0010441 | 0.0037787 | -0.00798 | 0.00729
Nitrite | NWQL | Kone Discrete Analyzer | Instrument Blank | mg/L | 2008-05-20 | 09/17/08a | 41 | -5e-05 | -8.8537e-05 | 0.00047336 | -0.00252 | 0.00056
|  |  | Blind Blank |  | 2008-05-16 | 09/17/08a | 21 | 0 | 0.00092381 | 0.0028325 | -0.003 | 0.011
Nitrate and nitrite | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | Instrument Blank | mg/L | 2008-05-16 | 09/17/08a | 84 | 0 | 0.00025833 | 0.0016804 | -0.004 | 0.005
|  |  | Blind Blank |  | 2008-05-16 | 09/17/08a | 21 | 0 | 0.00092381 | 0.0028325 | -0.003 | 0.011
Total Nitrogen, filtered | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | Instrument Blank | mg/L | 2008-05-22 | 09/10/08a | 84 | -0.004 | -0.0030714 | 0.012958 | -0.028 | 0.029
|  |  | Blind Blank |  | 2008-05-19 | 09/24/08a | 20 | -0.0135 | -0.015 | 0.012105 | -0.041 | 0.003
Total Nitrogen, unfiltered | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | Instrument Blank | mg/L | 2008-05-22 | 09/11/08a | 116 | -0.003 | -0.0013707 | 0.01574 | -0.042 | 0.029
|  |  | Blind Blank |  | 2008-05-19 | 09/24/081 | 20 | -0.0215 | -0.01835 | 0.019887 | -0.046 | 0.023
Ortho-phosphate | NWQL | Kone Discrete Analyzer | Instrument Blank | mg/L | 2008-05-20 | 09/19/081 | 46 | 0.002165 | 0.0018767 | 0.0012347 | -0.00261 | 0.00292
|  |  | Blind Blank |  | 2008-05-20 | 09/23/081 | 21 | 0.00207 | 0.0018129 | 0.0011845 | -0.0028 | 0.00312
Total Phosphorous, filtered | NWQL | Low Level, Continuous Flow Analyzer | Instrument Blank | mg/L | 2008-05-21 | 2008-07-29 | 3 | 0.0003 | 0.00063333 | 0.00066583 | 0.0002 | 0.0014
|  |  | Blind Blank |  | 2008-05-14 | 09/24/08a | 20 | 0.0004 | 0.00057 | 0.00085968 | -0.0007 | 0.0029
Total Phosphorous, unfiltered | NWQL | Low Level, Continuous Flow Analyzer | Instrument Blank | mg/L | 2008-07-24 | 09/04/08a | 7 | 0.0003 | 0.00041429 | 0.00054292 | -0.0001 | 0.0015
|  |  | Blind Blank |  | 2008-05-22 | 09/24/08a | 20 | 0.00085 | 0.000895 | 0.001062 | -0.0007 | 0.0031
Particulate Inorganic Carbon | NWQL | Elemental analyzer | Instrument Blank | mg/L | 5/13/082 | 8/24/2009a,b | 51 | -0.004 | -0.0030392 | 0.0069799 | -0.018 | 0.022
|  |  | Blind Blank |  | 2008-05-13 | 8/24/2009a,b | 62 | -0.001 | -0.0019516 | 0.011261 | -0.031 | 0.031
Total Carbon, particulate | NWQL | Elemental analyzer | Instrument Blank | mg/L | 5/13/082 | 8/24/2009a,b | 51 | 0.007 | 0.006 | 0.012075 | -0.019 | 0.045
|  |  | Blind Blank |  | 2008-05-13 | 8/24/2009a,b | 62 | 0.016 | 0.016226 | 0.020031 | -0.046 | 0.053
Total Nitrogen, particulate | NWQL | Elemental analyzer | Instrument Blank | mg/L | 2008-05-13 | 8/24/2009a,b | 51 | 0.003 | 0.0030784 | 0.0060658 | -0.014 | 0.015
|  |  | Blind Blank |  | 2008-05-13 | 8/24/2009a,b | 62 | 0.002 | 0.0017097 | 0.0068959 | -0.015 | 0.017
Sulfate | NWQL | IC | Instrument Blank | mg/L | 2008-05-17 | 07/22/08a | 381 | 0 | 0.0046955 | 0.046159 | -0.146 | 0.379
Chloride | NWQL | IC | Instrument Blank | mg/L | 2008-05-17 | 07/22/08a | 84 | -0.004 | -0.0030714 | 0.012958 | -0.028 | 0.029
a Statistics include two or three days from the end of January 2009
b Based on nominal volume of 150 mL

## Mercury Field Blanks
Appendix B. Total mercury and methylmercury in laboratory and field blank water, Watershed Sources Project, Sinclair Inlet, Washington December 2007 through May 2010.
[QA WA WSC PURELAB ULTRA PURE, Washington Water Science Center Purelab Ultra Pure water system, DD, dry dock; T, temporary well; (see Appendix A for complete description of station names and field Identifiers). USGS Site identifier: Unique number for each site based on latitude and longitude of the site. First six digits are latitude, next seven digits are longitude, and final two digits are a sequence number to uniquely identify each site. Abbreviations: ng/L, nanograms per liter]
Field Identifier | USGS Site Identifier | Date | Time | Concentration (ng/L)
Total Mercury
Laboratory water analyzed by National Water Quality Laboratory - Unfiltered
QA WA WSC PURELAB ULTRA PURE | 471516122262202 | 2007-12-10 | 1458 | <10
QA WA WSC PURELAB ULTRA PURE | 471516122262202 | 2007-12-10 | 1608 | <10
QA WA WSC PURELAB ULTRA PURE | 471516122262202 | 2008-01-22 | 1508 | <10
Laboratory water analyzed by Wisconsin Mercury Research Laboratory - Unfiltered
QA WA WSC PURELAB ULTRA PURE | 471516122262202 | 2008-06-25 | 1148 | <.04
QA WA WSC PURELAB ULTRA PURE | 471516122262202 | 2010-05-03 | 1258 | 0.22a
Field blank water analyzed by Wisconsin Mercury Research Laboratory - Filtered
ANNAPOLIS | 12072520 | 2008-07-15 | 0958 | 0.33
DD-6 | 473316122383301 | 2008-03-26 | 1308 | 0.08
DD-6 | 473316122383301 | 2008-04-25 | 1348 | 0.09
T-3 | 473321122390501 | 2008-06-05 | 0938 | 0.11
T-1 | 473321122390701 | 2008-06-05 | 0938 | 0.28
Methylmercury
Field blank water analyzed by Wisconsin Mercury Research Laboratory - Filtered
ANNAPOLIS | 12072520 | 2008-07-15 | 0958 | <.04
aTaken in association with the tumble core control experiment and after field operations.

## Laboratory Standards
Appendix B. Results of laboratory standards, Watershed Sources Project, Sinclair Inlet, Washington, January through November 2009.
[Abbreviations: NWQL, National Water Quality Laboratory; WMRL, Wisconsin Mercury Research Laboratory; CV-AFS, cold vapor atomic fluorescence spectrometry; ICP-OES, inductively coupled plasma optical emission spectrometry; CCV, continuous calibration verification sample; TPC, third party control; SRS, standard reference water sample; MS, matrix spike; MSD, matrix spike duplicate; CRM, certified reference material; SRM, standard reference material; DigCk, digestion check; LCS, laboratory control sample; LAC, laboratory accuracy check; ng/L, nanogram per liter; mg/L, milligram per liter; µg/L; microgram per liter;—, not applicable]
Analyte | Laboratory | Method | Type of Laboratory Measurement | Measure | Inclusive Dates |  | Units | Concentration Range | Number of samples | Median | Mean | Standard Deviation | Minimum | Maximum
Total Mercury, filtered | NWQL | CV-AFS | CCV | Percent Recovery | 2008-03-11 | 2008-07-29 | ng/L | 40.9 - 55.5 | 72 | 97.5 | 96.5972 | 6.3394 | 82 | 111
|  |  | CCV | Percent Recovery | 2008-03-11 | 2008-07-29 | ng/L | 81.5 - 114 | 58 | 96 | 94.2683 | 6.9091 | 82 | 114
|  |  | TPC | Percent Recovery | 2008-01-16 | 2008-07-03 | ng/L | 15 | 6 | 102.3 | 100.8 | 10.7 | 85 | 117.4
|  |  | TPC | Percent Recovery | 2008-01-16 | 2008-07-03 | ng/L | 75 | 6 | 102 | 103.4 | 4.2 | 99.2 | 110.2
|  |  | Blind SRS | Percent Recovery | 2008-01-24 | 2008-06-23 | ng/L | 14 | 18(1) | 97.9 | 95.0444 | 11.7 | 52.1 | 105.8
|  |  | Blind SRS | Percent Recovery | 2008-01-24 | 2008-07-09 | ng/L | 25 | 15 | 94.6 | 94.5 | 9.2473 | 84.8 | 122.7
|  |  | Blind SRS | Percent Recovery | 2008-02-05 | 2008-07-29 | ng/L | 34 | 20 | 104.8 | 105.965 | 8.0407 | 89.8 | 121.1
Total Mercury, filtered | WMRL | CV-AFS | MS | Percent Recovery | 2008-04-15 | 2008-09-23 | ng/L | 0.968 - 1.046 | 6 | 100.45 | 100.9 | 3.0659 | 96.8 | 104.6
|  |  | MS | Percent Recovery | 2008-04-15 | 2008-09-23 | ng/L | 2.25 - 2.6575 | 14 | 100.25 | 99.45 | 4.7022 | 90 | 106.3
|  |  | MS | Percent Recovery | 2008-04-15 | 2008-09-23 | ng/L | 4.76 - 4.98 | 4 | 98.85 | 98.125 | 1.9923 | 95.2 | 99.6
|  |  | CRM | Percent Recovery | 2008-05-27 | 09/24/082 | ng/L | 4.77 - 5.48 | 50 | 99.9842 | 99.968 | 1.1102 | 96.842 | 103.158
Total Mercury, particulate | WMRL | CV-AFS | MS | Percent Recovery | 2008-04-14 | 2008-08-14 | ng/L | 7.19 - 10.2 | 17 | 94.3098 | 88.9834 | 10.9861 | 71.9493 | 102.0296
|  |  | SRM | Percent Recovery | 2008-04-14 | 2008-08-14 | ng/g | 135.8 - 273.54 | 20 | 91.8988 | 97.5547 | 20.2517 | 80.8316 | 162.8242
Methyl-Mercury, filtered | WMRL | CV-AFS | MS | Percent Recovery | 2008-04-22 | 08/14/082 | ng/L | 0.8 - 0.98 | 18 | 101.4 | 102.9611 | 5.3892 | 96.5 | 117.6
|  |  | MS/MSD | RPD | 2008-04-22 | 08/14/082 | ng/L | 0.8 - 0.98 | 9 | 1.3381 | 2.1598 | 2.0992 | 0.66382 | 7.316
Total Iron | NWQL | ICP-OES | CCV | Percent Recovery | 2008-05-12 | 2009-08-25 | µg/L | 500 | 3117 | 99 | 99.6946 | 3.331 | 94 | 108
|  |  | TPC | Percent Recovery | 2008-05-09 | 2009-08-26 | µg/L | 500 | 12 | 100.4 | 100.7133 | 2.0486 | 97.92 | 105.48
Total Manganese | NWQL | ICP-OES | CCV | Percent Recovery | 2008-05-07 | 2009-08-27 | µg/L | 234 - 266 | 3116 | 101 | 100.8065 | 1.9847 | 93 | 107
|  |  | TPC | Percent Recovery | 2008-05-09 | 2009-08-26 | µg/L | 250 | 13 | 98.6 | 99.1385 | 1.5943 | 97.4 | 103.28
|  |  | Blind SRS | Percent Recovery | 2008-05-12 | 2009-08-12 | µg/L | 4.47 - 8.68 | 48 | 98.5928 | 98.6277 | 4.1369 | 86.3041 | 107.9695
|  |  | Blind SRS | Percent Recovery | 2008-05-16 | 2009-07-15 | µg/L | 16.48 - 36.9 | 42 | 97.0396 | 97.6532 | 3.3152 | 91.9802 | 106.7354
Dissolved Organic Carbon | NWQL | ultra violet (UV)-promoted persulfate oxidation and infrared spectrometry | CCV | Percent Recovery | 2008-05-13 | 2009-08-18 | mg/L | 4.04 - 5.80 | 112 | 101 | 102.0536 | 5.5521 | 83 | 119
|  |  | CCV | Percent Recovery | 2008-05-13 | 2009-08-18 | mg/L | 8.6 - 11.2 | 275 | 98 | 98.2218 | 5.3735 | 87 | 113
|  |  | CCV | Percent Recovery | 2008-05-13 | 2009-08-18 | mg/L | 18.3 - 22.3 | 176 | 101 | 100.8239 | 4.4462 | 91 | 113
|  |  | CCV | Percent Recovery | 2008-05-13 | 2009-08-18 | mg/L | 27.4 - 33.9 | 159 | 100 | 99.4906 | 3.8383 | 91 | 113
|  |  | TPC | Percent Recovery | 2008-05-13 | 2009-08-20 | mg/L | 4.2953 - 4.9621 | 47 | 93.898 | 93.3884 | 2.997 | 85.906 | 99.242
|  |  | Blind SRS | Percent Recovery | 2008-05-07 | 2009-11-15 | mg/L | 1.4469 - 12.7005 | 37 | 103.0504 | 103.997 | 6.1786 | 93.1444 | 116.2831
Ammonia | NWQL | Kone Discrete Analyzer | CCV | Percent Recovery | 2008-05-20 | 09/17/082 | mg/L | 0.56687 - 0.6176 | 83 | 98 | 98.506 | 1.9654 | 94 | 103
|  |  | TPC HIGH | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.56807 - 0.60191 | 72(3) | 97.3275 | 97.2849 | 1.5508 | 94.6783 | 100.3183
|  |  | TPC MED | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.27534 - 0.29619 | 72(4) | 94.7817 | 94.9421 | 1.5329 | 91.78 | 98.73
|  |  | TPC LOW | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.06476 - 0.09903 | 73 | 91.3 | 92.1589 | 5.2998 | 86.3 | 132
|  |  | Blind SRS | Percent Recovery | 2008-05-20 | 09/18/082 | mg/L | 0.22866 - 1.18252 | 48 | 96.8174 | 99.3123 | 9.7124 | 66.0667 | 125.3217
Nitrite | NWQL | Kone Discrete Analyzer | CCV | Percent Recovery | 2008-05-20 | 09/17/082 | mg/L | 0.11668 - 0.12554 | 79 | 100 | 100.443 | 1.3563 | 97 | 105
|  |  | TPC HIGH | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.16048 - 0.17029 | 59 | 101.9 | 102.2 | 1.4055 | 100.3 | 106.4
|  |  | TPC MED | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.07892 - 0.0847 | 59 | 101.8375 | 101.5248 | 1.717 | 98.65 | 105.875
|  |  | TPC LOW | Percent Recovery | 2008-09-09 | 2008-09-11 | mg/L | 0.01966 - 0.0214 | 59 | 102.3 | 101.7559 | 1.9828 | 98.3 | 107
Nitrate + Nitrite | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | CCV | Percent Recovery | 2008-05-16 | 08/25/082 | mg/L | 0.495 - 0.512 | 16 | 100.4 | 100.5037 | 1.0708 | 99 | 102.4
|  |  | LCS HIGH | Percent Recovery | 2008-05-16 | 08/25/082 | mg/L | 0.733 - 0.782 | 16 | 99.7867 | 99.9467 | 1.5976 | 97.7333 | 104.2667
|  |  | LCS MED | Percent Recovery | 2008-05-16 | 08/25/082 | mg/L | 0.484 - 0.519 | 59 | 100.6 | 100.5041 | 1.3912 | 96.8 | 103.8
|  |  | LCS LOW | Percent Recovery | 2008-05-16 | 08/25/082 | mg/L | 0.121 - 0.129 | 59 | 100.72 | 100.3105 | 1.6299 | 96.8 | 103.2
|  |  | Blind SRS HIGH | Percent Recovery | 2008-05-15 | 09/17/082 | mg/L | 0.2377 - 0.374 | 28 | 100.9722 | 100.6652 | 1.7186 | 96.6667 | 103.8889
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-15 | 09/17/082 | mg/L | 0.0706 - 0.108 | 32 | 98.2714 | 98.5088 | 2.1826 | 93.2692 | 103.8462
Orthophosphate | NWQL | Kone Discrete Analyzer | CCV | Percent Recovery | 2008-05-20 | 09/17/082 | mg/L | 0.11848 - 0.12928 | 84 | 103 | 103.3214 | 2.3396 | 99 | 108
|  |  | TPC HIGH | Percent Recovery | 9/9/2008 | 2008-09-11 | mg/L | 0.16112 - 0.17345 | 87 | 104.0875 | 103.9147 | 1.976 | 100.7 | 108.4062
|  |  | TPC MED | Percent Recovery | 9/9/2008 | 2008-09-11 | mg/L | 0.07853 - 0.0881 | 87 | 103.825 | 103.1477 | 2.839 | 98.1625 | 110.125
|  |  | TPC LOW | Percent Recovery | 9/9/2008 | 2008-09-11 | mg/L | 0.01704 - 0.02324 | 87 | 98.9 | 98.823 | 4.7614 | 85.2 | 116.2
|  |  | Blind SRS HIGH | Percent Recovery | 2008-05-20 | 09/16/082 | mg/L | 0.56604 - 0.64977 | 18 | 99.1397 | 99.2268 | 2.1995 | 95.2929 | 103.1381
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-20 | 09/16/082 | mg/L | 0.17597 - 0.27919 | 23 | 98.2185 | 97.656 | 5.2983 | 89.763 | 107.9911
Total Nitrogen, filtered | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | CCV | Percent Recovery | 2008-05-30 | 09/22/082 | mg/L | 3.525 - 3.975 | 84 | 100 | 100.0714 | 1.8743 | 94 | 106
|  |  | DigCk | Percent Recovery | 2008-05-20 | 09/17/082 | mg/L | 3.82 - 4.121 | 39 | 99 | 98.6923 | 1.88 | 96 | 103
|  |  | TPC HIGH | Percent Recovery | 2008-05-30 | 09/22/082 | mg/L | 3.587 - 3.715 | 20 | 99.7002 | 99.655 | 0.9151 | 98.0055 | 101.5027
|  |  | TPC MED | Percent Recovery | 2008-05-30 | 09/22/082 | mg/L | 1.745 - 1.978 | 72 | 98.8452 | 98.9151 | 1.8573 | 95.3552 | 108.0874
|  |  | TPC LOW | Percent Recovery | 2008-05-30 | 09/22/082 | mg/L | 0.345 - 0.398 | 70 | 99.7268 | 99.8602 | 3.02 | 94.2623 | 108.7432
Total Phosphorus, filtered | NWQL | Low Level, Continuous Flow Analyzer | CCV | Percent Recovery | 2008-05-06 | 09/18/082 | mg/L | 0.1434 - 0.1587 | 85 | 101 | 100.8706 | 1.3165 | 96 | 106
filtered total phosphorus |  |  | TPC HIGH | Percent Recovery | 2008-05-08 | 09/18/082 | mg/L | 0.1679 - 0.1799 | 128 | 102 | 101.6875 | 1.4513 | 99 | 106
|  |  | TPC MED | Percent Recovery | 2008-05-08 | 09/18/082 | mg/L | 0.1171 - 0.1279 | 127 | 101 | 100.6693 | 1.7596 | 97 | 106
|  |  | TPC LOW | Percent Recovery | 2008-05-08 | 09/18/082 | mg/L | 0.0221 - 0.0266 | 125 | 100 | 100.152 | 3.208 | 91 | 110
|  |  | Blind SRS HIGH | Percent Recovery | 2008-05-08 | 09/18/082 | mg/L | 0.2873 - 0.2929 | 3 | 101.2544 | 101.1382 | 0.98078 | 100.1045 | 102.0557
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-08 | 09/18/082 | mg/L | 0.0568 - 0.0656 | 30 | 101.0426 | 100.5798 | 2.2032 | 94.5092 | 104.3261
Total Nitrogen, unfiltered | NWQL | Colorimetry, ASF, cadmium reduction-diazotization | CCV | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 3.4875 - 3.9 | 128 | 100 | 99.8438 | 1.8377 | 93 | 104
unfiltered total nitrogen |  |  | DigCk | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 3.81 - 4.093 | 63 | 99 | 98.7302 | 1.4724 | 95 | 102
|  |  | TPC HIGH | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 3.478 - 3.835 | 125 | 99 | 99.192 | 1.9037 | 95 | 105
|  |  | TPC MED | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 1.725 - 1.898 | 437 | 98 | 98.3776 | 1.9575 | 94 | 104
|  |  | TPC LOW | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 0.324 - 0.415 | 441 | 99 | 99.4943 | 3.9226 | 89 | 113
|  |  | Blind SRS HIGH | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 1.523 - 1.806 | 25 | 98.4211 | 98.4508 | 3.7328 | 91.8129 | 105.614
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-19 | 09/11/082 | mg/L | 0.346 - 1.005 | 26 | 96.0135 | 96.7105 | 9.2081 | 85 | 135.8108
Total Phosphorus, unfiltered | NWQL | Low Level, Continuous Flow Analyzer | CCV | Percent Recovery | 2008-05-01 | 09/19/082 | mg/L | 0.1447 - 0.1645 | 196 | 100 | 100.6582 | 1.4144 | 96 | 110
unfiltered total phosphorus |  |  | TPC HIGH | Percent Recovery | 2008-05-13 | 09/17/082 | mg/L | 0.1617 - 0.1875 | 276 | 101 | 101.3913 | 1.8128 | 95 | 111
|  |  | TPC MED | Percent Recovery | 2008-05-13 | 09/17/082 | mg/L | 0.1163 - 0.1282 | 276 | 101 | 100.7899 | 1.7826 | 96 | 106
|  |  | TPC LOW | Percent Recovery | 2008-05-13 | 09/17/082 | mg/L | 0.0217 - 0.0261 | 272 | 99 | 98.9154 | 3.1465 | 90 | 108
|  |  | Blind SRS HIGH | Percent Recovery | 2008-05-13 | 09/17/082 | mg/L | 0.2867 - 0.419 | 5 | 102.0557 | 101.561 | 0.98564 | 99.8955 | 102.23
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-13 | 09/17/082 | mg/L | 0.0588 - 0.0663 | 30 | 101.0815 | 101.3071 | 2.3077 | 97.5155 | 107.3211
Total Carbon, particulate | NWQL | Elemental analyzer | CCV | Percent Recovery | 2008-06-10 | 2009-09-17 | mg/L | 68.6 - 72.3 | 733 | 100 | 99.9686 | 0.90255 | 97 | 102
|  |  | Long-term LAC | Percent Recovery | 2008-05-30 | 2009-11-03 | mg/L | 31.818 - 35.171 | 40 | 101.4364 | 101.1421 | 2.0684 | 94.2337 | 104.1641
Total Nitrogen, particulate | NWQL | Elemental analyzer | CCV | Percent Recovery | 2008-06-10 | 2009-09-17 | mg/L | 9.88 - 10.79 | 176 | 100 | 100.0682 | 1.5291 | 96 | 105
|  |  | SRM | Percent Recovery | 2008-05-30 | 2009-11-03 | mg/L | 4.407 - 7.917 | 26 | 98.2218 | 100.6526 | 13.4662 | 92.1967 | 165.6276
Chloride | NWQL | Ion Chromatography | Blind SRS HIGH | Percent Recovery | 2008-05-17 | 09/25/082 | mg/L | 41.904 - 51.555 | 19 | 96.3591 | 96.3338 | 1.334 | 94.0606 | 98.8117
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-17 | 09/25/082 | mg/L | 12.409 - 26.893 | 41 | 96.6601 | 97.0957 | 2.0453 | 92.7607 | 101
Sulfate | NWQL | Ion Chromatography | Blind SRS HIGH | Percent Recovery | 2008-05-17 | 09/25/082 | mg/L | 55.323 - 143.333 | 30 | 98.8097 | 98.3524 | 1.5461 | 95.1384 | 100.827
|  |  | Blind SRS LOW | Percent Recovery | 2008-05-17 | 09/25/082 | mg/L | 7.763 - 30.819 | 29 | 100.5354 | 100.1625 | 2.0486 | 96.3934 | 103.4187
1Includes one value of 51% recovery
2Statistics include two or three days from the end of January 2009
39/11/2008 20:30 value of 166500 not included in statistics
49/11/2008 20:30 value of 333000 not included in statistics

## Mercury SRMs
Appendix B. Total mercury and particulate mercury in standard reference samples; Watershed Sources Project, Sinclair Inlet, Washington, January through July 2008.
[OUA, Operable Unit A; OUBT, Operable Unit B Terrestrial (see Appendix A for complete description of station names and field Identifiers); Abbreviations: NWQL, National Water Quality Laboratory; WMRL, Wisconsin Mercury Research Laboratory; ng/L, nanograms per liter; MPV, most probable value; SRM, Standard reference material; ng/L , nanograms per liter; ng/mg, nanograms per milligram; %, percent].
Standard reference material in water
Standard reference sample | Type of standard | Site submitted under | Dates | Times | Laboratory analyzed by | Filtered mercury (ng/L) | MPV (ng/L) | Deviation (ng/L) | Percent recovery (%) |  | Deviation | Percent recovery
USGS BQS Hg-45 | SRS | OUA-204 | 2008-01-28 | 1254 | NWQL | 9 | 8 | 1 | 112.5 |  | 2 | 0
USGS BQS Hg-44 | SRS | Anderson | 2008-05-08 | 1144 | WMRL | 13.6 | 14 | 2 | 97.1429 |  | 1
Standard reference material on particulates
Standard reference sample |  | Site submitted under | Dates | Times | Laboratory analyzed by | Reported mercury mass on filter (ng) | SRM weight (mg) | MPV (ng/mg) | MPV Mercury Mass on filter (ng) | Percent recovery (%)
IAEA-405 | SRM | OUBT - 715R | 2008-05-06 | 0704 | WMRL | 5.18 | 5.75 | 0.81 | 4.658 | 111.2
IAEA-405 | SRM | Blackjack | 2008-05-06 | 1440 | WMRL | 1.54 | 2.34 | 0.81 | 1.895 | 81.3
IAEA-405 | SRM | Olney | 2008-07-21 | 1503 | WMRL | 6.434 | 9.18 | 0.81 | 7.436 | 86.5
PACS-2 | CRM | OUBT-722H | 2008-05-06 | 1204 | WMRL | 51.1 | 22.57 | 3.04 | 68.62 | 74.5
PACS-2 | CRM | Anderson | 2008-07-21 | 1203 | WMRL | 29 | 14.55 | 3.04 | 44.23 | 65.6

## Interlaboratory comparison
Appendix B. Interlaboratory comparison of total mercury in water samples, Watershed Sources Project, Sinclair Inlet, Washington, June 2008 through March 2010.
[OUBT, Operable Unit B Terrestrial; PSNS, Puget Sound Naval Shipyard (see Appendix A for complete description of station names and field Identifiers). Abbreviations: µg/L, micrograms per liter; ng/L, nanograms per liter; NWQL, National Water Quality Laboratory; WMRL, Wisconsin Mercury Research Laboratory]
Filtered Total Mercury
Field Identifer | Date | Time | NWQL (µg/L) | WMRL (ng/L) | Mean | Relative percent difference
OUBT-722 | 2008-06-04 | 1944 | 0.102 | --
|  | 1947 | -- | 105 | 103.5 | 0.72464
| 2008-06-05 | 930 | 2.19 | --
|  | 931 | -- | 1950 | 2070 | 2.8986
| 2008-06-05 | 1231 | 1.76 | --
|  | 1232 | -- | 1810 | 1785 | 0.70028
PSNS015 | 2009-12-29 | Composite | <0.010 | -- | Discrepancy between NWQL and WMRL results- all filtered total mercury data from study disregarded
|  |  | -- | 34.9
| 2010-03-30 | Composite | 0.017
|  |  | -- | 20.2 | 18.5 | 16.2

## FTHg Field RPD
Appendix B. Field replicates of total mercury in water samples, Watershed Sources Project, Sinclair Inlet, Washinton, February through July, 2008.
[OUBT, Operable Unit B Terrestrial; WWTP, waste water treatment plant; T, Temporary Well; (see Appendix A for complete description of station names and field Identifiers); Abbreviations: ng/L, nanograms per liter]
National Water Quality Laboratory |  |  |  |  | Wisconsin Mercury Research Laboratory
Field Identifier | Date | Time | Filtered total mercury (ng/L) | Relative percent difference | Field Identifier | Date | Time | Filtered total mercury (ng/L) | Relative percent difference
OUBT-722 | 2008-02-01 | 1300 | 453 | -- | OUBT-722 | 2008-04-23 | 1030 | 581 | --
|  | 1301 | 423 | 6.8493 |  |  | 1031 | 501 | 14.7874
[5 empty] | OUBT-724 | 2008-04-24 | 1200 | 6.52 | --
OUBT-722M | 2008-05-06 | 1440 | 342 | -- |  |  | 1201 | 7.80 | 17.8771
|  | 1441 | 328 | 4.1791 | Steam Plant | 2008-05-28 | 1350 | 108 | --
| 2008-05-06 | 1640 | 363 | -- |  |  | 1351 | 104 | 3.7736
|  | 1641 | 329 | 9.8266 |  | 2008-06-24 | 1220 | 15.8 | --
[7 empty] | 1221 | 15.1 | 4.5307
OUBT-722 | 2008-06-04 | 1300 | 2090 | -- | Bremerton WWTP | 2008-07-15 | 0650 | 1.06 | --
|  | 1301 | 1060 | 65.3968 |  |  | 0651 | .97 | 8.867
| 2008-06-05 | 930 | 2190 | --
|  | 0931 | 1950 | 11.5942 | T-1 | 2008-06-04 | 1500 | 1.86 | --
| 2008-06-05 | 1230 | 1910 | -- |  |  | 1501 | 2.11 | 12.5945
|  | 1231 | 1760 | 8.1744 |  | 2008-06-05 | 1220 | 1.78 | --
[7 empty] | 1221 | 1.57 | 12.5373
[5 empty] | T-2 | 2008-06-05 | 1430 | .39 | --
[7 empty] | 1431 | .36 | 8
[5 empty] | T-3 | 2008-06-04 | 1445 | .81 | --
[7 empty] | 1446 | .97 | 17.9775
[6 empty] | 2008-06-05 | 1350 | 1.45 | --
[7 empty] | 1351 | 1.72 | 17.0347
[5 empty] | T-4 | 2008-06-05 | 1410 | 2.33 | --
[7 empty] | 1411 | 2.15 | 8.0357
[5 empty] | Anderson | 2008-07-15 | 1200 | 0.57 | --
[7 empty] | 1201 | 0.34 | 50.5495
|  |  | Minimum | 4.2 |  |  |  | Minimum | 3.8
|  |  | Median | 9 |  |  |  | Median | 12.6
|  |  | Maximum | 65 |  |  |  | Maximum | 51

## FMHg Field RPD
Appendix B. Field replicates of methylmercury in water samples, Watershed Sources Project, Sinclair Inlet, Washington, April through July, 2008.
[OUBT, Operable Unit B Terrestrial; WS WWTP, West Sound waste water treatment plant, (see Appendix A for complete description of station names and field Identifiers). Abbreviations: ng/L, nanograms per liter; E, estimated; I, identical value at the lowest reported digit; --, not applicable]
Field Identifier | Date | Time | Filtered methyl-mercury (ng/L) | Average Filtered methyl-mercury (ng/L) | Relative percent difference
OUBT-722 | 2008-04-23 | 1030 | 0.1 | -- | --
|  | 1031 | 0.07 | 0.085 | 35.2941
OUBT-718 | 2008-04-21 | 1630 | E.05 | -- | --
|  | 1631 | <.04 | -- | --
Steam Plant | 2008-04-25 | 1030 | <.04 | -- | --
|  | 1031 | <.04 | -- | --
WS WWTP | 2008-07-21 | 810 | 0.15 | -- | --
|  | 811 | 0.15 | 0.15 | I

## PTHg RPD
Appendix B. Field replicates of particulate mercury in water samples, Watershed Sources Project, Sinclair Inlet, Washington, April through July, 2008.
[OUBT, Operable Unit B Terrestrial; WS WWTP, West Sound waste water treatment plant (see Appendix A for complete description of station names and field Identifiers). Abbreviations: ng/L; nanograms per liter ]
Field Identifier | Date | Time | Particulate total mercury (ng/L) | Relative percent difference
OUBT-722 | 2008-04-23 | 1030 | 56 | --
|  | 1031 | 57 | 1.7699
Steam Plant | 2008-04-25 | 1030 | 33.1 | --
|  | 1031 | 31.9 | 3.6923
| 2008-05-28 | 1350 | 22.8 | --
|  | 1351 | 24.9 | 8.805
WS WWTP | 2008-05-08 | 0810 | 12.5 | --
|  | 0811 | 7.84 | 45.821
Annapolis | 2008-05-09 | 1330 | 0.481 | --
|  | 1331 | 0.636 | 27.7529
| 2008-07-15 | 1200 | 0.926 | --
|  | 1201 | 0.696 | 28.36

## TSS RPD
Appendix B. Laboratory and field replicates of total suspended solids in water samples, Watershed Sources Project, Sinclair Inlet, Washington, January 2008 through January 2009.
[OUBT, Operable Unit B Terrestrial; OUNSC, Operable Unit Naval Supply Center; DD, Dry Dock; WS, West Sound; WWTP, waste water treatment plant; PSNS, Puget Sound Naval Shipyard (see Appendix A for complete description of station names and field Identifiers). Abbreviations: mg/L, milligrams per liter; Rep, replicate]
Total suspended solids (mg/L)
Field Identifier | Date | Time | Rep 1 | Rep 2 | Mean | Relative Percent Difference
Surface Water
Laboratory replicates
Steam Plant | 2008-03-26 | 1400 | 1.45 | 1.54 | 1.5 | 6
| 2008-05-28 | 1350 | 1.39 | 1.38 | 1.39 | 0.7
Bremerton WWTP | 2008-05-08 | 625 | 3.13 | 4.08 | 3.61 | 26.3
WS WWTP | 2008-05-08 | 0810 | 29.85 | 32.18 | 31.02 | 7.5
| 2008-07-21 | 0810 | 12.69 | 14.25 | 13.47 | 11.6
Navy City | 2009-01-07 | 2120 | 362.8 | 334.8 | 348.8 | 8
Field replicates
Anderson | 2008-07-15 | 1200 | 3.85 | 3.52 | 3.69 | 8.9
PSNS 015 | 2009-01-07 | 2110 | 151.8 | 139.5 | 145.65 | 8.4
Navy City | 2009-01-08 | 2120 | 348.8 | 398.7 | 373.75 | 13.4
Groundwater
Laboratory replicates
OUBT-715R | 2008-05-06 | 1330 | 9.4 | 9.58 | 9.49 | 1.9
| 2008-05-06 | 2110 | 4.62 | 11.12 | 7.87 | 82.6
OUBT-718 | 2008-02-01 | 1500 | 2.7 | 2.81 | 2.76 | 4
OUBT-722 | 2008-02-01 | 1300 | 1.34 | 0.51 | 0.93 | 89.2
DD-5 | 2008-03-26 | 1100 | 1.31 | 1.53 | 1.42 | 15.5
| 2008-06-24 | 1030 | 43.62 | 52.85 | 48.24 | 19.1
Field replicates
OUA-204 | 2008-01-28 | 1250 | 0.21 | 0.16 | 0.19 | 26.3
OUNSC-311 | 2008-01-31 | 1200 | 76.79 | 93.75 | 85.27 | 19.9
OUBT-722 | 2008-02-01 | 1300 | 1.34 | 1.69 | 1.52 | 23

## Ancillary Field Blanks
Appendix B. Selected constituents in field blank samples, Watershed Sources Project, Sinclair Inlet, Washington, April through July 2008.
[OUBT, Operable Unit B Terrestrial; OUNSC (see Appendix A for complete description of station names and field Identifiers). Abbreviations; N, nitrogen; P, phosphorus; SiO2, silicon dioxide; mg/L, milligram per liter; µg/L, microgram per liter; E, estimated; <, less than; -, not applicable].
|  |  | Nutrients |  |  |  |  | Trace metals |  | Carbon and Nitrogen
Field Identifier | Date | Time | Ammonia (mg/L as N) | Nitrate plus nitrite (mg/L as N) | Nitrite (mg/L as N) | Ortho-phosphate (mg/L as P) | Silica (mg/L as SiO2) | Manganese (µg/L) | Iron (µg/L) | Dissolved organic carbon (mg/L) | Total particulate carbon (mg/L) | Particulate inorganic carbon (mg/L) | Particulate organic carbon (mg/L) | Total particulate nitrogen (mg/L)
OUBT-709 | 2008-04-24 | 1538 | 0.002* | 0.002* | <.001* | <.001* | 0.033* | <.4 | <8 | E.35 | <.14 | <.04 | <.12 | <.04
Annapolis | 2008-07-15 | 958 | - | - | - | - | - | - | - | <0.4 | <.14 | <.04 | <.12 | <.04
* analysis perfomed at UW Oceanography laboratory

## Ancillary Field RPD
Appendix B. Field replicates of selected constituents in water samples, Watershed Sources Project, Sinclair Inlet, Washington, April through July 2008.
[OUBT, Operable Unit B terrestrial (see Appendix A for complete description of station names and field identifers). Abbreviations: Rep, replicate; N, nitrogen; P, phosphorus; SiO2, silica dioxide; mg/L, milligram per liter; µg/L, microgram per liter; E, estimated; <, less than; -, not applicable; I, duplicates values are identical to lowest digit of the reporting level].
OUBT-724 04/24/08 1200 | [5 empty] | Anderson 07/15/08 1200
Analyte | Unit | Rep 1 | Rep 2 | Mean | Relative percent difference | Analyte | Unit | Rep 1 | REP 2 | Mean | Relative percent difference
Trace metals
Manganese | (µg/L) | 0.411 | <8 | - | - | Manganese | (µg/L) | 1.298 | 1.328 | 1.313 | I
Iron | (µg/L) | <8 | <8 | <8 | - | Iron | (µg/L) | 7.63 | 8.014 | 7.822 | I
Filtered nutrients
Ammonia | (mg/L as N) | 3.199* | 4.068* | 3.64 | 23.9163 | Ammonia | (mg/L as N) | 0.04561 | E.010 | 0.02802 | -
Nitrate plus nitrite | (mg/L as N) | 1.867* | 1.38* | 1.63 | 30.2 | Nitrate plus nitrite | (mg/L as N) | 0.218 | 0.227 | 0.2225 | 4.0449
Nitrite | (mg/L as N) | 0.183* | 0.172* | 0.1775 | 6.1972 | Nitrite | (mg/L as N) | <.002 | <.002 | 0.001645 | -
Orthophosphate | (mg/L as P) | 1.843* | 1.761* | 1.802 | 4.4 | Orthophosphate | (mg/L as P) | 0.03629 | 0.03325 | 0.03477 | 8.7432
Silica | (mg/L as SiO2) | 24.943* | 23.088* | 24.0155 | 7.5 | Silica | (mg/L as SiO2) | 24.296 | 24.11 | 24.203 | 0.7685
Toal particulate carbon and nitrogen
Dissolved organic carbon | (mg/L) | 0.9875 | 0.8755 | 1 | 10.5 | Dissolved organic carbon | (mg/L) | 0.6451 | 0.6132 | 0.62915 | I
Total particulate carbon | (mg/L) | <.14 | <.14 | <.14 | - | Total particulate carbon | (mg/L) | 0.407 | 0.412 | 0.4095 | I
Particulate inorganic carbon | (mg/L) | <.04 | <.04 | <.04 | - | Particulate inorganic carbon | (mg/L) | <.04 | <.04 | <.04 | -
Particulate organic carbon | (mg/L) | <.12 | <.12 | <.12 | - | Particulate organic carbon | (mg/L) | 0.407 | 0.412 | 0.4095 | I
Total particulate nitrogen | (mg/L) | <.04 | <.04 | <.04 | - | Total particulate nitrogen | (mg/L) | <.04 | <.04 | <.04 | -
Total phosphorus and nitrogen
[6 empty] | Phosphorus filtered | (mg/L as P) | 0.0265 | 0.0263 | 0.027 | 3.8
[6 empty] | Phosphorus unfiltered | (mg/L as P) | 0.0422 | 0.0383 | 0.04025 | 10
[6 empty] | Total Nitrogen filtered | (mg/L) | 0.26 | 0.243 | 0.2515 | 8
[6 empty] | Total Nitrogen unfiltered | (mg/L) | 0.25 | 0.255 | 0.26 | 3.9
Major Ions
[6 empty] | Calcium | (mg/L) | 8.913 | 8.76 | 8.8365 | 1.7315
[6 empty] | Magnesium | (mg/L) | 4.559 | 4.438 | 4.4985 | 2.6898
[6 empty] | Potassium | (mg/L) | 0.525 | 0.491 | 0.508 | 7.8
[6 empty] | Sodium | (mg/L) | 4.169 | 3.905 | 4.037 | 6.4
[6 empty] | Chloride | (mg/L) | 2.626 | 2.023 | 2.33 | 26.2
[6 empty] | Fluoride | (mg/L) | <.12 | <.12 | <.12 | -
[6 empty] | Sulfate | (mg/L) | 1.44 | 1.455 | 1.4475 | 1.4
* analysis perfomed at UW Oceanography laboratory