## t15_esipos_response
Table 15. Summary of response factor and parameters of linear fit of response factor by nominal concentration for positive electrospray ionization analytes [Data are from ESI+ 7 method detection level (MDL) experiments with 10 concentration levels and 7 replicates at each level. The number of calibration standards included in the summary was different for each analyte because of different response factors and the corresponding lowest calibration standards that met the identification criteria. Slope of zero indicates response factors are constant over the concentration range. Slopes that are significantly different from zero are in bold and highlighted in red. Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; ESI+, positive electrospray ionization; USGS, U.S. Geological Survey; ng/L, nanograms per liter; N, number of samples; t, Student's t-distribution; Prob>, probably greater then; Lower 95 percent, lower 95th percentile confidence limit of slope estimate; Upper 95 percent, upper 95th percentile confidence limit of slope estimate; na, not available] | [14 empty] | These columns are not printed
Validation quality code | Analytical method group | Pesticide group | Analyte | USGS Parameter code | MDL (ng/L) | N | Median response factor | Mean response factor | Slope estimate | Std Error | t Ratio | Prob>¦t¦ | Lower 95% | Upper 95% | Slope contains zero
A | acetanilide & amide | Acetochlor | Acetochlor | 68520 | 5 | 70 | 159.1123 | 173.2443 | -0.0011246 | 0.0047365 | -0.24 | 0.813 | -0.010576 | 0.0083269 |  | -0.011 | 0.008
A |  |  | Hydroxyacetochlor | 68615 | 4 | 63 | 89.5288 | 88.4866 | 0.001661 | 0.0007617 | 2.18 | 0.0331 | 0.00013794 | 0.0031841 |  | 0 | 0.003
A |  | Acetochlor/Metolachlor | 2-chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 2 | 63 | 316.0019 | 319.6958 | -0.00075038 | 0.0014812 | -0.51 | 0.6142 | -0.0037121 | 0.0022114 |  | -0.004 | 0.002
A |  | Alachlor | 2-Chloro-2,6-diethylacetanilide | 68525 | 2 | 70 | 343.1588 | 335.0234 | 0.0032428 | 0.0022822 | 1.42 | 0.1599 | -0.0013113 | 0.0077969 |  | -0.001 | 0.008
A |  |  | Alachlor | 65064 | 4 | 63 | 241.2146 | 312.6136 | -0.014973 | 0.0058765 | -2.55 | 0.0134 | -0.026724 | -0.0032226 | 1 | -0.027 | -0.003
A |  |  | Hydroxyalachlor | 68616 | 4 | 56 | 147.0367 | 149.4508 | 0.0019495 | 0.0011333 | 1.72 | 0.0911 | -0.00032266 | 0.0042216 |  | 0 | 0.004
A |  | Dimethenamid | Dimethenamid | 68580 | 1 | 70 | 1202.6371 | 1182.2778 | 0.022475 | 0.0055024 | 4.08 | 0.0001 | 0.011495 | 0.033455 | 1 | 0.011 | 0.033
A |  | Metolachlor | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 2 | 63 | 612.46 | 594.6938 | 0.011429 | 0.0033614 | 3.4 | 0.0012 | 0.0047072 | 0.01815 | 1 | 0.005 | 0.018
A |  |  | Dechlorometolachlor | 68562 | 1 | 70 | 4781.6432 | 4736.5233 | -0.020723 | 0.019835 | -1.04 | 0.2998 | -0.060304 | 0.018858 |  | -0.06 | 0.019
A |  |  | Hydroxymetolachlor | 68622 | 1 | 70 | 1776.8556 | 1774.3663 | 0.0080488 | 0.0076669 | 1.05 | 0.2975 | -0.0072503 | 0.023348 |  | -0.007 | 0.023
A |  |  | Metolachlor | 65090 | 4 | 70 | 2068.5028 | 4091.4678 | -0.37551 | 0.1928 | -1.95 | 0.0556 | -0.76024 | 0.0092156 |  | -0.76 | 0.009
A |  |  | Metolachlor hydroxy morpholinone | 68649 | 3 | 63 | 340.6136 | 336.4 | 0.0057465 | 0.0019246 | 2.99 | 0.0041 | 0.001898 | 0.0095949 | 1 | 0.002 | 0.01
A |  | Pronamide | Pronamide | 67706 | 2 | 70 | 293.3998 | 281.8273 | 0.0070211 | 0.0022231 | 3.16 | 0.0024 | 0.0025851 | 0.011457 | 1 | 0.003 | 0.011
A |  | Propanil | Propanil | 66641 | 5 | 56 | 94.906 | 93.9482 | 0.00017514 | 0.00071411 | 0.25 | 0.8072 | -0.0012566 | 0.0016068 |  | -0.001 | 0.002
A | acid | Bentazone | 2-Amino-N-isopropylbenzamide | 68503 | 2 | 69 | 790.4767 | 789.5324 | 0.0010169 | 0.0048318 | 0.21 | 0.8339 | -0.0086274 | 0.010661 |  | -0.009 | 0.011
A | carbamate & thiocarbamate | Aldicarb | Aldicarb | 68528 | 4 | 63 | 164.8215 | 160.7341 | 0.0024964 | 0.0014025 | 1.78 | 0.0801 | -0.00030806 | 0.0053009 |  | 0 | 0.005
A |  |  | Aldicarb sulfone | 68529 | 10 | 69 | 9.6813 | 10.2512 | -6.3275e-05 | 0.00012759 | -0.5 | 0.6216 | -0.00031795 | 0.0001914 |  | 0 | 0
A |  |  | Aldicarb sulfoxide | 68530 | 1 | 70 | 308.6533 | 310.1081 | 0.0050252 | 0.0015164 | 3.31 | 0.0015 | 0.0019992 | 0.0080512 | 1 | 0.002 | 0.008
A |  | Butylate | Butylate | 65068 | 3 | 56 | 83.8675 | 81.6938 | 0.001594 | 0.00051227 | 3.11 | 0.003 | 0.00056697 | 0.002621 | 1 | 0.001 | 0.003
A |  | Carbaryl | Carbaryl | 65069 | 3 | 70 | 266.6861 | 264.8634 | -0.0029101 | 0.0021587 | -1.35 | 0.1821 | -0.0072177 | 0.0013975 |  | -0.007 | 0.001
A |  | Carbofuran | Carbofuran | 65070 | 2 | 70 | 1011.4129 | 1004.7074 | 0.010641 | 0.0039997 | 2.66 | 0.0097 | 0.0026597 | 0.018622 | 1 | 0.003 | 0.019
A |  | EPTC | EPTC | 65080 | 103 | 49 | 14.2818 | 15.1378 | 3.4652e-05 | 0.00022578 | 0.15 | 0.8787 | -0.00041956 | 0.00048886 |  | 0 | 0
A |  |  | EPTC R248722 | 68594 | 2 | 70 | 463.739 | 449.0929 | 0.0034164 | 0.0032136 | 1.06 | 0.2915 | -0.0029962 | 0.0098291 |  | -0.003 | 0.01
A |  | Methomyl | Methomyl | 68645 | 1 | 70 | 297.7084 | 293.8078 | 0.006675 | 0.001818 | 3.67 | 0.0005 | 0.0030472 | 0.010303 | 1 | 0.003 | 0.01
A |  |  | Methomyl oxime | 68646 | 39 | 28 | 3.4395 | 3.4238 | 6.2305e-06 | 2.0524e-05 | 0.3 | 0.7639 | -3.5957e-05 | 4.8418e-05 |  | 0 | 0
A |  | Molinate | 4-Hydroxy molinate | 68515 | 3 | 70 | 381.9381 | 423.794 | -0.0058578 | 0.0096692 | -0.61 | 0.5467 | -0.025152 | 0.013437 |  | -0.025 | 0.013
A |  |  | Carboxy molinate | 68549 | 11 | 42 | 138.5721 | 137.0083 | 0.0014254 | 0.00063141 | 2.26 | 0.0295 | 0.00014929 | 0.0027016 |  | 0 | 0.003
A |  |  | Molinate | 65091 | 9 | 49 | 48.6489 | 49.159 | 0.00078243 | 0.00034458 | 2.27 | 0.0278 | 8.9222e-05 | 0.0014756 |  | 0 | 0.001
A |  | Oxamyl | Oxamyl | 68664 | 1 | 70 | 529.0604 | 534.991 | 0.0092811 | 0.0019431 | 4.78 | <.0001 | 0.0054037 | 0.013158 | 1 | 0.005 | 0.013
A |  |  | Oxamyl oxime | 68665 | 1 | 70 | 1299.051 | 1305.702 | -0.00090896 | 0.0048668 | -0.19 | 0.8524 | -0.010621 | 0.0088026 |  | -0.011 | 0.009
A |  | Propoxur | Propoxur | 68679 | 2 | 70 | 721.1274 | 737.5324 | -0.00056755 | 0.0046873 | -0.12 | 0.904 | -0.009921 | 0.0087859 |  | -0.01 | 0.009
A |  | thiobencarb | 4-Chlorobenzylmethyl sulfoxide | 68514 | 2 | 70 | 517.5731 | 510.8344 | 0.0040471 | 0.0037167 | 1.09 | 0.28 | -0.0033694 | 0.011464 |  | -0.003 | 0.011
A |  | Thiobencarb | Thiobencarb | 65107 | 1 | 70 | 824.9773 | 814.6375 | 0.013389 | 0.0037913 | 3.53 | 0.0007 | 0.0058238 | 0.020955 | 1 | 0.006 | 0.021
A |  | Triallate | Triallate | 68710 | 8 | 63 | 37.4311 | 41.0812 | -0.00032103 | 0.00055782 | -0.58 | 0.5671 | -0.0014365 | 0.0007944 |  | -0.001 | 0.001
A | fungicide | Azoxystrobin | Azoxystrobin | 66589 | 1 | 70 | 3448.5322 | 3437.7762 | -0.011584 | 0.014647 | -0.79 | 0.4318 | -0.040811 | 0.017644 |  | -0.041 | 0.018
A |  | Benomyl | 2-Aminobenzimidazole | 68502 | 4 | 49 | 119.1874 | 118.7745 | 0.0027549 | 0.00073126 | 3.77 | 0.0005 | 0.0012838 | 0.004226 | 1 | 0.001 | 0.004
A |  | Fentin | Fentin | 68603 | 8 | 49 | 51.6315 | 52.1938 | -0.00082529 | 0.0002401 | -3.44 | 0.0012 | -0.0013083 | -0.00034228 |  | -0.001 | 0
A |  | Kresoxim-methyl | Kresoxim-methyl | 67670 | 2 | 70 | 258.1108 | 257.5439 | 0.00090999 | 0.0015677 | 0.58 | 0.5635 | -0.0022183 | 0.0040383 |  | -0.002 | 0.004
A |  | Metalaxyl | Metalaxyl | 68437 | 3 | 70 | 346.7088 | 400.0765 | -0.0085309 | 0.0061299 | -1.39 | 0.1686 | -0.020763 | 0.0037012 |  | -0.021 | 0.004
A |  | Metconazole | Metconazole | 66620 | 2 | 63 | 321.2105 | 323.1714 | -0.0020436 | 0.0014995 | -1.36 | 0.1779 | -0.005042 | 0.00095476 |  | -0.005 | 0.001
A |  | Myclobutanil | Myclobutanil | 66632 | 3 | 63 | 173.9619 | 179.3699 | 0.00018047 | 0.001517 | 0.12 | 0.9057 | -0.0028529 | 0.0032139 |  | -0.003 | 0.003
A |  | Propiconazole | Propiconazole | 66643 | 3 | 70 | 198.7424 | 244.1069 | -0.0085185 | 0.0050005 | -1.7 | 0.093 | -0.018497 | 0.0014599 |  | -0.018 | 0.001
A |  | Pyraclostrobin | Pyraclostrobin | 66646 | 1 | 70 | 757.4433 | 770.1976 | 0.005372 | 0.0041937 | 1.28 | 0.2046 | -0.0029964 | 0.01374 |  | -0.003 | 0.014
A |  | Tebuconazole | Tebuconazole | 66649 | 2 | 70 | 279.153 | 305.8289 | -0.0063223 | 0.0031581 | -2 | 0.0493 | -0.012624 | -2.0364e-05 |  | -0.013 | 0
A |  | Tetraconazole | Tetraconazole | 66654 | 3 | 70 | 213.9377 | 238.1124 | -0.0032383 | 0.0031296 | -1.03 | 0.3045 | -0.0094833 | 0.0030067 |  | -0.009 | 0.003
A |  | Trifloxystrobin | Trifloxystrobin | 66660 | 1 | 70 | 2113.844 | 2225.195 | -0.019234 | 0.017393 | -1.11 | 0.2727 | -0.053941 | 0.015474 |  | -0.054 | 0.015
A | miscellaneous | Bromacil | Bromacil | 68542 | 3 | 63 | 279.292 | 278.4028 | -0.0046611 | 0.0018389 | -2.53 | 0.0138 | -0.0083383 | -0.0009839 | 1 | -0.008 | -0.001
A |  | Etoxazole | Etoxazole | 68598 | 1 | 70 | 1813.3121 | 1835.3778 | 0.01065 | 0.007055 | 1.51 | 0.1358 | -0.0034276 | 0.024728 |  | -0.003 | 0.025
A |  | Flubendiamide | Desiodo flubendiamide | 68563 | 2 | 70 | 380.4346 | 408.8837 | -0.0040522 | 0.0039211 | -1.03 | 0.3051 | -0.011877 | 0.0037723 |  | -0.012 | 0.004
A |  | Imazamox | Imazamox | 68625 | 14 | 49 | 58.6171 | 52.7516 | 0.0011965 | 0.00074919 | 1.6 | 0.117 | -0.00031072 | 0.0027036 |  | 0 | 0.003
A |  | Imazaquin | Imazaquin | 61682 | 9 | 63 | 112.1412 | 106.0721 | 0.0020542 | 0.0013373 | 1.54 | 0.1297 | -0.00061997 | 0.0047284 |  | -0.001 | 0.005
A |  | Imazethapyr | Imazethapyr | 61683 | 10 | 56 | 137.3196 | 132.5489 | 0.0004976 | 0.00118 | 0.42 | 0.6749 | -0.0018681 | 0.0028633 |  | -0.002 | 0.003
A |  | Imidacloprid | Imidacloprid | 68426 | 5 | 49 | 84.7991 | 84.8259 | 0.0011484 | 0.00044589 | 2.58 | 0.0132 | 0.0002514 | 0.0020454 |  | 0 | 0.002
A |  | Indoxacarb | Indoxacarb | 68627 | 4 | 56 | 45.3532 | 44.8426 | 0.00069797 | 0.00020493 | 3.41 | 0.0013 | 0.0002871 | 0.0011088 |  | 0 | 0.001
A |  | Isoxaflutole | Isoxaflutole | 68632 | 6 | 70 | 77.3594 | 85.9869 | -0.00055151 | 0.0023061 | -0.24 | 0.8117 | -0.0051532 | 0.0040502 |  | -0.005 | 0.004
A |  | Lactofen | Lactofen | 68638 | 1 | 63 | 356.6775 | 355.5865 | 0.0048343 | 0.0014883 | 3.25 | 0.0019 | 0.0018583 | 0.0078103 | 1 | 0.002 | 0.008
A |  | Methoxyfenozide | Methoxyfenozide | 68647 | 1 | 70 | 1571.9653 | 1551.0744 | 0.021182 | 0.0060702 | 3.49 | 0.0009 | 0.0090688 | 0.033295 | 1 | 0.009 | 0.033
A |  | Norflurazon | Demethyl norflurazon | 68567 | 2 | 70 | 636.3158 | 739.1801 | -0.022623 | 0.010656 | -2.12 | 0.0374 | -0.043886 | -0.0013598 | 1 | -0.044 | -0.001
A |  |  | Norflurazon | 67685 | 2 | 70 | 573.1133 | 562.1703 | 0.0068972 | 0.002743 | 2.51 | 0.0143 | 0.0014236 | 0.012371 | 1 | 0.001 | 0.012
A |  | Pendimethalin | Pendimethalin | 65098 | 1 | 56 | 237.7889 | 237.1734 | 0.0028929 | 0.000866 | 3.34 | 0.0015 | 0.0011567 | 0.0046292 | 1 | 0.001 | 0.005
A |  | Propargite | Propargite | 68677 | 1 | 70 | 2445.9585 | 2413.5417 | -0.018655 | 0.0083931 | -2.22 | 0.0296 | -0.035403 | -0.0019066 | 1 | -0.035 | -0.002
A |  | Pymetrozine | Pymetrozine | 68680 | 1 | 70 | 408.1934 | 406.1785 | 0.0057822 | 0.0017058 | 3.39 | 0.0012 | 0.0023783 | 0.0091861 | 1 | 0.002 | 0.009
A |  | Pyridaben | Pyridaben | 68682 | 1 | 70 | 1950.3804 | 1942.2588 | -0.0039344 | 0.0084506 | -0.47 | 0.643 | -0.020797 | 0.012928 |  | -0.021 | 0.013
A |  | Pyriproxyfen | Pyriproxyfen | 68683 | 1 | 70 | 1400.593 | 1408.7041 | 0.0014121 | 0.0046762 | 0.3 | 0.7636 | -0.0079191 | 0.010743 |  | -0.008 | 0.011
A |  | Tebufenozide | Tebufenozide | 68692 | 1 | 70 | 2185.1616 | 2170.56 | 0.010439 | 0.0092926 | 1.12 | 0.2652 | -0.0081037 | 0.028982 |  | -0.008 | 0.029
A |  | Terbacil | Terbacil | 68698 | 10 | 49 | 56.4077 | 57.6592 | -0.0013693 | 0.00032989 | -4.15 | 0.0001 | -0.0020329 | -0.00070564 | 1 | -0.002 | -0.001
A | organophosphate | Acephate | Acephate | 68519 | 5 | 70 | 450.1968 | 449.4792 | 0.0043421 | 0.0019853 | 2.19 | 0.0322 | 0.00038039 | 0.0083037 |  | 0 | 0.008
A |  | Azinphos-methyl | Azinphos-methyl | 65066 | 4 | 70 | 167.7614 | 179.6757 | -0.0028174 | 0.0021644 | -1.3 | 0.1974 | -0.0071364 | 0.0015015 |  | -0.007 | 0.002
A |  |  | Azinphos-methyl oxon | 68211 | 7 | 49 | 183.0077 | 178.7282 | 0.00022878 | 0.0011877 | 0.19 | 0.8481 | -0.0021605 | 0.0026181 |  | -0.002 | 0.003
A |  | Chlorpyrifos | Chlorpyrifos | 65072 | 2 | 70 | 248.4274 | 251.5324 | 0.0022037 | 0.0016079 | 1.37 | 0.175 | -0.0010047 | 0.0054122 |  | -0.001 | 0.005
A |  |  | Chlorpyrifos oxon | 68216 | 2 | 70 | 385.0442 | 369.4766 | 0.006049 | 0.0027755 | 2.18 | 0.0328 | 0.00051054 | 0.011588 | 1 | 0.001 | 0.012
A |  | Diazinon | Diazinon | 65078 | 1 | 70 | 645.6313 | 664.419 | 0.0033739 | 0.0040061 | 0.84 | 0.4026 | -0.0046201 | 0.011368 |  | -0.005 | 0.011
A |  |  | Hydroxydiazinon | 68618 | 5 | 56 | 245.5184 | 244.882 | 0.0042271 | 0.0013661 | 3.09 | 0.0031 | 0.0014882 | 0.006966 | 1 | 0.001 | 0.007
A |  |  | Pyrimidinol | 68505 | 10 | 49 | 639.6864 | 1497.6269 | -0.18211 | 0.065857 | -2.77 | 0.0081 | -0.31459 | -0.049618 | 1 | -0.315 | -0.05
A |  | Dicrotophos | Dicrotophos | 68573 | 2 | 56 | 328.9452 | 328.4265 | 0.0059706 | 0.0012614 | 4.73 | <.0001 | 0.0034418 | 0.0084995 | 1 | 0.003 | 0.008
A |  | Dimethoate | Dimethoate | 66596 | 1 | 70 | 530.4864 | 527.3823 | 0.0075622 | 0.0029113 | 2.6 | 0.0115 | 0.0017529 | 0.013372 | 1 | 0.002 | 0.013
A |  |  | Dimethoate oxon | 68661 | 1 | 70 | 403.4694 | 399.2159 | 0.0059538 | 0.0021548 | 2.76 | 0.0074 | 0.001654 | 0.010254 | 1 | 0.002 | 0.01
A |  | Disulfoton | Disulfoton | 67595 | 6 | 49 | 47.7502 | 47.0956 | 0.00074674 | 0.00026843 | 2.78 | 0.0078 | 0.00020673 | 0.0012867 |  | 0 | 0.001
A |  |  | Disulfoton oxon | 68586 | 1 | 70 | 1167.7125 | 1153.1655 | 0.013532 | 0.0053332 | 2.54 | 0.0135 | 0.00289 | 0.024174 | 1 | 0.003 | 0.024
A |  |  | Disulfoton oxon sulfone | 68588 | 3 | 63 | 264.3717 | 260.4169 | 0.0026535 | 0.0017208 | 1.54 | 0.1282 | -0.0007875 | 0.0060945 |  | -0.001 | 0.006
A |  |  | Disulfoton oxon sulfoxide | 68587 | 3 | 70 | 501.8385 | 462.3649 | 0.0070826 | 0.0057995 | 1.22 | 0.2262 | -0.0044902 | 0.018655 |  | -0.004 | 0.019
A |  |  | Disulfoton sulfone | 68589 | 4 | 56 | 125.6761 | 128.5741 | -0.00036186 | 0.00091722 | -0.39 | 0.6948 | -0.0022008 | 0.0014771 |  | -0.002 | 0.001
A |  |  | Disulfoton sulfoxide | 68590 | 2 | 70 | 665.1872 | 676.8073 | -0.0048912 | 0.0033847 | -1.45 | 0.153 | -0.011645 | 0.0018629 |  | -0.012 | 0.002
A |  | Ethoprop | Ethoprop | 68596 | 2 | 70 | 332.3617 | 350.4678 | -0.0028387 | 0.0036896 | -0.77 | 0.4443 | -0.010201 | 0.0045238 |  | -0.01 | 0.005
A |  |  | O-Ethyl-O-methyl-S-propyl phosphorothioate | 68597 | 2 | 70 | 540.6631 | 553.1543 | -0.0010541 | 0.0038187 | -0.28 | 0.7834 | -0.0086743 | 0.0065661 |  | -0.009 | 0.007
A |  |  | O-ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 2 | 70 | 418.2232 | 425.7075 | 0.0036399 | 0.0035649 | 1.02 | 0.3109 | -0.0034737 | 0.010754 |  | -0.003 | 0.011
A |  |  | O-ethyl-S-propyl phosphorothioate | 68658 | 32 | 35 | 9.4157 | 9.546 | 0.00010747 | 3.3109e-05 | 3.25 | 0.0027 | 4.0112e-05 | 0.00017483 |  | 0 | 0
A |  | Fenamiphos | Fenamiphos | 68599 | 1 | 70 | 1273.9603 | 1271.0887 | 0.0070758 | 0.0051164 | 1.38 | 0.1712 | -0.0031339 | 0.017285 |  | -0.003 | 0.017
A |  |  | Fenamiphos sulfone | 68600 | 2 | 70 | 486.1764 | 483.3912 | 0.0036941 | 0.003274 | 1.13 | 0.2631 | -0.002839 | 0.010227 |  | -0.003 | 0.01
A |  |  | Fenamiphos sulfoxide | 68601 | 2 | 70 | 393.7607 | 417.8553 | -0.0070927 | 0.004601 | -1.54 | 0.1278 | -0.016274 | 0.0020884 |  | -0.016 | 0.002
A |  | Fonofos | Fonofos | 65084 | 5 | 56 | 115.6104 | 117.6099 | 0.0017725 | 0.001045 | 1.7 | 0.0956 | -0.00032263 | 0.0038677 |  | 0 | 0.004
A |  | Malathion | Malaoxon | 68240 | 1 | 70 | 966.8095 | 963.7226 | 0.004899 | 0.0048968 | 1 | 0.3206 | -0.0048724 | 0.01467 |  | -0.005 | 0.015
A |  |  | Malathion | 65087 | 2 | 63 | 396.5732 | 397.0725 | 0.0044853 | 0.0016654 | 2.69 | 0.0091 | 0.001155 | 0.0078156 | 1 | 0.001 | 0.008
A |  | Methamidophos | Methamidophos | 68644 | 1 | 63 | 315.9994 | 318.748 | -0.0004933 | 0.0013444 | -0.37 | 0.7149 | -0.0031817 | 0.0021951 |  | -0.003 | 0.002
A |  | Methidathion | Methidathion | 65088 | 5 | 70 | 420.454 | 412.9815 | 0.0043419 | 0.0022086 | 1.97 | 0.0534 | -6.5235e-05 | 0.0087491 |  | 0 | 0.009
A |  | Naled | Dichlorvos | 68572 | 26 | 49 | 66.7899 | 66.5231 | 0.001129 | 0.00056982 | 1.98 | 0.0534 | -1.7321e-05 | 0.0022753 |  | 0 | 0.002
A |  | Parathion-ethyl | Paraoxon | 68666 | 2 | 70 | 598.4308 | 588.6386 | -0.0032382 | 0.0028901 | -1.12 | 0.2665 | -0.0090052 | 0.0025288 |  | -0.009 | 0.003
A |  | Parathion-methyl | Paraoxon-methyl | 68648 | 9 | 49 | 102.9395 | 102.1846 | -0.0013564 | 0.00063647 | -2.13 | 0.0383 | -0.0026368 | -7.5932e-05 |  | -0.003 | 0
A |  | Phorate | Phorate | 68668 | 3 | 63 | 67.7477 | 67.6432 | 0.0012423 | 0.00059186 | 2.1 | 0.04 | 5.879e-05 | 0.0024258 |  | 0 | 0.002
A |  |  | Phorate oxon | 68669 | 2 | 28 | 645.6385 | 630.4533 | -0.0014424 | 0.0030851 | -0.47 | 0.644 | -0.0077839 | 0.0048991 |  | -0.008 | 0.005
A |  |  | Phorate oxon sulfone | 68670 | 10 | 49 | 110.8919 | 111.8162 | -0.00096098 | 0.00067176 | -1.43 | 0.1592 | -0.0023124 | 0.00039044 |  | -0.002 | 0
A |  |  | Phorate oxon sulfoxide | 68671 | 3 | 56 | 189.8522 | 185.764 | 0.0023054 | 0.0011057 | 2.09 | 0.0418 | 8.863e-05 | 0.0045222 |  | 0 | 0.005
A |  |  | Phorate sulfone | 68672 | 4 | 63 | 85.5332 | 83.8497 | -7.501e-05 | 0.00084325 | -0.09 | 0.9294 | -0.0017612 | 0.0016112 |  | -0.002 | 0.002
A |  |  | Phorate sulfoxide | 68673 | 2 | 70 | 457.1094 | 450.6661 | 0.00096516 | 0.0030919 | 0.31 | 0.7559 | -0.0052047 | 0.007135 |  | -0.005 | 0.007
A |  | Profenofos | Profenofos | 68676 | 1 | 70 | 286.7657 | 287.3398 | 0.0013097 | 0.001597 | 0.82 | 0.415 | -0.0018771 | 0.0044965 |  | -0.002 | 0.004
A |  | Tebupirimfos | Tebupirimfos | 68693 | 1 | 70 | 1262.6916 | 1257.7303 | 0.024033 | 0.0051733 | 4.65 | <.0001 | 0.01371 | 0.034356 | 1 | 0.014 | 0.034
A |  |  | Tebupirimfos oxon | 68694 | 1 | 70 | 1721.5633 | 1708.248 | 0.0099737 | 0.007481 | 1.33 | 0.1869 | -0.0049544 | 0.024902 |  | -0.005 | 0.025
A |  | Terbufos | Terbufos | 68699 | 2 | 63 | 206.1371 | 201.5727 | 0.0044605 | 0.0011079 | 4.03 | 0.0002 | 0.0022451 | 0.0066759 | 1 | 0.002 | 0.007
A |  |  | Terbufos oxon | 68700 | 2 | 70 | 695.0723 | 667.0432 | 0.0077798 | 0.0054534 | 1.43 | 0.1583 | -0.0031022 | 0.018662 |  | -0.003 | 0.019
A |  |  | Terbufos oxon sulfone | 68701 | 5 | 70 | 281.1739 | 260.3474 | 0.0069002 | 0.0034906 | 1.98 | 0.0521 | -6.5166e-05 | 0.013866 |  | 0 | 0.014
A |  |  | Terbufos oxon sulfoxide | 68702 | 2 | 69 | 780.6609 | 761.1323 | 0.0090766 | 0.0046031 | 1.97 | 0.0528 | -0.00011116 | 0.018264 |  | 0 | 0.018
A |  |  | Terbufos sulfone | 68703 | 5 | 49 | 65.9638 | 65.4148 | 0.00015035 | 0.0003495 | 0.43 | 0.669 | -0.00055276 | 0.00085345 |  | -0.001 | 0.001
A |  |  | Terbufos sulfoxide | 68704 | 1 | 70 | 1032.2929 | 1025.8399 | -0.0042419 | 0.003875 | -1.09 | 0.2775 | -0.011974 | 0.0034907 |  | -0.012 | 0.003
A |  | Tribufos | Tribufos | 68711 | 1 | 70 | 889.0027 | 878.8281 | -0.001428 | 0.0033507 | -0.43 | 0.6713 | -0.0081142 | 0.0052583 |  | -0.008 | 0.005
A | pyrethroid & organochlorine& phenylpyrazine | Permethrin | cis-Permethrin | 68769 | 1 | 35 | 5.3503 | 8.8316 | na | na | na | na | na |  | #VALUE! | #VALUE! | 0
A |  |  | trans-Permethrin | 68708 | 2 | 35 | 5.3503 | 8.8282 | na | na | na | na | na |  | #VALUE! | #VALUE! | 0
A |  | Piperonyl butoxide | Piperonyl butoxide | 65102 | 2 | 70 | 3803.8071 | 4334.2602 | -0.18341 | 0.055791 | -3.29 | 0.0016 | -0.29474 | -0.072085 | 1 | -0.295 | -0.072
A | sulfonylurea & urea | Chlorimuron-ethyl | Chlorimuron-ethyl | 68872 | 4 | 49 | 4.8855 | 4.9647 | 1.9971e-06 | 5.8812e-05 | 0.03 | 0.9731 | -0.00011632 | 0.00012031 |  | 0 | 0
A |  | Chlorsulfuron | Chlorsulfuron | 61678 | 25 | 28 | 2.6341 | 2.5439 | 5.6617e-05 | 2.1479e-05 | 2.64 | 0.014 | 1.2466e-05 | 0.00010077 |  | 0 | 0
A |  | Diflubenzuron | Diflubenzuron | 68576 | 3 | 70 | 219.9428 | 240.4832 | -0.0055883 | 0.0029494 | -1.89 | 0.0624 | -0.011474 | 0.00029715 |  | -0.011 | 0
A |  | Diflufenzopyr | Hydroxyphthalazinone | 68623 | 23 | 49 | 27.8938 | 27.9662 | 0.00035322 | 0.00018667 | 1.89 | 0.0646 | -2.2304e-05 | 0.00072874 |  | 0 | 0.001
A |  |  | Phthalazinone | 68675 | 3 | 70 | 57.4829 | 57.6744 | 0.00066743 | 0.00050278 | 1.33 | 0.1888 | -0.00033586 | 0.0016707 |  | 0 | 0.002
A |  | Diuron | 3,4-Dichlorophenylurea | 68226 | 72 | 56 | 40.7428 | 42.0992 | -0.00051289 | 0.00051464 | -1 | 0.3234 | -0.0015447 | 0.00051889 |  | -0.002 | 0.001
A |  |  | Diuron | 66598 | 2 | 70 | 618.9832 | 761.5712 | -0.035408 | 0.01291 | -2.74 | 0.0078 | -0.06117 | -0.009646 | 1 | -0.061 | -0.01
A |  |  | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 2 | 70 | 312.8342 | 367.26 | -0.0091613 | 0.0058649 | -1.56 | 0.1229 | -0.020864 | 0.0025419 |  | -0.021 | 0.003
A |  | Fluometuron | 4-Hydroxy-tert-fluometuron | 68619 | 10 | 49 | 12.6321 | 13.1612 | 7.978e-05 | 9.8075e-05 | 0.81 | 0.4201 | -0.00011752 | 0.00027708 |  | 0 | 0
A |  |  | Demethyl fluometuron | 68591 | 1 | 70 | 2659.727 | 2594.2811 | -0.062507 | 0.01043 | -5.99 | <.0001 | -0.08332 | -0.041693 | 1 | -0.083 | -0.042
A |  |  | Fluometuron | 68608 | 1 | 63 | 1091.767 | 1088.8754 | 0.0047802 | 0.0038213 | 1.25 | 0.2157 | -0.0028609 | 0.012421 |  | -0.003 | 0.012
A |  |  | Hydroxy mono demethyl fluometuron | 68617 | 5 | 63 | 65.3282 | 66.3534 | -0.00050811 | 0.00066137 | -0.77 | 0.4453 | -0.0018306 | 0.00081438 |  | -0.002 | 0.001
A |  |  | Hydroxyfluometuron | 68620 | 4 | 56 | 225.3608 | 221.9579 | -0.0005262 | 0.0016077 | -0.33 | 0.7447 | -0.0037495 | 0.0026971 |  | -0.004 | 0.003
A |  | Halosulfuron-methyl | Halosulfuron-methyl | 61680 | 11 | 56 | 37.3295 | 37.5611 | 0.00058894 | 0.00040949 | 1.44 | 0.1561 | -0.00023202 | 0.0014099 |  | 0 | 0.001
A |  | Linuron | Linuron | 68639 | 3 | 70 | 181.97 | 180.1118 | 0.0011154 | 0.0013401 | 0.83 | 0.4081 | -0.0015587 | 0.0037895 |  | -0.002 | 0.004
A |  | Nicosulfuron | Nicosulfuron | 61685 | 7 | 56 | 40.2174 | 41.4271 | 0.00016073 | 0.000678 | 0.24 | 0.8135 | -0.0011986 | 0.00152 |  | -0.001 | 0.002
A |  | Orthosulfamuron | Orthosulfamuron | 68662 | 3 | 42 | 8.5947 | 8.5201 | 7.435e-06 | 9.6167e-05 | 0.08 | 0.9388 | -0.00018693 | 0.0002018 |  | 0 | 0
A |  | Siduron | Siduron | 68686 | 2 | 70 | 396.207 | 407.2215 | -0.010645 | 0.0029479 | -3.61 | 0.0006 | -0.016528 | -0.0047626 | 1 | -0.017 | -0.005
A |  | Sulfometuron-methyl | Sulfometuron-methyl | 68688 | 2 | 63 | 450.6747 | 454.6041 | -0.00082292 | 0.0022794 | -0.36 | 0.7193 | -0.0053809 | 0.0037351 |  | -0.005 | 0.004
A |  | Sulfosulfuron | Sulfosulfuron | 68689 | 4 | 70 | 142.9728 | 161.9626 | -0.0015606 | 0.0026895 | -0.58 | 0.5637 | -0.0069273 | 0.0038062 |  | -0.007 | 0.004
A |  |  | Sulfosulfuron ethyl sulfone | 68690 | 1 | 70 | 528.4112 | 526.383 | 0.0082901 | 0.0022821 | 3.63 | 0.0005 | 0.0037362 | 0.012844 | 1 | 0.004 | 0.013
A |  | Tebuthiuron | Hydroxytebuthurion | 68621 | 5 | 49 | 165.4292 | 163.6347 | 0.0017242 | 0.00082597 | 2.09 | 0.0423 | 6.2586e-05 | 0.0033859 |  | 0 | 0.003
A |  |  | Tebuthiuron | 68695 | 1 | 70 | 1297.6912 | 1408.575 | -0.040042 | 0.013392 | -2.99 | 0.0039 | -0.066765 | -0.013319 | 1 | -0.067 | -0.013
A |  |  | Tebuthiuron TP 104 | 68575 | 3 | 63 | 392.7889 | 395.3301 | 0.0012216 | 0.0026047 | 0.47 | 0.6407 | -0.0039868 | 0.00643 |  | -0.004 | 0.006
A |  |  | Tebuthiuron TP 106 | 68714 | 38 | 49 | 57.3594 | 56.5774 | 0.00049014 | 0.00056534 | 0.87 | 0.3904 | -0.00064717 | 0.0016275 |  | -0.001 | 0.002
A |  |  | Tebuthiuron TP 108 | 68696 | 5 | 49 | 201.5347 | 200.8549 | -0.0031836 | 0.00085002 | -3.75 | 0.0005 | -0.0048936 | -0.0014736 | 1 | -0.005 | -0.001
A |  |  | Tebuthiuron TP 109 (OH) | 68697 | 8 | 49 | 28.1354 | 28.1968 | 0.00039008 | 0.00017927 | 2.18 | 0.0346 | 2.9435e-05 | 0.00075073 |  | 0 | 0.001
A | triazine | Ametryn | Ametryn | 68533 | 1 | 70 | 877.8638 | 874.0095 | 0.010189 | 0.0044536 | 2.29 | 0.0253 | 0.0013021 | 0.019076 | 1 | 0.001 | 0.019
A |  | Atrazine | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 2 | 70 | 255.8272 | 262.4679 | 0.0018144 | 0.0023668 | 0.77 | 0.446 | -0.0029085 | 0.0065374 |  | -0.003 | 0.007
A |  |  | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 3 | 63 | 89.3565 | 91.0629 | -0.0014613 | 0.00044292 | -3.3 | 0.0016 | -0.002347 | -0.00057566 | 1 | -0.002 | -0.001
A |  |  | 2-Hydroxyatrazine | 68660 | 4 | 56 | 332.7283 | 319.8194 | 0.0011097 | 0.0036303 | 0.31 | 0.761 | -0.0061687 | 0.008388 |  | -0.006 | 0.008
A |  |  | Atrazine | 65065 | 2 | 63 | 445.4125 | 493.6025 | -0.013373 | 0.0049073 | -2.73 | 0.0084 | -0.023185 | -0.0035598 | 1 | -0.023 | -0.004
A |  |  | Deethylatrazine | 68552 | 5 | 55 | 103.8363 | 102.3469 | -0.0017461 | 0.00055217 | -3.16 | 0.0026 | -0.0028536 | -0.00063854 | 1 | -0.003 | -0.001
A |  |  | Deisopropylatrazine | 68550 | 10 | 70 | 26.4353 | 27.6697 | -0.00024719 | 0.00031454 | -0.79 | 0.4347 | -0.00087485 | 0.00038048 |  | -0.001 | 0
A |  |  | Didealkylatrazine | 68547 | 12 | 42 | 13.1069 | 14.11 | -0.00085399 | 0.0001625 | -5.26 | <.0001 | -0.0011824 | -0.00052556 | 1 | -0.001 | -0.001
A |  | Cyanazine | Cyanazine | 66592 | 8 | 49 | 119.3488 | 117.6354 | 0.0010546 | 0.00077672 | 1.36 | 0.181 | -0.00050794 | 0.0026172 |  | -0.001 | 0.003
A |  | Hexazinone | 4-Hydroxyhexazinone A | 68517 | 1 | 70 | 381.9051 | 379.9567 | 0.0027081 | 0.0023516 | 1.15 | 0.2535 | -0.0019843 | 0.0074006 |  | -0.002 | 0.007
A |  |  | Demethyl hexazinone B | 68566 | 1 | 70 | 1486.837 | 1482.2402 | -0.0064357 | 0.0074774 | -0.86 | 0.3924 | -0.021357 | 0.0084851 |  | -0.021 | 0.008
A |  |  | Hexazinone | 65085 | 1 | 70 | 1408.5666 | 1400.014 | -0.001061 | 0.0067536 | -0.16 | 0.8756 | -0.014538 | 0.012416 |  | -0.015 | 0.012
A |  |  | Hexazinone TP C | 68612 | 1 | 70 | 597.8776 | 600.6105 | 0.0066407 | 0.0024917 | 2.67 | 0.0096 | 0.0016685 | 0.011613 | 1 | 0.002 | 0.012
A |  |  | Hexazinone TP F | 68574 | 5 | 56 | 206.7725 | 205.7932 | -0.00015816 | 0.0017294 | -0.09 | 0.9275 | -0.0036253 | 0.003309 |  | -0.004 | 0.003
A |  |  | Hexazinone TP G | 68713 | 11 | 49 | 64.6043 | 63.9349 | -0.00022114 | 0.0005457 | -0.41 | 0.6871 | -0.0013189 | 0.00087666 |  | -0.001 | 0.001
A |  | Metribuzin | Metribuzin | 68652 | 10 | 63 | 71.3057 | 75.0714 | -0.001268 | 0.0012376 | -1.02 | 0.3096 | -0.0037428 | 0.0012067 |  | -0.004 | 0.001
A |  |  | Metribuzin DA | 68568 | 4 | 49 | 110.482 | 108.9277 | 0.00019674 | 0.00059722 | 0.33 | 0.7433 | -0.0010047 | 0.0013982 |  | -0.001 | 0.001
A |  | Prometon | Prometon | 67702 | 2 | 70 | 801.4259 | 812.3243 | 0.0035239 | 0.0055695 | 0.63 | 0.529 | -0.0075899 | 0.014638 |  | -0.008 | 0.015
A |  | Prometryn | Deisopropyl prometryn | 68564 | 1 | 70 | 1645.6962 | 1621.7184 | 0.010237 | 0.0068854 | 1.49 | 0.1417 | -0.0035025 | 0.023977 |  | -0.004 | 0.024
A |  |  | Prometryn | 65103 | 1 | 70 | 1177.8614 | 1167.8791 | 0.015474 | 0.0043216 | 3.58 | 0.0006 | 0.0068505 | 0.024098 | 1 | 0.007 | 0.024
A |  | Propazine | Propazine | 68678 | 2 | 70 | 473.4439 | 481.1926 | -0.0017991 | 0.0034283 | -0.52 | 0.6014 | -0.0086401 | 0.0050419 |  | -0.009 | 0.005
A |  | Simazine | Hydroxysimazine | 68624 | 1 | 70 | 304.6442 | 296.3065 | 0.0069729 | 0.0024809 | 2.81 | 0.0065 | 0.0020224 | 0.011923 | 1 | 0.002 | 0.012
A |  |  | Simazine | 65105 | 5 | 56 | 125.8395 | 124.9072 | 0.00019102 | 0.00084696 | 0.23 | 0.8224 | -0.001507 | 0.0018891 |  | -0.002 | 0.002
A |  | Terbuthylazine | Terbuthylazine | 66651 | 1 | 63 | 1086.1355 | 1072.302 | 0.012663 | 0.0042478 | 2.98 | 0.0041 | 0.004169 | 0.021157 | 1 | 0.004 | 0.021
E | acetanilide & amide | Acetochlor/Metolachlor | 2-(1-hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 47 | 35 | 89.1149 | 80.6928 | 0.0051022 | 0.0012956 | 3.94 | 0.0004 | 0.0024664 | 0.0077381 | 1 | 0.002 | 0.008
E | carbamate & thiocarbamate | Asulam | Asulam | 68536 | 14 | 42 | 7.5837 | 7.7372 | -1.0451e-05 | 5.9423e-05 | -0.18 | 0.8613 | -0.00013055 | 0.00010965 |  | 0 | 0
E |  | Carbofuran | 3-Hydroxycarbofuran | 68508 | 8 | 49 | 96.9612 | 99.1123 | -0.00085132 | 0.0006937 | -1.23 | 0.2259 | -0.0022469 | 0.00054423 |  | -0.002 | 0.001
E | fungicide | Benomyl | Carbendazim | 68548 | 1 | 70 | 1218.4761 | 1237.0047 | 0.01238 | 0.011726 | 1.06 | 0.2948 | -0.011019 | 0.03578 |  | -0.011 | 0.036
E |  | Famoxadone | Famoxadone | 67609 | 30 | 49 | 18.7219 | 18.5616 | 0.00026879 | 0.00017877 | 1.5 | 0.1394 | -9.0854e-05 | 0.00062844 |  | 0 | 0.001
E |  | Propiconazole | 1H-1,2,4-Triazole | 68498 | 11 | 42 | 3.573 | 4.3364 | -0.00060188 | 0.00010051 | -5.99 | <.0001 | -0.00080502 | -0.00039874 |  | -0.001 | 0
E | miscellaneous | Butralin | Butralin | 68545 | 2 | 70 | 210.4298 | 208.9724 | 0.0035859 | 0.0012751 | 2.81 | 0.0064 | 0.0010414 | 0.0061304 | 1 | 0.001 | 0.006
E |  | Fenbutatin oxide | Fenbutatin oxide | 68602 | 10 | 70 | 232.2762 | 353.0671 | -0.021959 | 0.025982 | -0.85 | 0.401 | -0.073806 | 0.029888 |  | -0.074 | 0.03
E |  | Flumetsulam | Flumetsulam | 61679 | 8 | 49 | 73.6669 | 75.0086 | 0.0020527 | 0.00032359 | 6.34 | <.0001 | 0.0014017 | 0.0027037 | 1 | 0.001 | 0.003
E |  | Oxyfluorfen | Oxyfluorfen | 65093 | 250 | 42 | 4.2038 | 4.2086 | -3.1472e-05 | 3.2104e-05 | -0.98 | 0.3328 | -9.6356e-05 | 3.3412e-05 |  | 0 | 0
E | organophosphate | Diazinon | Diazinon oxon | 68236 | 2 | 35 | 1.4221 | 1.639 | -6.3152e-05 | 3.2031e-05 | -1.97 | 0.0571 | -0.00012832 | 2.0165e-06 |  | 0 | 0
E |  | Naled | Naled | 68654 | 15 | 42 | 67.8443 | 68.0715 | 0.00088698 | 0.00033128 | 2.68 | 0.0107 | 0.00021744 | 0.0015565 |  | 0 | 0.002
E |  | Parathion-methyl | Parathion-methyl | 65089 | 250 | 28 | 0.9995 | 1.0286 | 3.8295e-06 | 1.7949e-05 | 0.21 | 0.8327 | -3.3066e-05 | 4.0725e-05 |  | 0 | 0
E | pyrethroid & phenylpyrazine | Bifenthrin | Bifenthrin | 65067 | 9 | 70 | 429.7769 | 514.5686 | -0.020694 | 0.012286 | -1.68 | 0.0967 | -0.04521 | 0.0038218 |  | -0.045 | 0.004
E | sulfonylurea & urea | Novaluron | Novaluron | 68655 | 4 | 70 | 90.4645 | 87.8388 | 0.00037895 | 0.0010339 | 0.37 | 0.7151 | -0.0016842 | 0.0024421 |  | -0.002 | 0.002
9031 | acetanilide & amide | Dimethenamid | Dimethenamid OA | 68581 | 42 | 35 | 35.8548 | 34.3807 | 0.0011311 | 0.00029781 | 3.8 | 0.0006 | 0.0005252 | 0.001737 | 1 | 0.001 | 0.002
9031 |  |  | Dimethenamid SA | 68582 | 39 | 42 | 33.7868 | 31.3199 | 0.0015434 | 0.00043787 | 3.52 | 0.0011 | 0.00065842 | 0.0024284 | 1 | 0.001 | 0.002
9031 |  |  | Dimethenamid SAA | 68583 | 94 | 35 | 35.2428 | 35.1006 | 0.00078489 | 0.00018227 | 4.31 | 0.0001 | 0.00041407 | 0.0011557 |  | 0 | 0.001
9031 | miscellaneous | Isoxaflutole | Diketonitrile isoxaflutole | 68578 | 31 | 28 | 3.4501 | 3.4398 | 6.3286e-06 | 2.6834e-05 | 0.24 | 0.8154 | -4.883e-05 | 6.1487e-05 |  | 0 | 0
9031 | pyrethroid & phenylpyrazine | Fipronil | Fipronil | 66604 | 2 | 42 | 6.7204 | 7.1943 | -0.00020381 | 6.5983e-05 | -3.09 | 0.0036 | -0.00033716 | -7.0449e-05 |  | 0 | 0
9031 |  |  | Fipronil amide | 68604 | 4 | 49 | 14.5653 | 15.1692 | -2.7047e-05 | 0.00023138 | -0.12 | 0.9074 | -0.00049252 | 0.00043843 |  | 0 | 0
9031 |  |  | Fipronil sulfide | 66610 | 2 | 56 | 12.7052 | 22.5203 | -0.0019327 | 0.00083174 | -2.32 | 0.0239 | -0.0036003 | -0.00026521 |  | -0.004 | 0
9031 | sulfonylurea & urea | Diflufenzopyr | Diflufenzopyr | 68577 | 44 | 42 | 13.4892 | 13.2823 | 0.00019417 | 0.00011287 | 1.72 | 0.0931 | -3.3948e-05 | 0.00042228 |  | 0 | 0
9031 |  | Sulfentrazone | Sulfentrazone | 68687 | 5 | 42 | 38.0095 | 38.5484 | 0.0011384 | 0.00016567 | 6.87 | <.0001 | 0.00080359 | 0.0014732 | 1 | 0.001 | 0.001
9031 | triazine | Metribuzin | Metribuzin DK | 68653 | 100 | 42 | 11.8061 | 12.321 | -9.1516e-05 | 0.00014855 | -0.62 | 0.5413 | -0.00039174 | 0.00020871 |  | 0 | 0
D | carbamate & thiocarbamate | Carbofuran | 3-Ketocarbofuran | 68509 | 88 | 49 | 64.296 | 64.1505 | 0.00048483 | 0.00037882 | 1.28 | 0.2069 | -0.00027724 | 0.0012469 |  | 0 | 0.001
D |  |  | 7-Hydroxycarbofuran | 68518 | 250 | 28 | 5.2313 | 5.8229 | -0.00036133 | 6.3332e-05 | -5.71 | <.0001 | -0.00049151 | -0.00023115 |  | 0 | 0
D |  | Formetanate | Formetanate | 68609 | 1 | 70 | 194.32 | 191.5069 | 0.0060301 | 0.0016273 | 3.71 | 0.0004 | 0.0027828 | 0.0092774 | 1 | 0.003 | 0.009
D | miscellaneous | Aminopyralid | Aminopyralid | 68534 | 250 | 42 | 2.2302 | 2.3624 | -0.00015364 | 3.0754e-05 | -5 | <.0001 | -0.00021579 | -9.148e-05 |  | 0 | 0
D |  | Bifenazate | Bifenazate | 68540 | 5 | 56 | 138.6279 | 185.6147 | 0.018755 | 0.0053005 | 3.54 | 0.0008 | 0.0081281 | 0.029382 | 1 | 0.008 | 0.029
D |  |  | Bifenazate diazene | 68541 | 2 | 70 | 1161.2099 | 1155.8371 | 0.00081216 | 0.0080456 | 0.1 | 0.9199 | -0.015243 | 0.016867 |  | -0.015 | 0.017
D |  | Flumiclorac | Flumiclorac | 68607 | 2 | 70 | 151.3465 | 146.1978 | 0.002992 | 0.0015017 | 1.99 | 0.0503 | -4.5401e-06 | 0.0059886 |  | 0 | 0.006
D | organophosphate | Phosmet | Phosmet oxon | 68674 | 10 | 56 | 596.7702 | 610.9871 | 0.029645 | 0.011646 | 2.55 | 0.0138 | 0.0062955 | 0.052995 | 1 | 0.006 | 0.053
Std Error Lists the estimates of the standard errors of the parameter estimates of the linear fit. They are used in constructing tests and confidence intervals.
t Ratio Lists the test statistics for the hypothesis that each parameter is zero. It is the ratio of the parameter estimate of the linear fit to its standard error. If the hypothesis is true, then this statistic has a Student’s t-distribution.
Prob >t Lists the observed significance probability calculated from each t-ratio. It is the probability of getting, by chance alone, a t-ratio greater (in absolute value) than the computed value, given a true null hypothesis. Often, a value below 0.05 (or sometimes 0.01) is interpreted as evidence that the parameter is significantly different from zero
Lower 95% is the lower endpoint of the 95% confidence interval for the parameter estimate.
Upper 95% is the upper endpoint of the 95% confidence interval for the parameter estimate
[6 empty] | >1000 | 30
[6 empty] | <10 | 17

## t16_esineg_response
Table 16. Summary of response factor and parameters of linear fit of response factor by nominal concentration for negative electrospray ionization mode analytes [Data are from ESI+ 7 method detection level (MDL) experiments with 10 concentration levels and 7 replicates at each level. The number of calibration standards included in the summary was different for each analyte because of different response factors and the corresponding lowest calibration standards that met the identification criteria. Slope of zero indicates response factors are constant over the concentration range. Slopes that are significantly different from zero are in bold and highlighted in red. Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; ESI-, negative electrospray ionization; USGS, U.S. Geological Survey; ng/L, nanograms per liter; N, number of samples; t, Student's t-distribution; Prob>, probably greater then; Lower 95 percent, lower 95th percentile confidence limit of slope estimate; Upper 95 percent, upper 95th percentile confidence limit of slope estimate; na, not available]
Validation quality code | Analytical method group | Pesticide group | Analyte | USGS Parameter code | MDL (ng/L) | N | Median response factor | Mean response factor | Slope estimate | Std Error | t Ratio | Prob>¦t¦ | Lower 95% | Upper 95%
A | acetanilide & amide | Acetochlor | Acetochlor OA | 68522 | 45 | 52 | 8.8258 | 9.2773 | 0.0001012 | 0.00010709 | 0.94493 | 0.34924 | -0.00011391 | 0.0003163
A | acetanilide & amide | Acetochlor | Acetochlor SA | 68523 | 50 | 69 | 16.9749 | 17.1018 | 0.00030672 | 0.00019783 | 1.5504 | 0.12576 | -8.8157e-05 | 0.00070159
A | acetanilide & amide | Acetochlor | Acetochlor SAA | 68524 | 88 | 58 | 2.2616 | 2.4353 | 8.0477e-06 | 3.7458e-05 | 0.21485 | 0.83067 | -6.6989e-05 | 8.3084e-05
A | acetanilide & amide | Acetochlor/Metolachlor | sec-Acetochlor-OA | 68684 | 26 | 57 | 9.9903 | 9.6787 | 0.00015015 | 9.5593e-05 | 1.5708 | 0.12198 | -4.142e-05 | 0.00034172
A | acetanilide & amide | Alachlor | Alachlor OA | 68526 | 42 | 57 | 9.6145 | 10.1684 | 0.00012423 | 0.00013063 | 0.95104 | 0.34574 | -0.00013755 | 0.00038602
E | acetanilide & amide | Alachlor | Alachlor SA | 68871 | 100 | 70 | 17.6597 | 17.6668 | 0.00020571 | 0.00019372 | 1.0619 | 0.29204 | -0.00018086 | 0.00059228
A | acetanilide & amide | Alachlor | Alachlor SAA | 68527 | 84 | 56 | 2.7563 | 2.8905 | 3.3216e-05 | 3.9222e-05 | 0.84688 | 0.4008 | -4.5419e-05 | 0.00011185
A | acetanilide & amide | Alachlor | sec-Alachlor-OA | 68685 | 68 | 49 | 7.2871 | 7.3837 | 0.00011973 | 9.367e-05 | 1.2782 | 0.20745 | -6.871e-05 | 0.00030817
A | acetanilide & amide | Dimethenamid | Dimethenamid OA | 68581 | 42 | 56 | 7.8272 | 7.7252 | 0.00013814 | 9.8978e-05 | 1.3957 | 0.16852 | -6.0295e-05 | 0.00033658
A | acetanilide & amide | Dimethenamid | Dimethenamid SA | 68582 | 39 | 58 | 2.2471 | 2.4883 | 6.4868e-06 | 3.6841e-05 | 0.17608 | 0.86087 | -6.7314e-05 | 8.0287e-05
A | acetanilide & amide | Dimethenamid | Dimethenamid SAA | 68583 | 94 | 50 | 1.3242 | 1.4702 | 1.3388e-05 | 2.7575e-05 | 0.48553 | 0.62951 | -4.2054e-05 | 6.8831e-05
A | acetanilide & amide | Metolachlor | Metolachlor OA | 68650 | 74 | 52 | 10.9215 | 11.0272 | 0.00018522 | 0.00012329 | 1.5023 | 0.1393 | -6.2409e-05 | 0.00043285
A | acetanilide & amide | Metolachlor | Metolachlor SA | 68651 | 34 | 62 | 3.4506 | 3.4706 | -8.112e-06 | 4.728e-05 | -0.17157 | 0.86435 | -0.00010269 | 8.6462e-05
A | acid | 2,4-D | 2,4-D | 68500 | 31 | 59 | 8.639 | 8.7629 | -6.5279e-06 | 0.00011365 | -0.057437 | 0.9544 | -0.00023411 | 0.00022106
A | acid | Bentazone | Bentazon | 68538 | 5 | 70 | 50.0877 | 52.798 | -0.0024534 | 0.00057122 | -4.295 | 5.6859e-05 | -0.0035933 | -0.0013135
A | acid | Bromoxynil | Bromoxynil | 68543 | 39 | 61 | 5.5365 | 6.1902 | -0.00022483 | 8.1067e-05 | -2.7733 | 0.0074158 | -0.00038704 | -6.2611e-05
A | acid | MCPA | MCPA | 68641 | 48 | 51 | 9.264 | 9.3728 | 1.6359e-05 | 9.9251e-05 | 0.16483 | 0.86976 | -0.00018309 | 0.00021581
A | acid | Triclopyr | Triclopyr | 68712 | 44 | 56 | 4.4935 | 4.5761 | 1.9814e-05 | 6.3061e-05 | 0.3142 | 0.75458 | -0.00010662 | 0.00014624
A | carbamate & thiocarbamate | Triallate | 2,3,3-trichloroprop-2-ene-1-SA | 68691 | 27 | 59 | 3.0804 | 3.1192 | 1.3146e-05 | 2.043e-05 | 0.64345 | 0.52251 | -2.7765e-05 | 5.4056e-05
A | fungicide | Chlorothalonil | 4-Hydroxychlorothalonil | 68336 | 49 | 55 | 1.5053 | 1.5143 | 1.5613e-05 | 1.5426e-05 | 1.0121 | 0.31609 | -1.5328e-05 | 4.6553e-05
A | miscellaneous | Flubendiamide | Flubendiamide | 68606 | 2 | 78 | 103.1533 | 105.7769 | -0.0017728 | 0.00074549 | -2.3781 | 0.019918 | -0.0032576 | -0.00028804
A | miscellaneous | Isoxaflutole | Diketonitrile isoxaflutole | 68578 | 31 | 74 | 38.1084 | 45.6631 | -0.0021375 | 0.00072046 | -2.9669 | 0.0040807 | -0.0035737 | -0.0007013
A | miscellaneous | Isoxaflutole | Isoxaflutole acid RPA 203328 | 68633 | 4 | 68 | 9.9191 | 10.1423 | 0.00012026 | 9.5555e-05 | 1.2585 | 0.21264 | -7.0523e-05 | 0.00031104
A | miscellaneous | Oryzalin | Oryzalin | 68663 | 5 | 67 | 10.4468 | 11.7809 | -0.00055964 | 0.00014469 | -3.8678 | 0.00025668 | -0.0008486 | -0.00027067
A | miscellaneous | Pendimethalin | 4-(Hydroxymethyl)pendimethalin | 68511 | 106 | 49 | 0.69156 | 0.6844 | -8.717e-06 | 9.8203e-06 | -0.88765 | 0.37925 | -2.8473e-05 | 1.1039e-05
A | pyrethroid & phenylpyrazine | Bifenthrin; lambda-cyhalothrin; Tefluthrin | cis-Cyhalothric acid | 68553 | 250 | 33 | 0.45882 | 0.45612 | 7.4974e-06 | 4.8908e-06 | 1.533 | 0.13543 | -2.4774e-06 | 1.7472e-05
A | pyrethroid & phenylpyrazine | Fipronil | Dechlorofipronil | 68561 | 2 | 80 | 113.5082 | 111.8413 | 0.00031895 | 0.0007592 | 0.42011 | 0.67556 | -0.0011925 | 0.0018304
A | pyrethroid & phenylpyrazine | Fipronil | Desulfinylfipronil | 66607 | 2 | 77 | 229.0135 | 227.5587 | 0.0012926 | 0.0014101 | 0.91668 | 0.36225 | -0.0015164 | 0.0041015
A | pyrethroid & phenylpyrazine | Fipronil | Desulfinylfipronil amide | 68570 | 4 | 76 | 33.1232 | 34.845 | -0.00082369 | 0.00026741 | -3.0803 | 0.0029025 | -0.0013565 | -0.00029087
A | pyrethroid & phenylpyrazine | Fipronil | Fipronil | 66604 | 2 | 78 | 128.6606 | 129.0449 | -0.00056582 | 0.00074011 | -0.76451 | 0.44693 | -0.0020399 | 0.00090824
A | pyrethroid & phenylpyrazine | Fipronil | Fipronil amide | 68604 | 4 | 79 | 17.2615 | 18.0315 | -0.00037446 | 0.00017813 | -2.1023 | 0.0388 | -0.00072916 | -1.9772e-05
A | pyrethroid & phenylpyrazine | Fipronil | Fipronil sulfide | 66610 | 2 | 80 | 107.9369 | 109.8315 | -0.00084916 | 0.0006896 | -1.2314 | 0.22188 | -0.002222 | 0.00052373
A | pyrethroid & phenylpyrazine | Fipronil | Fipronil sulfonate | 68605 | 48 | 59 | 9.2303 | 9.0807 | 0.00014275 | 7.7413e-05 | 1.844 | 0.07038 | -1.2266e-05 | 0.00029777
A | pyrethroid & phenylpyrazine | Fipronil | Fipronil sulfone | 66613 | 2 | 80 | 157.6688 | 159.2148 | 1.2377e-05 | 0.000867 | 0.014275 | 0.98865 | -0.0017137 | 0.0017384
A | pyrethroid & phenylpyrazine | Permethrin | 3-Phenoxybenzoic acid | 68873 | 30 | 59 | 7.1221 | 8.1152 | -0.00023686 | 0.00015869 | -1.4926 | 0.14106 | -0.00055464 | 8.0913e-05
A | sulfonylurea & urea | Diflufenzopyr | Diflufenzopyr | 68577 | 44 | 61 | 7.7027 | 7.7892 | -8.8498e-05 | 9.5785e-05 | -0.92392 | 0.35929 | -0.00028016 | 0.00010317
A | sulfonylurea & urea | Prosulfuron | Prosulfuron | 61687 | 5 | 78 | 44.6386 | 43.1051 | -0.0014527 | 0.0005272 | -2.7554 | 0.0073324 | -0.0025027 | -0.00040266
A | sulfonylurea & urea | Sulfentrazone | Sulfentrazone | 68687 | 5 | 68 | 47.1622 | 46.3149 | -0.0029372 | 0.00057131 | -5.1413 | 2.6325e-06 | -0.0040779 | -0.0017966
A | triazine | Hexazinone | Hexazinone TP E | 68614 | 38 | 57 | 0.72754 | 0.74876 | -2.3082e-07 | 6.7348e-06 | -0.034272 | 0.97278 | -1.3728e-05 | 1.3266e-05
A | triazine | Metribuzin | Metribuzin DADK | 68569 | 100 | 52 | 0.78739 | 0.87524 | -2.9138e-05 | 1.4564e-05 | -2.0006 | 0.050876 | -5.8392e-05 | 1.1538e-07
A | triazine | Metribuzin | Metribuzin DK | 68653 | 100 | 47 | 0.92038 | 1.0053 | -2.8737e-05 | 2.0148e-05 | -1.4263 | 0.16069 | -6.9316e-05 | 1.1843e-05
E | acid | Dacthal | Chlorthal-monomethyl | 68560 | 250 | 45 | 0.49343 | 0.529 | -1.036e-05 | 8.4857e-06 | -1.2209 | 0.22878 | -2.7473e-05 | 6.7529e-06
E | acid | Dicamba | Dicamba | 68571 | 250 | 41 | 1.5422 | 1.5385 | -7.5905e-06 | 2.0038e-05 | -0.3788 | 0.70689 | -4.8121e-05 | 3.294e-05
E | sulfonylurea & urea | Halosulfuron-methyl | Chlorosulfonamide acid | 68551 | 37 | 57 | 5.2064 | 5.4482 | -1.11e-05 | 6.2535e-05 | -0.1775 | 0.85976 | -0.00013642 | 0.00011422
E | triazine | Hexazinone | Hexazinone TP D | 68613 | 100 | 47 | 1.3087 | 1.4372 | -4.6285e-05 | 2.3436e-05 | -1.9749 | 0.05443 | -9.3488e-05 | 9.1791e-07
9030 | acetanilide & amide | Propanil | Propanil | L903F | 29 | 67 | 21.6528 | 22.4063 | -0.00064215 | 0.00017323 | -3.707 | 0.00043583 | -0.0009881 | -0.00029619
9030 | sulfonylurea & urea | Chlorimuron-ethyl | Chlorimuron-ethyl | L903C | 6 | 67 | 12.3075 | 12.3759 | -0.00032669 | 0.00013325 | -2.4517 | 0.016912 | -0.0005928 | -6.0576e-05
9030 | sulfonylurea & urea | Chlorsulfuron | Chlorsulfuron | L903R | 48 | 60 | 9.6312 | 10.0423 | -0.00039605 | 0.00016198 | -2.4451 | 0.017539 | -0.00072028 | -7.1818e-05
9030 | sulfonylurea & urea | Novaluron | Novaluron | L903J | 5 | 78 | 29.4055 | 28.6203 | -0.001396 | 0.00025027 | -5.578 | 3.5689e-07 | -0.0018945 | -0.00089758
D | fungicide | Iprodione | Iprodione | 66617 | 1000 | 30 | 0.041162 | 0.053369 | -4.4568e-06 | 1.7169e-06 | -2.5958 | 0.01486 | -7.9737e-06 | -9.3989e-07
D | triazine | Atrazine | Ammelide | 68535 | 37 | 58 | 0.62409 | 0.73716 | -6.6802e-05 | 1.349e-05 | -4.9521 | 7.1088e-06 | -9.3826e-05 | -3.9779e-05
[6 empty] | >100 | 6
[6 empty] | <1 | 8

## T17_lims_2437_mdl_mrl
Table 17. Summary of method detection levels, laboratory reporting levels, low and high calibration standards for pesticide compounds determined by direct aqueous-injection liquid chromatography-tandem mass spectrometry [Table is sorted by lab code and Validation quality code. ESI+, positive electrospray ionization mode; ESI-, negative electrospray ionization mode; Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; USGS, U.S. Geological Survey; MDL, method detection level; ng/L, nanograms per liter; LRL, laboratory reporting level; std., standard]
Acquisition mode | Validation quality code | Lab code | Pesticide compound | USGS Parameter code | MDL (ng/L) | LRL (ng/L) | Low std (ng/L) | High std (ng/L)
ESI+ | A | 9030 | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 2.448 | 5 | 2.5 | 10000
ESI+ | A | 9030 | 2-Amino-N-isopropylbenzamide | 68503 | 1.872 | 4 | 1 | 10000
ESI+ | A | 9030 | 2-Aminobenzimidazole | 68502 | 4.08 | 9 | 5 | 10000
ESI+ | A | 9030 | 2-Chloro-2,6-diethylacetanilide | 68525 | 2.256 | 5 | 2.5 | 10000
ESI+ | A | 9030 | 2-Chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 2.232 | 5 | 2.5 | 10000
ESI+ | A | 9030 | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 1.8 | 4 | 1 | 10000
ESI+ | A | 9030 | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 2.808 | 6 | 2.5 | 10000
ESI+ | A | 9030 | 2-Hydroxyatrazine | 68660 | 4 | 8 | 2.5 | 10000
ESI+ | A | 9030 | 3,4-Dichlorophenylurea | 68226 | 72 | 144 | 10 | 10000
ESI+ | A | 9030 | 4-Chlorobenzylmethyl sulfoxide | 68514 | 2 | 4 | 1 | 10000
ESI+ | A | 9030 | 4-Hydroxy molinate | 68515 | 3.144 | 7 | 2.5 | 10000
ESI+ | A | 9030 | 4-Hydroxy-tert-fluometuron | 68619 | 9.936 | 20 | 10 | 10000
ESI+ | A | 9030 | 4-Hydroxyhexazinone A | 68517 | 1.296 | 3 | 1 | 10000
ESI+ | A | 9030 | Acephate | 68519 | 5 | 10 | 1 | 10000
ESI+ | A | 9030 | Acetochlor | 68520 | 4.728 | 10 | 5 | 10000
ESI+ | A | 9030 | Alachlor | 65064 | 4.368 | 9 | 5 | 10000
ESI+ | A | 9030 | Aldicarb | 68528 | 3.936 | 8 | 1 | 10000
ESI+ | A | 9030 | Aldicarb sulfone | 68529 | 10 | 20 | 5 | 10000
ESI+ | A | 9030 | Aldicarb sulfoxide | 68530 | 1.128 | 3 | 1 | 10000
ESI+ | A | 9030 | Ametryn | 68533 | 1.32 | 3 | 1 | 10000
ESI+ | A | 9030 | Atrazine | 65065 | 2.472 | 5 | 1 | 10000
ESI+ | A | 9030 | Azinphos-methyl | 65066 | 3.624 | 8 | 1 | 10000
ESI+ | A | 9030 | Azinphos-methyl oxon | 68211 | 7.104 | 15 | 1 | 10000
ESI+ | A | 9030 | Azoxystrobin | 66589 | 1.392 | 3 | 1 | 10000
ESI+ | A | 9030 | Bromacil | 68542 | 2.904 | 6 | 1 | 10000
ESI+ | A | 9030 | Butylate | 65068 | 2.736 | 6 | 5 | 10000
ESI+ | A | 9030 | Carbaryl | 65069 | 3 | 6 | 2.5 | 10000
ESI+ | A | 9030 | Carbofuran | 65070 | 1.632 | 4 | 1 | 10000
ESI+ | A | 9030 | Carboxy molinate | 68549 | 11 | 22 | 2.5 | 10000
ESI+ | A | 9030 | Chlorimuron-ethyl | 68872 | 4 | 8 | 2.5 | 10000
ESI+ | A | 9030 | Chlorpyrifos | 65072 | 1.512 | 4 | 1 | 10000
ESI+ | A | 9030 | Chlorpyrifos oxon | 68216 | 1.512 | 4 | 1 | 10000
ESI+ | A | 9030 | Chlorsulfuron | 61678 | 25 | 50 | 10 | 10000
ESI+ | A | 9030 | cis-Permethrin | 68769 | 1 | 2 | 1 | 10000
ESI+ | A | 9030 | Cyanazine | 66592 | 7.728 | 16 | 5 | 10000
ESI+ | A | 9030 | Dechlorometolachlor | 68562 | 0.552 | 2 | 1 | 10000
ESI+ | A | 9030 | Deethylatrazine | 68552 | 5.208 | 11 | 1 | 10000
ESI+ | A | 9030 | Desiodo flubendiamide | 68563 | 1.848 | 4 | 5 | 10000
ESI+ | A | 9030 | Deisopropyl prometryn | 68564 | 1.248 | 3 | 1 | 10000
ESI+ | A | 9030 | Deisopropylatrazine | 68550 | 10 | 20 | 5 | 10000
ESI+ | A | 9030 | Demethyl fluometuron | 68591 | 0.816 | 2 | 1 | 10000
ESI+ | A | 9030 | Demethyl hexazinone B | 68566 | 1.224 | 3 | 1 | 10000
ESI+ | A | 9030 | Demethyl norflurazon | 68567 | 1.944 | 4 | 2.5 | 10000
ESI+ | A | 9030 | Diazinon | 65078 | 1.2 | 3 | 1 | 10000
ESI+ | A | 9030 | Dichlorvos | 68572 | 26 | 52 | 10 | 10000
ESI+ | A | 9030 | Dicrotophos | 68573 | 1.896 | 4 | 1 | 10000
ESI+ | A | 9030 | Didealkylatrazine | 68547 | 12 | 24 | 5 | 10000
ESI+ | A | 9030 | Diflubenzuron | 68576 | 2.568 | 6 | 1 | 10000
ESI+ | A | 9030 | Dimethenamid | 68580 | 1.224 | 3 | 1 | 10000
ESI+ | A | 9030 | Dimethoate | 66596 | 1.44 | 3 | 1 | 10000
ESI+ | A | 9030 | Dimethoate oxon | 68661 | 0.864 | 2 | 1 | 10000
ESI+ | A | 9030 | Disulfoton | 67595 | 6.24 | 13 | 2.5 | 10000
ESI+ | A | 9030 | Disulfoton oxon | 68586 | 1.224 | 3 | 1 | 10000
ESI+ | A | 9030 | Disulfoton oxon sulfone | 68588 | 3.312 | 7 | 1 | 10000
ESI+ | A | 9030 | Disulfoton oxon sulfoxide | 68587 | 3.144 | 7 | 1 | 10000
ESI+ | A | 9030 | Disulfoton sulfone | 68589 | 4.272 | 9 | 5 | 10000
ESI+ | A | 9030 | Disulfoton sulfoxide | 68590 | 1.512 | 4 | 1 | 10000
ESI+ | A | 9030 | Diuron | 66598 | 2.304 | 5 | 1 | 10000
ESI+ | A | 9030 | EPTC | 65080 | 103 | 206 | 10 | 10000
ESI+ | A | 9030 | EPTC R248722 | 68594 | 1.704 | 4 | 2.5 | 10000
ESI+ | A | 9030 | Ethoprop | 68596 | 2.064 | 5 | 1 | 10000
ESI+ | A | 9030 | Etoxazole | 68598 | 0.768 | 2 | 1 | 10000
ESI+ | A | 9030 | Famoxadone | 67609 | 30 | 60 | 10 | 10000
ESI+ | A | 9030 | Fenamiphos | 68599 | 0.984 | 2 | 1 | 10000
ESI+ | A | 9030 | Fenamiphos sulfone | 68600 | 2.064 | 5 | 1 | 10000
ESI+ | A | 9030 | Fenamiphos sulfoxide | 68601 | 2.472 | 5 | 1 | 10000
ESI+ | A | 9030 | Fentin | 68603 | 7.848 | 16 | 5 | 10000
ESI+ | A | 9030 | Fluometuron | 68608 | 1.368 | 3 | 1 | 10000
ESI+ | A | 9030 | Fonofos | 65084 | 5.208 | 11 | 2.5 | 10000
ESI+ | A | 9030 | Halosulfuron-methyl | 61680 | 10.704 | 22 | 2.5 | 10000
ESI+ | A | 9030 | Hexazinone | 65085 | 1.32 | 3 | 1 | 10000
ESI+ | A | 9030 | Hexazinone TP C | 68612 | 0.912 | 2 | 1 | 10000
ESI+ | A | 9030 | Hexazinone TP F | 68574 | 4.92 | 10 | 5 | 10000
ESI+ | A | 9030 | Hexazinone TP G | 68713 | 10.776 | 22 | 5 | 10000
ESI+ | A | 9030 | Hydroxy mono demethyl fluometuron | 68617 | 4.968 | 10 | 2.5 | 10000
ESI+ | A | 9030 | Hydroxyacetochlor | 68615 | 4.344 | 9 | 2.5 | 10000
ESI+ | A | 9030 | Hydroxyalachlor | 68616 | 3.768 | 8 | 2.5 | 10000
ESI+ | A | 9030 | Hydroxydiazinon | 68618 | 5.472 | 11 | 1 | 10000
ESI+ | A | 9030 | Hydroxyfluometuron | 68620 | 3.672 | 8 | 1 | 10000
ESI+ | A | 9030 | Hydroxymetolachlor | 68622 | 0.888 | 2 | 1 | 10000
ESI+ | A | 9030 | Hydroxyphthalazinone | 68623 | 23 | 46 | 10 | 10000
ESI+ | A | 9030 | Hydroxysimazine | 68624 | 1.248 | 3 | 1 | 10000
ESI+ | A | 9030 | Hydroxytebuthurion | 68621 | 5.352 | 11 | 1 | 10000
ESI+ | A | 9030 | Imazamox | 68625 | 14 | 28 | 2.5 | 10000
ESI+ | A | 9030 | Imazaquin | 61682 | 9.072 | 19 | 5 | 10000
ESI+ | A | 9030 | Imazethapyr | 61683 | 9.528 | 20 | 2.5 | 10000
ESI+ | A | 9030 | Imidacloprid | 68426 | 5.4 | 11 | 2.5 | 10000
ESI+ | A | 9030 | Indoxacarb | 68627 | 3.816 | 8 | 2.5 | 10000
ESI+ | A | 9030 | Isoxaflutole | 68632 | 6.36 | 13 | 10 | 10000
ESI+ | A | 9030 | Kresoxim-methyl | 67670 | 2.28 | 5 | 1 | 10000
ESI+ | A | 9030 | Lactofen | 68638 | 1.44 | 3 | 1 | 10000
ESI+ | A | 9030 | Linuron | 68639 | 2.568 | 6 | 1 | 10000
ESI+ | A | 9030 | Malaoxon | 68240 | 1.224 | 3 | 1 | 10000
ESI+ | A | 9030 | Malathion | 65087 | 1.896 | 6 | 5 | 10000
ESI+ | A | 9030 | Metalaxyl | 68437 | 2.76 | 6 | 1 | 10000
ESI+ | A | 9030 | Metconazole | 66620 | 2.16 | 5 | 1 | 10000
ESI+ | A | 9030 | Methamidophos | 68644 | 1.272 | 3 | 1 | 10000
ESI+ | A | 9030 | Methidathion | 65088 | 5 | 10 | 1 | 10000
ESI+ | A | 9030 | Methomyl | 68645 | 1.08 | 3 | 1 | 10000
ESI+ | A | 9030 | Methomyl oxime | 68646 | 39 | 78 | 25 | 10000
ESI+ | A | 9030 | Methoxyfenozide | 68647 | 0.96 | 2 | 1 | 10000
ESI+ | A | 9030 | Metolachlor | 65090 | 4.2 | 9 | 1 | 10000
ESI+ | A | 9030 | Metolachlor hydroxy morpholinone | 68649 | 3.048 | 7 | 1 | 10000
ESI+ | A | 9030 | Metribuzin | 68652 | 9.744 | 20 | 10 | 10000
ESI+ | A | 9030 | Metribuzin DA | 68568 | 4.392 | 9 | 2.5 | 10000
ESI+ | A | 9030 | Molinate | 65091 | 8.544 | 18 | 5 | 10000
ESI+ | A | 9030 | Myclobutanil | 66632 | 3.384 | 7 | 1 | 10000
ESI+ | A | 9030 | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 2.496 | 5 | 5 | 10000
ESI+ | A | 9030 | Nicosulfuron | 61685 | 7 | 14 | 5 | 10000
ESI+ | A | 9030 | Norflurazon | 67685 | 1.512 | 4 | 1 | 10000
ESI+ | A | 9030 | O-Ethyl-O-methyl-S-propyl phosphorothioate | 68597 | 2.28 | 5 | 1 | 10000
ESI+ | A | 9030 | O-Ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 2.04 | 5 | 1 | 10000
ESI+ | A | 9030 | O-Ethyl-S-propyl phosphorothioate | 68658 | 32 | 64 | 25 | 10000
ESI+ | A | 9030 | Orthosulfamuron | 68662 | 3 | 6 | 1 | 10000
ESI+ | A | 9030 | Oxamyl | 68664 | 0.816 | 2 | 1 | 10000
ESI+ | A | 9030 | Oxamyl oxime | 68665 | 0.696 | 2 | 1 | 10000
ESI+ | A | 9030 | Paraoxon | 68666 | 1.536 | 4 | 1 | 10000
ESI+ | A | 9030 | Paraoxon-methyl | 68648 | 9.312 | 19 | 10 | 10000
ESI+ | A | 9030 | Pendimethalin | 65098 | 1.416 | 3 | 1 | 10000
ESI+ | A | 9030 | Phorate | 68668 | 3 | 6 | 5 | 10000
ESI+ | A | 9030 | Phorate oxon | 68669 | 2 | 4 | 1 | 10000
ESI+ | A | 9030 | Phorate oxon sulfone | 68670 | 9.72 | 20 | 2.5 | 10000
ESI+ | A | 9030 | Phorate oxon sulfoxide | 68671 | 3.432 | 7 | 1 | 10000
ESI+ | A | 9030 | Phorate sulfone | 68672 | 4.44 | 9 | 5 | 10000
ESI+ | A | 9030 | Phorate sulfoxide | 68673 | 2.424 | 5 | 1 | 10000
ESI+ | A | 9030 | Phthalazinone | 68675 | 2.712 | 6 | 2.5 | 10000
ESI+ | A | 9030 | Piperonyl butoxide | 65102 | 1.752 | 4 | 1 | 10000
ESI+ | A | 9030 | Profenofos | 68676 | 1.248 | 3 | 1 | 10000
ESI+ | A | 9030 | Prometon | 67702 | 1.56 | 4 | 1 | 10000
ESI+ | A | 9030 | Prometryn | 65103 | 0.96 | 2 | 1 | 10000
ESI+ | A | 9030 | Pronamide | 67706 | 1.92 | 4 | 1 | 10000
ESI+ | A | 9030 | Propanil | 66641 | 5.352 | 11 | 5 | 10000
ESI+ | A | 9030 | Propargite | 68677 | 0.792 | 2 | 1 | 10000
ESI+ | A | 9030 | Propazine | 68678 | 1.632 | 4 | 1 | 10000
ESI+ | A | 9030 | Propiconazole | 66643 | 2.856 | 6 | 1 | 10000
ESI+ | A | 9030 | Propoxur | 68679 | 1.656 | 4 | 2.5 | 10000
ESI+ | A | 9030 | Prosulfuron | 61687 | 5 | 10 | 5 | 10000
ESI+ | A | 9030 | Pymetrozine | 68680 | 0.936 | 2 | 1 | 10000
ESI+ | A | 9030 | Pyraclostrobin | 66646 | 1.128 | 3 | 1 | 10000
ESI+ | A | 9030 | Pyridaben | 68682 | 1.08 | 3 | 1 | 10000
ESI+ | A | 9030 | Pyrimidinol | 68505 | 9.816 | 20 | 2.5 | 10000
ESI+ | A | 9030 | Pyriproxyfen | 68683 | 1.248 | 3 | 1 | 10000
ESI+ | A | 9030 | Siduron | 68686 | 2.28 | 5 | 1 | 10000
ESI+ | A | 9030 | Simazine | 65105 | 4.632 | 10 | 5 | 10000
ESI+ | A | 9030 | Sulfometuron-methyl | 68688 | 1.824 | 4 | 1 | 10000
ESI+ | A | 9030 | Sulfosulfuron | 68689 | 4.392 | 9 | 2.5 | 10000
ESI+ | A | 9030 | Sulfosulfuron ethyl sulfone | 68690 | 1.08 | 3 | 1 | 10000
ESI+ | A | 9030 | tebuconazole | 66649 | 2.472 | 5 | 1 | 10000
ESI+ | A | 9030 | Tebufenozide | 68692 | 0.792 | 2 | 1 | 10000
ESI+ | A | 9030 | Tebupirimfos | 68693 | 0.84 | 2 | 1 | 10000
ESI+ | A | 9030 | Tebupirimfos oxon | 68694 | 0.768 | 2 | 1 | 10000
ESI+ | A | 9030 | Tebuthiuron | 68695 | 1.44 | 3 | 1 | 10000
ESI+ | A | 9030 | Tebuthiuron TP 104 | 68575 | 2.736 | 6 | 1 | 10000
ESI+ | A | 9030 | Tebuthiuron TP 106 | 68714 | 38 | 76 | 5 | 10000
ESI+ | A | 9030 | Tebuthiuron TP 108 | 68696 | 4.536 | 10 | 1 | 10000
ESI+ | A | 9030 | Tebuthiuron TP 109 (OH) | 68697 | 8.328 | 17 | 10 | 10000
ESI+ | A | 9030 | Terbacil | 68698 | 10.44 | 21 | 10 | 10000
ESI+ | A | 9030 | Terbufos | 68699 | 1.632 | 4 | 1 | 10000
ESI+ | A | 9030 | Terbufos oxon | 68700 | 1.512 | 4 | 1 | 10000
ESI+ | A | 9030 | Terbufos oxon sulfone | 68701 | 5.016 | 11 | 2.5 | 10000
ESI+ | A | 9030 | Terbufos oxon sulfoxide | 68702 | 1.872 | 4 | 1 | 10000
ESI+ | A | 9030 | Terbufos sulfone | 68703 | 5.352 | 11 | 2.5 | 10000
ESI+ | A | 9030 | Terbufos sulfoxide | 68704 | 1.032 | 3 | 1 | 10000
ESI+ | A | 9030 | Terbuthylazine | 66651 | 1.032 | 3 | 1 | 10000
ESI+ | A | 9030 | Tetraconazole | 66654 | 2.928 | 6 | 1 | 10000
ESI+ | A | 9030 | Thiobencarb | 65107 | 0.768 | 2 | 1 | 10000
ESI+ | A | 9030 | trans-Permethrin | 68708 | 2 | 4 | 1 | 10000
ESI+ | A | 9030 | Triallate | 68710 | 7.512 | 16 | 1 | 10000
ESI+ | A | 9030 | Tribufos | 68711 | 0.816 | 2 | 1 | 10000
ESI+ | A | 9030 | Trifloxystrobin | 66660 | 0.84 | 2 | 1 | 10000
ESI+ | E | 9030 | Asulam | 68536 | 14 | 28 | 10 | 10000
ESI+ | E | 9030 | Carbendazim | 68548 | 1.248 | 3 | 1 | 10000
ESI+ | E | 9030 | Diazinon oxon | 68236 | 2 | 4 | 1 | 10000
ESI+ | E | 9030 | Fenbutatin oxide | 68602 | 10 | 20 | 1 | 10000
ESI+ | E | 9030 | Flumetsulam | 61679 | 8.112 | 17 | 5 | 10000
ESI+ | E | 9030 | Naled | 68654 | 15 | 30 | 5 | 10000
ESI+ | E | 9030 | Novaluron | 68655 | 4.032 | 9 | 5 | 10000
ESI+ | E | 9030 | Oxyfluorfen | 65093 | 250 | 500 | 50 | 10000
ESI+ | E | 9030 | 1H-1,2,4-Triazole | 68498 | 11 | 22 | 10 | 1000
ESI+ | E | 9030 | 2-(1-hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 47 | 94 | 50 | 10000
ESI+ | E | 9030 | 3-Hydroxycarbofuran | 68508 | 7.848 | 16 | 5 | 10000
ESI+ | E | 9030 | Bifenthrin | 65067 | 9.072 | 19 | 1 | 10000
ESI+ | E | 9030 | Butralin | 68545 | 2.088 | 5 | 1 | 10000
ESI+ | E | 9030 | Parathion-methyl | 65089 | 250 | 500 | 250 | 10000
ESI+ | E | 9030 | Phosmet | 65101 | 10 | 20 | 5 | 10000
ESI+ | D | 9030 | 7-Hydroxycarbofuran | 68518 | 250 | 500 | 50 | 10000
ESI- | A | 9031 | 2,3,3-trichloroprop-2-ene-1-SA | 68691 | 26.592 | 54 | 5 | 10000
ESI- | A | 9031 | 2,4-D | 68500 | 30.768 | 62 | 50 | 10000
ESI- | A | 9031 | 3-Phenoxybenzoic acid | 68873 | 30.096 | 61 | 10 | 10000
ESI- | A | 9031 | 4-(Hydroxymethyl)pendimethalin | 68511 | 106.296 | 213 | 100 | 10000
ESI- | A | 9031 | 4-Hydroxychlorothalonil | 68336 | 48.816 | 98 | 25 | 10000
ESI- | A | 9031 | Acetochlor OA | 68522 | 44.976 | 90 | 10 | 10000
ESI- | A | 9031 | Acetochlor SA | 68523 | 35.712 | 72 | 5 | 10000
ESI- | A | 9031 | Acetochlor SAA | 68524 | 87.744 | 176 | 50 | 10000
ESI- | A | 9031 | Alachlor OA | 68526 | 41.712 | 84 | 10 | 10000
ESI- | E | 9031 | Alachlor SA | 68871 | 35.496 | 71 | 10 | 10000
ESI- | A | 9031 | Alachlor SAA | 68527 | 84.408 | 169 | 50 | 10000
ESI- | A | 9031 | Bentazone | 68538 | 5 | 10 | 5 | 10000
ESI- | A | 9031 | Bromoxynil | 68543 | 39.048 | 79 | 10 | 10000
ESI- | A | 9031 | cis-Cyhalothric acid | 68553 | 43 | 88 | 50 | 10000
ESI- | A | 9031 | Dechlorofipronil | 68561 | 1.752 | 4 | 1 | 10000
ESI- | A | 9031 | Desulfinylfipronil | 66607 | 1.848 | 4 | 1 | 10000
ESI- | A | 9031 | Desulfinylfipronil amide | 68570 | 4.152 | 9 | 5 | 10000
ESI- | A | 9031 | Diflufenzopyr | 68577 | 43.968 | 88 | 25 | 10000
ESI- | A | 9031 | Diketonitrile isoxaflutole | 68578 | 30.72 | 62 | 10 | 10000
ESI- | A | 9031 | Dimethenamid OA | 68581 | 42.144 | 85 | 25 | 10000
ESI- | A | 9031 | Dimethenamid SA | 68582 | 39.432 | 79 | 25 | 10000
ESI- | A | 9031 | Dimethenamid SAA | 68583 | 94.464 | 189 | 50 | 10000
ESI- | A | 9031 | Fipronil | 66604 | 1.992 | 4 | 1 | 10000
ESI- | A | 9031 | Fipronil amide | 68604 | 4.008 | 9 | 5 | 10000
ESI- | A | 9031 | Fipronil sulfide | 66610 | 1.992 | 4 | 1 | 10000
ESI- | A | 9031 | Fipronil sulfonate | 68605 | 48.24 | 97 | 25 | 10000
ESI- | A | 9031 | Fipronil sulfone | 66613 | 1.848 | 4 | 2.5 | 10000
ESI- | A | 9031 | Flubendiamide | 68606 | 2.232 | 5 | 1 | 10000
ESI- | A | 9031 | Hexazinone TP E | 68614 | 37.512 | 76 | 10 | 10000
ESI- | A | 9031 | Isoxaflutole acid RPA 203328 | 68633 | 4.488 | 9 | 2.5 | 10000
ESI- | A | 9031 | MCPA | 68641 | 47.664 | 96 | 25 | 10000
ESI- | A | 9031 | Metolachlor OA | 68650 | 74.304 | 149 | 25 | 10000
ESI- | A | 9031 | Metolachlor SA | 68651 | 33.672 | 68 | 10 | 10000
ESI- | A | 9031 | Metribuzin DADK | 68569 | 100 | 200 | 100 | 10000
ESI- | A | 9031 | Metribuzin DK | 68653 | 100 | 200 | 100 | 10000
ESI- | A | 9031 | Oryzalin | 68663 | 5.184 | 11 | 5 | 10000
ESI- | A | 9031 | sec-Acetochlor OA | 68684 | 25.56 | 52 | 5 | 10000
ESI- | A | 9031 | sec-Alachlor OA | 68685 | 67.632 | 136 | 50 | 10000
ESI- | A | 9031 | Sulfentrazone | 68687 | 5.448 | 11 | 10 | 10000
ESI- | A | 9031 | Triclopyr | 68712 | 43.512 | 88 | 10 | 10000
ESI- | E | 9031 | Chlorosulfonamide acid | 68551 | 37.176 | 75 | 10 | 10000
ESI- | E | 9031 | Hexazinone TP D | 68613 | 100 | 200 | 50 | 10000
ESI- | E | 9031 | Chlorthal-monomethyl | 68560 | 250 | 500 | 400 | 40000
ESI- | E | 9031 | Dicamba | 68571 | 250 | 500 | 400 | 40000
ESI- | 9030 | 9031 | Chlorsulfuron | 61678 | 48.336 | 97 | 10 | 10000
ESI- | 9030 | 9031 | Chlorimuron-ethyl | L903C | 5.376 | 11 | 10 | 10000
ESI- | 9030 | 9031 | Novaluron | L903J | 5 | 10 | 5 | 10000
ESI- | D | 9031 | Ammelide | 68535 | 37.248 | 75 | 25 | 10000
ESI- | D | 9031 | Deisopropyliprodione | 68565 | 250 | 500 | 100 | 10000
ESI+ | A | 170
| E | 15
| D | 1
| 9030 | 0
| 9031 | 0
ESI- | A | 39
| E | 5
| D | 2
| 9030 | 3
| 9031 | 0

## T18_esipos_mdl_low_summary
Table 18. Estimated method detection levels for pesticide compounds determined by direct aqueous-injection liquid chromatography-tandem mass spectrometry in positive electrospray ionization mode calculated according to U.S. Environmental Protection Agency procedure from analysis of seven replicate 5 and 10 nanograms per liter standards determined in seven consecutive analytical batches [Method detection level (MDL) is calculated as standard deviation times t, where t is Student's t value for 49 replicates (t= 2.4); valid estimated MDLs must be greater than nominal spike concentration and have a detection frequency <50 percent. MDL selected was the lowest valid MDL of the two concentration levels; pesticide compounds are shown in bold if calculated MDL is invalid because it was greater than nominal spike concentration or a detection frequency <50 percent. <, less than; Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; USGS, U.S. Geological Survey; N, number of replicates; Std. Dev., standard deviation; RSD, relative standard deviation; min, minimum; na, not available] | [10 empty] | Note: These columns used to select valid MDL
[6 empty] | Concentration (ng/L) |  |  | Detections
Validation quality code | Lab code | Pesticide compound | USGS Parameter code | Nominal concentration of spike sample (ng/L) | N | Mean (ng/L) | Std Dev (ng/L) | MDL (ng/L) | Number of detections | Detection frequency (%) | Valid MDL (ng/L) | Valid MDL nominal conc. | Calculated MDL (ng/L)
E | 9030 | 1H-1,2,4-Triazole | 68498 | 5 | 49 | 102.3 | 186.34 |  | 17 | 34.6939 |  |  | 447.216
|  |  |  | 10 | 49 | 9.7 | 23.8 |  | 12 | 24.4898 |  |  | 57.12
E | 9030 | 2-(1-hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 5 | 49 | 28.9 | 23 |  | 37 | 75.5102 |  |  | 55.2
|  |  |  | 10 | 49 | 31.8 | 25.04 |  | 39 | 79.5918 |  |  | 60.096
A | 9030 | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 5 | 49 | 5.5 | 1.02 | 2.448 | 49 | 100 | 2.448 | 5 | 2.448 | 2.448
|  |  |  | 10 | 49 | 10.6 | 6.93 |  | 49 | 100 |  |  | 16.632
A | 9030 | 2-Amino-N-isopropylbenzamide | 68503 | 5 | 49 | 4.7 | 0.78 | 1.872 | 49 | 100 | 1.872 | 5 | 1.872 | 1.872
|  |  |  | 10 | 49 | 8.6 | 1.99 |  | 47 | 95.9184 | 4.776 |  | 4.776
A | 9030 | 2-Aminobenzimidazole | 68502 | 5 | 49 | 9.9 | 1.7 | 4.08 | 49 | 100 | 4.08 | 5 | 4.08 | 4.08
|  |  |  | 10 | 49 | 14.5 | 2.19 |  | 49 | 100 | 5.256 |  | 5.256
A | 9030 | 2-Chloro-2,6-diethylacetanilide | 68525 | 5 | 49 | 4.8 | 1.16 | 2.784 | 48 | 97.9592 | 2.784 | 5 | 2.784 | 2.784
|  |  |  | 10 | 49 | 9.3 | 1.53 |  | 49 | 100 | 3.672 |  | 3.672
A | 9030 | 2-Chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 5 | 49 | 4.5 | 1.13 | 2.712 | 48 | 97.9592 | 2.712 | 5 | 2.712 | 2.712
|  |  |  | 10 | 49 | 9.9 | 4.09 |  | 48 | 97.9592 | 9.816 |  | 9.816
A | 9030 | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 5 | 49 | 4.9 | 0.75 | 1.8 | 49 | 100 | 1.8 | 5 | 1.8 | 1.8
|  |  |  | 10 | 49 | 9.2 | 1.14 |  | 49 | 100 | 2.736 |  | 2.736
A | 9030 | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 5 | 49 | 5.4 | 1.17 | 2.808 | 49 | 100 | 2.808 | 5 | 2.808 | 2.808
|  |  |  | 10 | 49 | 10.1 | 2.02 |  | 49 | 100 | 4.848 |  | 4.848
A | 9030 | 2-Hydroxyatrazine | 68660 | 5 | 49 | 8 | 3.93 |  | 47 | 95.9184 |  |  | 9.432
|  |  |  | 10 | 49 | 12 | 5.01 |  | 47 | 95.9184 |  |  | 12.024
E | 9030 | 3-Hydroxycarbofuran | 68508 | 5 | 49 | 7.5 | 4.04 |  | 47 | 95.9184 |  |  | 9.696
|  |  |  | 10 | 49 | 11.7 | 3.27 | 7.848 | 49 | 100 | 7.848 | 10 | 7.848 | 7.848
A | 9030 | 3,4-Dichlorophenylurea | 68226 | 5 | 49 | 6.2 | 3.19 |  | 49 | 100 |  |  | 7.656
|  |  |  | 10 | 49 | 11 | 4.93 |  | 49 | 100 |  |  | 11.832
A | 9030 | 4-Chlorobenzylmethyl sulfoxide | 68514 | 5 | 49 | 4.7 | 0.98 | 2.352 | 49 | 100 | 2.352 | 5 | 2.352 | 2.352
|  |  |  | 10 | 49 | 8.5 | 1.32 |  | 49 | 100 | 3.168 |  | 3.168
A | 9030 | 4-Hydroxy molinate | 68515 | 5 | 49 | 3.8 | 1.41 | 3.384 | 48 | 97.9592 | 3.384 | 5 | 3.384 | 3.384
|  |  |  | 10 | 49 | 8.4 | 2.04 |  | 49 | 100 | 4.896 |  | 4.896
A | 9030 | 4-Hydroxy-tert-fluometuron | 68619 | 5 | 49 | 7.9 | 3.72 |  | 46 | 93.8776 |  |  | 8.928
|  |  |  | 10 | 49 | 13.5 | 4.14 | 9.936 | 49 | 100 | 9.936 | 10 | 9.936 | 9.936
A | 9030 | 4-Hydroxyhexazinone A | 68517 | 5 | 49 | 4.9 | 0.54 | 1.296 | 49 | 100 | 1.296 | 5 | 1.296 | 1.296
|  |  |  | 10 | 49 | 9.3 | 1.08 |  | 49 | 100 | 2.592 |  | 2.592
A | 9030 | Acephate | 68519 | 5 | 49 | 4.8 | 0.44 | 1.056 | 49 | 100 | 1.056 | 5 | 1.056 | 1.056
|  |  |  | 10 | 49 | 9 | 0.6 |  | 49 | 100 | 1.44 |  | 1.44
A | 9030 | Acetochlor | 68520 | 5 | 49 | 4.2 | 1.97 | 4.728 | 48 | 97.9592 | 4.728 | 5 | 4.728 | 4.728
|  |  |  | 10 | 49 | 8.9 | 2.23 |  | 49 | 100 | 5.352 |  | 5.352
A | 9030 | Alachlor | 65064 | 5 | 49 | 5.6 | 1.82 | 4.368 | 49 | 100 | 4.368 | 5 | 4.368 | 4.368
|  |  |  | 10 | 49 | 9.8 | 2.11 |  | 49 | 100 | 5.064 |  | 5.064
A | 9030 | Aldicarb | 68528 | 5 | 49 | 5.3 | 1.64 | 3.936 | 49 | 100 | 3.936 | 5 | 3.936 | 3.936
|  |  |  | 10 | 49 | 9.3 | 2.17 |  | 49 | 100 | 5.208 |  | 5.208
A | 9030 | Aldicarb sulfoxide | 68530 | 5 | 49 | 4.8 | 0.47 | 1.128 | 49 | 100 | 1.128 | 5 | 1.128 | 1.128
|  |  |  | 10 | 49 | 9.1 | 0.9 |  | 49 | 100 | 2.16 |  | 2.16
A | 9030 | Ametryn | 68533 | 5 | 49 | 4.8 | 0.55 | 1.32 | 49 | 100 | 1.32 | 5 | 1.32 | 1.32
|  |  |  | 10 | 49 | 8.9 | 0.99 |  | 49 | 100 | 2.376 |  | 2.376
E | 9030 | Asulam | 68536 | 5 | 49 | 13.4 | 9.21 |  | 41 | 83.6735 |  |  | 22.104
|  |  |  | 10 | 49 | 17.5 | 9.91 |  | 43 | 87.7551 |  |  | 23.784
A | 9030 | Atrazine | 65065 | 5 | 49 | 5.2 | 1.27 | 3.048 | 48 | 97.9592 | 3.048 | 5 | 3.048 | 3.048
|  |  |  | 10 | 49 | 9 | 1.93 |  | 48 | 97.9592 | 4.632 |  | 4.632
A | 9030 | Azinphos-methyl | 65066 | 5 | 49 | 4.7 | 1.51 | 3.624 | 49 | 100 | 3.624 | 5 | 3.624 | 3.624
|  |  |  | 10 | 49 | 9.2 | 1.94 |  | 49 | 100 | 4.656 |  | 4.656
A | 9030 | Azinphos-methyl oxon | 68211 | 5 | 49 | 8.2 | 2.41 |  | 49 | 100 |  |  | 5.784
|  |  |  | 10 | 49 | 13.4 | 2.96 | 7.104 | 49 | 100 | 7.104 | 10 | 7.104 | 7.104
A | 9030 | Azoxystrobin | 66589 | 5 | 49 | 5.5 | 2.05 | 4.92 | 49 | 100 | 4.92 | 5 | 4.92 | 4.92
|  |  |  | 10 | 49 | 8.7 | 0.58 | 1.392 | 49 | 100 | 1.392 | 10 | 1.392 | 1.392
E | 9030 | Bifenthrin | 65067 | 5 | 49 | 4.9 | 2.45 |  | 49 | 100 |  |  | 5.88
|  |  |  | 10 | 49 | 8 | 3.78 | 9.072 | 49 | 100 | 9.072 | 10 | 9.072 | 9.072
A | 9030 | Bromacil | 68542 | 5 | 49 | 5.1 | 1.21 | 2.904 | 49 | 100 | 2.904 | 5 | 2.904 | 2.904
|  |  |  | 10 | 49 | 9.4 | 1.74 |  | 49 | 100 | 4.176 |  | 4.176
E | 9030 | Butralin | 68545 | 5 | 49 | 4.8 | 0.87 | 2.088 | 49 | 100 | 2.088 | 5 | 2.088 | 2.088
|  |  |  | 10 | 49 | 8.9 | 1.13 |  | 49 | 100 | 2.712 |  | 2.712
A | 9030 | Butylate | 65068 | 5 | 49 | 7.1 | 1.14 | 2.736 | 49 | 100 | 2.736 | 5 | 2.736 | 2.736
|  |  |  | 10 | 49 | 11.3 | 1.82 |  | 49 | 100 | 4.368 |  | 4.368
A | 9030 | Carbaryl | 65069 | 5 | 49 | 4.5 | 1.25 | 3 | 49 | 100 | 3 | 5 | 3 | 3
|  |  |  | 10 | 49 | 8.6 | 1.48 |  | 49 | 100 | 3.552 |  | 3.552
E | 9030 | Carbendazim | 68548 | 5 | 49 | 4.6 | 0.52 | 1.248 | 49 | 100 | 1.248 | 5 | 1.248 | 1.248
|  |  |  | 10 | 49 | 8.7 | 0.78 |  | 49 | 100 | 1.872 |  | 1.872
A | 9030 | Carbofuran | 65070 | 5 | 49 | 4.7 | 0.68 | 1.632 | 49 | 100 | 1.632 | 5 | 1.632 | 1.632
|  |  |  | 10 | 49 | 8.9 | 0.87 |  | 49 | 100 | 2.088 |  | 2.088
A | 9030 | Carboxy molinate | 68549 | 5 | 49 | 18.9 | 6.1 |  | 49 | 100 |  |  | 14.64
|  |  |  | 10 | 49 | 21.9 | 6.9 |  | 49 | 100 |  |  | 16.56
A | 9030 | Chlorimuron-ethyl | 68872 | 5 | 49 | 11.1 | 8.94 |  | 42 | 85.7143 |  |  | 21.456
|  |  |  | 10 | 49 | 15.9 | 13.4 |  | 44 | 89.7959 |  |  | 32.16
A | 9030 | Chlorpyrifos | 65072 | 5 | 49 | 4.8 | 0.63 | 1.512 | 49 | 100 | 1.512 | 5 | 1.512 | 1.512
|  |  |  | 10 | 49 | 8.9 | 1.02 |  | 49 | 100 | 2.448 |  | 2.448
A | 9030 | Chlorpyrifos oxon | 68216 | 5 | 49 | 5 | 0.63 | 1.512 | 49 | 100 | 1.512 | 5 | 1.512 | 1.512
|  |  |  | 10 | 49 | 8.8 | 0.85 |  | 49 | 100 | 2.04 |  | 2.04
A | 9030 | Chlorsulfuron | 61678 | 5 | 49 | 182.6 | 141.19 |  | 37 | 75.5102 |  |  | 338.856
|  |  |  | 10 | 49 | 176.2 | 171.23 |  | 31 | 63.2653 |  |  | 410.952
A | 9030 | Cyanazine | 66592 | 5 | 49 | 8.2 | 3.64 |  | 47 | 95.9184 |  |  | 8.736
|  |  |  | 10 | 49 | 12.8 | 3.22 | 7.728 | 49 | 100 | 7.728 | 10 | 7.728 | 7.728
A | 9030 | Dechlorometolachlor | 68562 | 5 | 49 | 4.6 | 0.23 | 0.552 | 49 | 100 | 0.552 | 5 | 0.552 | 0.552
|  |  |  | 10 | 49 | 8.7 | 0.46 |  | 49 | 100 | 1.104 |  | 1.104
A | 9030 | Deethylatrazine | 68552 | 5 | 49 | 6.5 | 2.35 |  | 48 | 97.9592 |  |  | 5.64
|  |  |  | 10 | 49 | 10.6 | 4.05 | 9.72 | 46 | 93.8776 | 9.72 | 10 | 9.72 | 9.72
A | 9030 | Desiodo flubendiamide | 68563 | 5 | 49 | 4.8 | 0.77 | 1.848 | 49 | 100 | 1.848 | 5 | 1.848 | 1.848
|  |  |  | 10 | 49 | 9.4 | 1.3 |  | 49 | 100 | 3.12 |  | 3.12
A | 9030 | Deisopropyl prometryn | 68564 | 5 | 49 | 4.6 | 0.52 | 1.248 | 49 | 100 | 1.248 | 5 | 1.248 | 1.248
|  |  |  | 10 | 49 | 9.1 | 0.79 |  | 49 | 100 | 1.896 |  | 1.896
A | 9030 | Deisopropylatrazine | 68550 | 5 | 49 | 4.8 | 1.34 | 3.216 | 49 | 100 | 3.216 | 5 | 3.216 | 3.216
|  |  |  | 10 | 49 | 10.3 | 2.18 |  | 49 | 100 | 5.232 |  | 5.232
A | 9030 | Demethyl fluometuron | 68591 | 5 | 49 | 4.7 | 0.34 | 0.816 | 49 | 100 | 0.816 | 5 | 0.816 | 0.816
|  |  |  | 10 | 49 | 9.1 | 0.66 |  | 49 | 100 | 1.584 |  | 1.584
A | 9030 | Demethyl hexazinone B | 68566 | 5 | 49 | 4.7 | 0.51 | 1.224 | 49 | 100 | 1.224 | 5 | 1.224 | 1.224
|  |  |  | 10 | 49 | 8.7 | 1.48 |  | 48 | 97.9592 | 3.552 |  | 3.552
A | 9030 | Demethyl norflurazon | 68567 | 5 | 49 | 4.6 | 0.81 | 1.944 | 49 | 100 | 1.944 | 5 | 1.944 | 1.944
|  |  |  | 10 | 49 | 8.7 | 1.32 |  | 49 | 100 | 3.168 |  | 3.168
A | 9030 | Diazinon | 65078 | 5 | 49 | 4.8 | 0.5 | 1.2 | 49 | 100 | 1.2 | 5 | 1.2 | 1.2
|  |  |  | 10 | 49 | 8.9 | 0.8 |  | 49 | 100 | 1.92 |  | 1.92
E | 9030 | Diazinon oxon | 68236 | 5 | 49 | 23.5 | 25.61 |  | 29 | 59.1837 |  |  | 61.464
|  |  |  | 10 | 49 | 30.9 | 52.02 |  | 25 | 51.0204 |  |  | 124.848
A | 9030 | Dichlorvos | 68572 | 5 | 49 | 8.8 | 5.9 |  | 46 | 93.8776 |  |  | 14.16
|  |  |  | 10 | 49 | 14.3 | 6.03 |  | 49 | 100 |  |  | 14.472
A | 9030 | Dicrotophos | 68573 | 5 | 49 | 6.6 | 0.79 | 1.896 | 49 | 100 | 1.896 | 5 | 1.896 | 1.896
|  |  |  | 10 | 49 | 10.7 | 0.99 |  | 49 | 100 | 2.376 |  | 2.376
A | 9030 | Didealkylatrazine | 68547 | 5 | 49 | 7.9 | 7.72 |  | 36 | 73.4694 |  |  | 18.528
|  |  |  | 10 | 49 | 10.1 | 8.2 |  | 41 | 83.6735 |  |  | 19.68
A | 9030 | Diflubenzuron | 68576 | 5 | 49 | 5.1 | 1.07 | 2.568 | 49 | 100 | 2.568 | 5 | 2.568 | 2.568
|  |  |  | 10 | 49 | 9.8 | 1.6 |  | 49 | 100 | 3.84 |  | 3.84
A | 9030 | Dimethenamid | 68580 | 5 | 49 | 4.7 | 0.51 | 1.224 | 49 | 100 | 1.224 | 5 | 1.224 | 1.224
|  |  |  | 10 | 49 | 9.2 | 0.8 |  | 49 | 100 | 1.92 |  | 1.92
A | 9030 | Dimethoate | 66596 | 5 | 49 | 4.9 | 0.6 | 1.44 | 49 | 100 | 1.44 | 5 | 1.44 | 1.44
|  |  |  | 10 | 49 | 9.1 | 0.84 |  | 49 | 100 | 2.016 |  | 2.016
A | 9030 | Dimethoate oxon | 68661 | 5 | 49 | 4.8 | 0.36 | 0.864 | 49 | 100 | 0.864 | 5 | 0.864 | 0.864
|  |  |  | 10 | 49 | 9.2 | 0.67 |  | 49 | 100 | 1.608 |  | 1.608
A | 9030 | Disulfoton | 67595 | 5 | 49 | 8.6 | 2.76 |  | 49 | 100 |  |  | 6.624
|  |  |  | 10 | 49 | 12.2 | 2.6 | 6.24 | 49 | 100 | 6.24 | 10 | 6.24 | 6.24
A | 9030 | Disulfoton oxon | 68586 | 5 | 49 | 4.8 | 0.51 | 1.224 | 49 | 100 | 1.224 | 5 | 1.224 | 1.224
|  |  |  | 10 | 49 | 8.9 | 1.56 |  | 48 | 97.9592 | 3.744 |  | 3.744
A | 9030 | Disulfoton oxon sulfone | 68588 | 5 | 49 | 5.2 | 1.38 | 3.312 | 49 | 100 | 3.312 | 5 | 3.312 | 3.312
|  |  |  | 10 | 49 | 9.8 | 1.83 |  | 49 | 100 | 4.392 |  | 4.392
A | 9030 | Disulfoton oxon sulfoxide | 68587 | 5 | 49 | 4.7 | 1.79 | 4.296 | 49 | 100 | 4.296 | 5 | 4.296 | 4.296
|  |  |  | 10 | 49 | 9.4 | 1.31 | 3.144 | 49 | 100 | 3.144 | 10 | 3.144 | 3.144
A | 9030 | Disulfoton sulfone | 68589 | 5 | 49 | 6.2 | 1.78 | 4.272 | 49 | 100 | 4.272 | 5 | 4.272 | 4.272
|  |  |  | 10 | 49 | 10.3 | 3.05 |  | 49 | 100 | 7.32 |  | 7.32
A | 9030 | Disulfoton sulfoxide | 68590 | 5 | 49 | 4.7 | 0.92 | 2.208 | 48 | 97.9592 | 2.208 | 5 | 2.208 | 2.208
|  |  |  | 10 | 49 | 8.9 | 2.06 |  | 47 | 95.9184 | 4.944 |  | 4.944
A | 9030 | Diuron | 66598 | 5 | 49 | 4.7 | 1.54 | 3.696 | 46 | 93.8776 | 3.696 | 5 | 3.696 | 3.696
|  |  |  | 10 | 49 | 8.4 | 2.4 |  | 46 | 93.8776 | 5.76 |  | 5.76
A | 9030 | EPTC | 65080 | 5 | 49 | 9.1 | 7.83 |  | 36 | 73.4694 |  |  | 18.792
|  |  |  | 10 | 49 | 13 | 10.53 |  | 43 | 87.7551 |  |  | 25.272
A | 9030 | EPTC R248722 | 68594 | 5 | 49 | 4.9 | 0.71 | 1.704 | 49 | 100 | 1.704 | 5 | 1.704 | 1.704
|  |  |  | 10 | 49 | 9.1 | 0.92 |  | 49 | 100 | 2.208 |  | 2.208
A | 9030 | Ethoprop | 68596 | 5 | 49 | 4.8 | 0.86 | 2.064 | 49 | 100 | 2.064 | 5 | 2.064 | 2.064
|  |  |  | 10 | 49 | 8.9 | 1.18 |  | 49 | 100 | 2.832 |  | 2.832
A | 9030 | Etoxazole | 68598 | 5 | 49 | 4.7 | 0.32 | 0.768 | 49 | 100 | 0.768 | 5 | 0.768 | 0.768
|  |  |  | 10 | 49 | 8.7 | 0.7 |  | 49 | 100 | 1.68 |  | 1.68
E | 9031 | Famoxadone | 67609 | 5 | 49 | 8.1 | 5.08 |  | 42 | 85.7143 |  |  | 12.192
|  |  |  | 10 | 49 | 12.5 | 6.44 |  | 48 | 97.9592 |  |  | 15.456
A | 9030 | Fenamiphos | 68599 | 5 | 49 | 4.6 | 0.79 | 1.896 | 48 | 97.9592 | 1.896 | 5 | 1.896 | 1.896
|  |  |  | 10 | 49 | 8.8 | 0.63 | 1.512 | 49 | 100 | 1.512 | 10 | 1.512 | 1.512
A | 9030 | Fenamiphos sulfone | 68600 | 5 | 49 | 4.5 | 0.86 | 2.064 | 49 | 100 | 2.064 | 5 | 2.064 | 2.064
|  |  |  | 10 | 49 | 8.8 | 1.21 |  | 49 | 100 | 2.904 |  | 2.904
A | 9030 | Fenamiphos sulfoxide | 68601 | 5 | 49 | 5 | 1.03 | 2.472 | 49 | 100 | 2.472 | 5 | 2.472 | 2.472
|  |  |  | 10 | 49 | 9.1 | 1.61 |  | 49 | 100 | 3.864 |  | 3.864
E | 9030 | Fenbutatin oxide | 68602 | 5 | 49 | 18.8 | 21.59 |  | 42 | 85.7143 |  |  | 51.816
|  |  |  | 10 | 49 | 9.5 | 8.27 |  | 45 | 91.8367 |  |  | 19.848
A | 9030 | Fentin | 68603 | 5 | 49 | 5.9 | 2.8 |  | 48 | 97.9592 |  |  | 6.72
|  |  |  | 10 | 49 | 10.2 | 3.27 | 7.848 | 49 | 100 | 7.848 | 10 | 7.848 | 7.848
E | 9030 | Flumetsulam | 61679 | 5 | 49 | 7.8 | 2.41 |  | 49 | 100 |  |  | 5.784
|  |  |  | 10 | 49 | 12 | 3.38 | 8.112 | 49 | 100 | 8.112 | 10 | 8.112 | 8.112
A | 9030 | Fluometuron | 68608 | 5 | 49 | 5.1 | 1.2 | 2.88 | 47 | 95.9184 | 2.88 | 5 | 2.88 | 2.88
|  |  |  | 10 | 49 | 9.6 | 0.98 | 2.352 | 49 | 100 | 2.352 | 10 | 2.352 | 2.352
A | 9030 | Fonofos | 65084 | 5 | 49 | 6.3 | 2.17 | 5.208 | 49 | 100 | 5.208 | 5 | 5.208 | 5.208
|  |  |  | 10 | 49 | 10.7 | 2.51 |  | 49 | 100 | 6.024 |  | 6.024
A | 9030 | Halosulfuron-methyl | 61680 | 5 | 49 | 6.4 | 3.42 |  | 48 | 97.9592 |  |  | 8.208
|  |  |  | 10 | 49 | 10.3 | 4.46 | 10.704 | 49 | 100 | 10.704 | 10 | 10.704 | 10.704
A | 9030 | Hexazinone | 65085 | 5 | 49 | 4.8 | 0.55 | 1.32 | 49 | 100 | 1.32 | 5 | 1.32 | 1.32
|  |  |  | 10 | 49 | 8.9 | 2.02 |  | 47 | 95.9184 | 4.848 |  | 4.848
A | 9030 | Hexazinone TP C | 68612 | 5 | 49 | 4.8 | 0.38 | 0.912 | 49 | 100 | 0.912 | 5 | 0.912 | 0.912
|  |  |  | 10 | 49 | 9.1 | 0.89 |  | 49 | 100 | 2.136 |  | 2.136
A | 9030 | Hexazinone TP F | 68574 | 5 | 49 | 5.7 | 2.19 | 5.256 | 48 | 97.9592 | 5.256 | 5 | 5.256 | 5.256
|  |  |  | 10 | 49 | 10.6 | 2.56 |  | 49 | 100 | 6.144 |  | 6.144
A | 9030 | Hexazinone TP G | 68713 | 5 | 49 | 7.5 | 4.2 |  | 45 | 91.8367 |  |  | 10.08
|  |  |  | 10 | 49 | 10.8 | 4.71 |  | 47 | 95.9184 |  |  | 11.304
A | 9030 | Hydroxy mono demethyl fluometuron | 68617 | 5 | 49 | 4.8 | 2.07 | 4.968 | 49 | 100 | 4.968 | 5 | 4.968 | 4.968
|  |  |  | 10 | 49 | 9.5 | 2.6 |  | 48 | 97.9592 | 6.24 |  | 6.24
A | 9030 | Hydroxyacetochlor | 68615 | 5 | 49 | 5.6 | 1.97 | 4.728 | 48 | 97.9592 | 4.728 | 5 | 4.728 | 4.728
|  |  |  | 10 | 49 | 9.1 | 3.31 |  | 46 | 93.8776 | 7.944 |  | 7.944
A | 9030 | Hydroxyalachlor | 68616 | 5 | 49 | 5.4 | 1.57 | 3.768 | 49 | 100 | 3.768 | 5 | 3.768 | 3.768
|  |  |  | 10 | 49 | 9.9 | 2.86 |  | 49 | 100 | 6.864 |  | 6.864
A | 9030 | Hydroxydiazinon | 68618 | 5 | 49 | 6.3 | 2.28 | 5.472 | 49 | 100 | 5.472 | 5 | 5.472 | 5.472
|  |  |  | 10 | 49 | 10.8 | 3.69 |  | 49 | 100 | 8.856 |  | 8.856
A | 9030 | Hydroxyfluometuron | 68620 | 5 | 49 | 5.8 | 1.92 | 4.608 | 47 | 95.9184 | 4.608 | 5 | 4.608 | 4.608
|  |  |  | 10 | 49 | 10.3 | 2.36 |  | 48 | 97.9592 | 5.664 |  | 5.664
A | 9030 | Hydroxymetolachlor | 68622 | 5 | 49 | 4.7 | 0.37 | 0.888 | 49 | 100 | 0.888 | 5 | 0.888 | 0.888
|  |  |  | 10 | 49 | 8.8 | 0.78 |  | 49 | 100 | 1.872 |  | 1.872
A | 9030 | Hydroxyphthalazinone | 68623 | 5 | 49 | 6.7 | 3.7 |  | 45 | 91.8367 |  |  | 8.88
|  |  |  | 10 | 49 | 11.8 | 4.89 |  | 49 | 100 |  |  | 11.736
A | 9030 | Hydroxysimazine | 68624 | 5 | 49 | 5.1 | 0.52 | 1.248 | 49 | 100 | 1.248 | 5 | 1.248 | 1.248
|  |  |  | 10 | 49 | 9.4 | 0.91 |  | 49 | 100 | 2.184 |  | 2.184
A | 9030 | Hydroxytebuthurion | 68621 | 5 | 49 | 8.4 | 2.23 | 5.352 | 49 | 100 | 5.352 | 5 | 5.352 | 5.352
|  |  |  | 10 | 49 | 12.8 | 3.06 |  | 49 | 100 | 7.344 |  | 7.344
A | 9030 | Imazamox | 68625 | 5 | 49 | 11.2 | 3.8 |  | 49 | 100 |  |  | 9.12
|  |  |  | 10 | 49 | 12.4 | 5.46 |  | 48 | 97.9592 |  |  | 13.104
A | 9030 | Imazaquin | 61682 | 5 | 49 | 4.9 | 2.58 |  | 49 | 100 |  |  | 6.192
|  |  |  | 10 | 49 | 9.2 | 3.78 | 9.072 | 49 | 100 | 9.072 | 10 | 9.072 | 9.072
A | 9030 | Imazethapyr | 61683 | 5 | 49 | 6.2 | 3.01 |  | 49 | 100 |  |  | 7.224
|  |  |  | 10 | 49 | 10.5 | 3.97 | 9.528 | 49 | 100 | 9.528 | 10 | 9.528 | 9.528
A | 9030 | Imidacloprid | 68426 | 5 | 49 | 8 | 2.25 | 5.4 | 49 | 100 | 5.4 | 5 | 5.4 | 5.4
|  |  |  | 10 | 49 | 12.3 | 3.17 |  | 49 | 100 | 7.608 |  | 7.608
A | 9030 | Indoxacarb | 68627 | 5 | 49 | 6.7 | 1.59 | 3.816 | 49 | 100 | 3.816 | 5 | 3.816 | 3.816
|  |  |  | 10 | 49 | 10.6 | 1.88 |  | 49 | 100 | 4.512 |  | 4.512
A | 9030 | Isoxaflutole | 68632 | 5 | 49 | 4.3 | 2.65 |  | 48 | 97.9592 |  |  | 6.36
|  |  |  | 10 | 49 | 8.8 | 3.17 | 7.608 | 47 | 95.9184 | 7.608 | 10 | 7.608 | 7.608
A | 9030 | Kresoxim-methyl | 67670 | 5 | 49 | 5 | 0.95 | 2.28 | 49 | 100 | 2.28 | 5 | 2.28 | 2.28
|  |  |  | 10 | 49 | 9.2 | 1.79 |  | 49 | 100 | 4.296 |  | 4.296
A | 9030 | Lactofen | 68638 | 5 | 49 | 5.3 | 0.6 | 1.44 | 49 | 100 | 1.44 | 5 | 1.44 | 1.44
|  |  |  | 10 | 49 | 9.1 | 1 |  | 49 | 100 | 2.4 |  | 2.4
A | 9030 | Linuron | 68639 | 5 | 49 | 4.5 | 1.07 | 2.568 | 49 | 100 | 2.568 | 5 | 2.568 | 2.568
|  |  |  | 10 | 49 | 8.9 | 1.45 |  | 49 | 100 | 3.48 |  | 3.48
A | 9030 | Malaoxon | 68240 | 5 | 49 | 4.7 | 0.51 | 1.224 | 49 | 100 | 1.224 | 5 | 1.224 | 1.224
|  |  |  | 10 | 49 | 8.9 | 0.94 |  | 49 | 100 | 2.256 |  | 2.256
A | 9030 | Malathion | 65087 | 5 | 49 | 5.3 | 1.1 | 2.64 | 48 | 97.9592 | 2.64 | 5 | 2.64 | 2.64
|  |  |  | 10 | 49 | 9.5 | 0.91 | 2.184 | 49 | 100 | 2.184 | 10 | 2.184 | 2.184
A | 9030 | Metalaxyl | 68437 | 5 | 49 | 4.7 | 1.33 | 3.192 | 48 | 97.9592 | 3.192 | 5 | 3.192 | 3.192
|  |  |  | 10 | 49 | 8.7 | 2 |  | 48 | 97.9592 | 4.8 |  | 4.8
A | 9030 | Metconazole | 66620 | 5 | 49 | 5.8 | 0.9 | 2.16 | 49 | 100 | 2.16 | 5 | 2.16 | 2.16
|  |  |  | 10 | 49 | 9.9 | 0.93 |  | 49 | 100 | 2.232 |  | 2.232
A | 9030 | Methamidophos | 68644 | 5 | 49 | 5.1 | 0.53 | 1.272 | 49 | 100 | 1.272 | 5 | 1.272 | 1.272
|  |  |  | 10 | 49 | 9.3 | 0.67 |  | 49 | 100 | 1.608 |  | 1.608
A | 9030 | Methidathion | 65088 | 5 | 49 | 4.8 | 0.83 | 1.992 | 49 | 100 | 1.992 | 5 | 1.992 | 1.992
|  |  |  | 10 | 49 | 9.3 | 1.05 |  | 49 | 100 | 2.52 |  | 2.52
A | 9030 | Methomyl | 68645 | 5 | 49 | 5 | 0.45 | 1.08 | 49 | 100 | 1.08 | 5 | 1.08 | 1.08
|  |  |  | 10 | 49 | 9.2 | 0.74 |  | 49 | 100 | 1.776 |  | 1.776
A | 9030 | Methomyl oxime | 68646 | 5 | 49 | 57.2 | 62.94 |  | 30 | 61.2245 |  |  | 151.056
|  |  |  | 10 | 49 | 63 | 60.49 |  | 33 | 67.3469 |  |  | 145.176
A | 9030 | Methoxyfenozide | 68647 | 5 | 49 | 4.8 | 0.4 | 0.96 | 49 | 100 | 0.96 | 5 | 0.96 | 0.96
|  |  |  | 10 | 49 | 9.2 | 0.85 |  | 49 | 100 | 2.04 |  | 2.04
A | 9030 | Metolachlor | 65090 | 5 | 49 | 3.7 | 1.75 | 4.2 | 42 | 85.7143 | 4.2 | 5 | 4.2 | 4.2
|  |  |  | 10 | 49 | 7.5 | 3.77 |  | 44 | 89.7959 | 9.048 |  | 9.048
A | 9030 | Metolachlor hydroxy morpholinone | 68649 | 5 | 49 | 4.5 | 1.94 | 4.656 | 44 | 89.7959 | 4.656 | 5 | 4.656 | 4.656
|  |  |  | 10 | 49 | 9.2 | 2.15 |  | 48 | 97.9592 | 5.16 |  | 5.16
A | 9030 | Metribuzin | 68652 | 5 | 49 | 4.8 | 2.87 |  | 45 | 91.8367 |  |  | 6.888
|  |  |  | 10 | 49 | 8.4 | 4.06 | 9.744 | 47 | 95.9184 | 9.744 | 10 | 9.744 | 9.744
A | 9030 | Metribuzin DA | 68568 | 5 | 49 | 8.7 | 1.83 | 4.392 | 49 | 100 | 4.392 | 5 | 4.392 | 4.392
|  |  |  | 10 | 49 | 11.8 | 3.5 |  | 48 | 97.9592 | 8.4 |  | 8.4
A | 9030 | Molinate | 65091 | 5 | 49 | 7.5 | 3.27 |  | 48 | 97.9592 |  |  | 7.848
|  |  |  | 10 | 49 | 12 | 3.56 | 8.544 | 49 | 100 | 8.544 | 10 | 8.544 | 8.544
A | 9030 | Myclobutanil | 66632 | 5 | 49 | 5.5 | 1.41 | 3.384 | 49 | 100 | 3.384 | 5 | 3.384 | 3.384
|  |  |  | 10 | 49 | 10.2 | 1.77 |  | 49 | 100 | 4.248 |  | 4.248
A | 9030 | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 5 | 49 | 4.5 | 1.55 | 3.72 | 46 | 93.8776 | 3.72 | 5 | 3.72 | 3.72
|  |  |  | 10 | 49 | 8.2 | 2.34 |  | 47 | 95.9184 | 5.616 |  | 5.616
E | 9030 | Naled | 68654 | 5 | 49 | 15.4 | 6.86 |  | 43 | 87.7551 |  |  | 16.464
|  |  |  | 10 | 49 | 18.7 | 8.37 |  | 48 | 97.9592 |  |  | 20.088
A | 9030 | Nicosulfuron | 61685 | 5 | 49 | 6.9 | 5.51 |  | 43 | 87.7551 |  |  | 13.224
|  |  |  | 10 | 49 | 10.3 | 6.39 |  | 48 | 97.9592 |  |  | 15.336
A | 9030 | Norflurazon | 67685 | 5 | 49 | 4.7 | 0.63 | 1.512 | 49 | 100 | 1.512 | 5 | 1.512 | 1.512
|  |  |  | 10 | 49 | 8.8 | 0.95 |  | 49 | 100 | 2.28 |  | 2.28
E | 9030 | Novaluron | 68655 | 5 | 49 | 5 | 1.68 | 4.032 | 49 | 100 | 4.032 | 5 | 4.032 | 4.032
|  |  |  | 10 | 49 | 9 | 3.18 |  | 49 | 100 | 7.632 |  | 7.632
A | 9030 | O-Ethyl-O-methyl-S-propyl phosphorothioate | 68597 | 5 | 49 | 4.8 | 1.17 | 2.808 | 48 | 97.9592 | 2.808 | 5 | 2.808 | 2.808
|  |  |  | 10 | 49 | 9.2 | 1.11 | 2.664 | 49 | 100 | 2.664 | 10 | 2.664 | 2.664
A | 9030 | O-Ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 5 | 49 | 4.9 | 0.85 | 2.04 | 49 | 100 | 2.04 | 5 | 2.04 | 2.04
|  |  |  | 10 | 49 | 9.1 | 1.08 |  | 49 | 100 | 2.592 |  | 2.592
A | 9030 | O-Ethyl-S-propyl phosphorothioate | 68658 | 5 | 49 | 15.4 | 10.49 |  | 45 | 91.8367 |  |  | 25.176
|  |  |  | 10 | 49 | 16.9 | 10.32 |  | 46 | 93.8776 |  |  | 24.768
A | 9030 | Orthosulfamuron | 68662 | 5 | 49 | 17.3 | 19.6 |  | 34 | 69.3878 |  |  | 47.04
|  |  |  | 10 | 49 | 26.3 | 49.6 |  | 41 | 83.6735 |  |  | 119.04
A | 9030 | Oxamyl | 68664 | 5 | 49 | 4.8 | 0.34 | 0.816 | 49 | 100 | 0.816 | 5 | 0.816 | 0.816
|  |  |  | 10 | 49 | 9 | 0.58 |  | 49 | 100 | 1.392 |  | 1.392
A | 9030 | Oxamyl oxime | 68665 | 5 | 49 | 4.8 | 0.29 | 0.696 | 49 | 100 | 0.696 | 5 | 0.696 | 0.696
|  |  |  | 10 | 49 | 9.1 | 0.59 |  | 49 | 100 | 1.416 |  | 1.416
E | 9030 | Oxyfluorfen | 65093 | 5 | 49 | 12.3 | 10.78 |  | 36 | 73.4694 |  |  | 25.872
|  |  |  | 10 | 49 | 16.6 | 13.48 |  | 38 | 77.551 |  |  | 32.352
A | 9030 | Paraoxon | 68666 | 5 | 49 | 4.6 | 0.64 | 1.536 | 49 | 100 | 1.536 | 5 | 1.536 | 1.536
|  |  |  | 10 | 49 | 8.9 | 0.88 |  | 49 | 100 | 2.112 |  | 2.112
A | 9030 | Paraoxon-methyl | 68648 | 5 | 49 | 7.9 | 3.61 |  | 49 | 100 |  |  | 8.664
|  |  |  | 10 | 49 | 11.5 | 4.19 | 10.056 | 48 | 97.9592 | 10.056 | 10 | 10.056 | 10.056
E | 9030 | Parathion-methyl | 65089 | 5 | 49 | 183.8 | 171.59 |  | 33 | 67.3469 |  |  | 411.816
|  |  |  | 10 | 49 | 186.1 | 162.38 |  | 37 | 75.5102 |  |  | 389.712
A | 9030 | Pendimethalin | 65098 | 5 | 49 | 6.5 | 0.59 | 1.416 | 49 | 100 | 1.416 | 5 | 1.416 | 1.416
|  |  |  | 10 | 49 | 10.9 | 1.24 |  | 49 | 100 | 2.976 |  | 2.976
A | 9030 | Phorate | 68668 | 5 | 49 | 6.1 | 1.25 | 3 | 49 | 100 | 3 | 5 | 3 | 3
|  |  |  | 10 | 49 | 9.7 | 1.61 |  | 49 | 100 | 3.864 |  | 3.864
A | 9030 | Phorate oxon | 68669 | 5 | 49 | 73.2 | 17.9 |  | 49 | 100 |  |  | 42.96
|  |  |  | 10 | 49 | 77 | 17.87 |  | 49 | 100 |  |  | 42.888
A | 9030 | Phorate oxon sulfone | 68670 | 5 | 49 | 6.5 | 2.81 |  | 49 | 100 |  |  | 6.744
|  |  |  | 10 | 49 | 9.9 | 4.05 | 9.72 | 49 | 100 | 9.72 | 10 | 9.72 | 9.72
A | 9030 | Phorate oxon sulfoxide | 68671 | 5 | 49 | 7 | 1.43 | 3.432 | 49 | 100 | 3.432 | 5 | 3.432 | 3.432
|  |  |  | 10 | 49 | 11.3 | 1.76 |  | 49 | 100 | 4.224 |  | 4.224
A | 9030 | Phorate sulfone | 68672 | 5 | 49 | 5.4 | 1.85 | 4.44 | 49 | 100 | 4.44 | 5 | 4.44 | 4.44
|  |  |  | 10 | 49 | 9 | 2.66 |  | 49 | 100 | 6.384 |  | 6.384
A | 9030 | Phorate sulfoxide | 68673 | 5 | 49 | 4.7 | 1.21 | 2.904 | 48 | 97.9592 | 2.904 | 5 | 2.904 | 2.904
|  |  |  | 10 | 49 | 9 | 1.79 |  | 48 | 97.9592 | 4.296 |  | 4.296
A | 9030 | Phthalazinone | 68675 | 5 | 49 | 4.9 | 1.13 | 2.712 | 49 | 100 | 2.712 | 5 | 2.712 | 2.712
|  |  |  | 10 | 49 | 9.6 | 1.76 |  | 49 | 100 | 4.224 |  | 4.224
A | 9030 | Piperonyl butoxide | 65102 | 5 | 49 | 5 | 0.73 | 1.752 | 49 | 100 | 1.752 | 5 | 1.752 | 1.752
|  |  |  | 10 | 49 | 9.4 | 3.84 |  | 49 | 100 | 9.216 |  | 9.216
A | 9030 | Profenofos | 68676 | 5 | 49 | 4.6 | 0.52 | 1.248 | 49 | 100 | 1.248 | 5 | 1.248 | 1.248
|  |  |  | 10 | 49 | 8.9 | 0.96 |  | 49 | 100 | 2.304 |  | 2.304
A | 9030 | Prometon | 67702 | 5 | 49 | 4.6 | 0.92 | 2.208 | 48 | 97.9592 | 2.208 | 5 | 2.208 | 2.208
|  |  |  | 10 | 49 | 9 | 1.43 |  | 49 | 100 | 3.432 |  | 3.432
A | 9030 | Prometryn | 65103 | 5 | 49 | 4.6 | 0.4 | 0.96 | 49 | 100 | 0.96 | 5 | 0.96 | 0.96
|  |  |  | 10 | 49 | 8.7 | 0.73 |  | 49 | 100 | 1.752 |  | 1.752
A | 9030 | Pronamide | 67706 | 5 | 49 | 5 | 0.8 | 1.92 | 49 | 100 | 1.92 | 5 | 1.92 | 1.92
|  |  |  | 10 | 49 | 9.2 | 1.18 |  | 49 | 100 | 2.832 |  | 2.832
A | 9030 | Propanil | 66641 | 5 | 49 | 6.1 | 2.23 | 5.352 | 49 | 100 | 5.352 | 5 | 5.352 | 5.352
|  |  |  | 10 | 49 | 9.9 | 2.36 |  | 49 | 100 | 5.664 |  | 5.664
A | 9030 | Propargite | 68677 | 5 | 49 | 4.7 | 0.33 | 0.792 | 49 | 100 | 0.792 | 5 | 0.792 | 0.792
|  |  |  | 10 | 49 | 8.8 | 0.74 |  | 49 | 100 | 1.776 |  | 1.776
A | 9030 | Propazine | 68678 | 5 | 49 | 4.6 | 0.68 | 1.632 | 49 | 100 | 1.632 | 5 | 1.632 | 1.632
|  |  |  | 10 | 49 | 8.8 | 1.05 |  | 49 | 100 | 2.52 |  | 2.52
A | 9030 | Propiconazole | 66643 | 5 | 49 | 4.9 | 1.19 | 2.856 | 49 | 100 | 2.856 | 5 | 2.856 | 2.856
|  |  |  | 10 | 49 | 9.4 | 2.49 |  | 49 | 100 | 5.976 |  | 5.976
A | 9030 | Propoxur | 68679 | 5 | 49 | 4.6 | 1.17 | 2.808 | 47 | 95.9184 | 2.808 | 5 | 2.808 | 2.808
|  |  |  | 10 | 49 | 8.9 | 1.04 | 2.496 | 49 | 100 | 2.496 | 10 | 2.496 | 2.496
A | 9030 | Pymetrozine | 68680 | 5 | 49 | 4.7 | 0.39 | 0.936 | 49 | 100 | 0.936 | 5 | 0.936 | 0.936
|  |  |  | 10 | 49 | 9 | 0.68 |  | 49 | 100 | 1.632 |  | 1.632
A | 9030 | Pyraclostrobin | 66646 | 5 | 49 | 4.8 | 0.47 | 1.128 | 49 | 100 | 1.128 | 5 | 1.128 | 1.128
|  |  |  | 10 | 49 | 8.8 | 0.75 |  | 49 | 100 | 1.8 |  | 1.8
A | 9030 | Pyridaben | 68682 | 5 | 49 | 4.7 | 0.45 | 1.08 | 49 | 100 | 1.08 | 5 | 1.08 | 1.08
|  |  |  | 10 | 49 | 8.6 | 0.88 |  | 49 | 100 | 2.112 |  | 2.112
A | 9030 | Pyrimidinol | 68505 | 5 | 49 | 7.3 | 2.31 |  | 49 | 100 |  |  | 5.544
|  |  |  | 10 | 49 | 12.5 | 4.09 | 9.816 | 49 | 100 | 9.816 | 10 | 9.816 | 9.816
A | 9030 | Pyriproxyfen | 68683 | 5 | 49 | 5 | 0.63 | 1.512 | 49 | 100 | 1.512 | 5 | 1.512 | 1.512
|  |  |  | 10 | 49 | 9 | 0.52 | 1.248 | 49 | 100 | 1.248 | 10 | 1.248 | 1.248
A | 9030 | Siduron | 68686 | 5 | 49 | 5 | 1.19 | 2.856 | 48 | 97.9592 | 2.856 | 5 | 2.856 | 2.856
|  |  |  | 10 | 49 | 9.4 | 1.34 |  | 49 | 100 | 3.216 |  | 3.216
A | 9030 | Simazine | 65105 | 5 | 49 | 5.9 | 1.93 | 4.632 | 49 | 100 | 4.632 | 5 | 4.632 | 4.632
|  |  |  | 10 | 49 | 10.8 | 2.03 |  | 49 | 100 | 4.872 |  | 4.872
A | 9030 | Sulfometuron-methyl | 68688 | 5 | 49 | 5.6 | 0.76 | 1.824 | 49 | 100 | 1.824 | 5 | 1.824 | 1.824
|  |  |  | 10 | 49 | 9.6 | 1.21 |  | 49 | 100 | 2.904 |  | 2.904
A | 9030 | Sulfosulfuron | 68689 | 5 | 49 | 5.2 | 2.21 | 5.304 | 46 | 93.8776 | 5.304 | 5 | 5.304 | 5.304
|  |  |  | 10 | 49 | 8.7 | 2.97 |  | 47 | 95.9184 | 7.128 |  | 7.128
A | 9030 | Sulfosulfuron ethyl sulfone | 68690 | 5 | 49 | 4.7 | 0.45 | 1.08 | 49 | 100 | 1.08 | 5 | 1.08 | 1.08
|  |  |  | 10 | 49 | 9 | 0.64 |  | 49 | 100 | 1.536 |  | 1.536
A | 9030 | tebuconazole | 66649 | 5 | 49 | 5 | 1.03 | 2.472 | 49 | 100 | 2.472 | 5 | 2.472 | 2.472
|  |  |  | 10 | 49 | 9.3 | 1.33 |  | 49 | 100 | 3.192 |  | 3.192
A | 9030 | Tebufenozide | 68692 | 5 | 49 | 4.7 | 0.33 | 0.792 | 49 | 100 | 0.792 | 5 | 0.792 | 0.792
|  |  |  | 10 | 49 | 8.9 | 0.61 |  | 49 | 100 | 1.464 |  | 1.464
A | 9030 | Tebupirimfos | 68693 | 5 | 49 | 4.7 | 0.35 | 0.84 | 49 | 100 | 0.84 | 5 | 0.84 | 0.84
|  |  |  | 10 | 49 | 8.7 | 0.58 |  | 49 | 100 | 1.392 |  | 1.392
A | 9030 | Tebupirimfos oxon | 68694 | 5 | 49 | 4.8 | 0.32 | 0.768 | 49 | 100 | 0.768 | 5 | 0.768 | 0.768
|  |  |  | 10 | 49 | 8.9 | 0.56 |  | 49 | 100 | 1.344 |  | 1.344
A | 9030 | Tebuthiuron | 68695 | 5 | 49 | 4.6 | 0.6 | 1.44 | 49 | 100 | 1.44 | 5 | 1.44 | 1.44
|  |  |  | 10 | 49 | 8.7 | 0.96 |  | 49 | 100 | 2.304 |  | 2.304
A | 9030 | Tebuthiuron TP 104 | 68575 | 5 | 49 | 5 | 1.14 | 2.736 | 49 | 100 | 2.736 | 5 | 2.736 | 2.736
|  |  |  | 10 | 49 | 9.8 | 1.48 |  | 49 | 100 | 3.552 |  | 3.552
A | 9030 | Tebuthiuron TP 106 | 68714 | 5 | 49 | 8.3 | 4.48 |  | 49 | 100 |  |  | 10.752
|  |  |  | 10 | 49 | 12 | 5.61 |  | 48 | 97.9592 |  |  | 13.464
A | 9030 | Tebuthiuron TP 108 | 68696 | 5 | 49 | 7.4 | 1.89 | 4.536 | 49 | 100 | 4.536 | 5 | 4.536 | 4.536
|  |  |  | 10 | 49 | 11.1 | 1.95 |  | 49 | 100 | 4.68 |  | 4.68
A | 9030 | Tebuthiuron TP 109 (OH) | 68697 | 5 | 49 | 7.7 | 3.33 |  | 49 | 100 |  |  | 7.992
|  |  |  | 10 | 49 | 12.3 | 3.47 | 8.328 | 49 | 100 | 8.328 | 10 | 8.328 | 8.328
A | 9030 | Terbacil | 68698 | 5 | 49 | 6.3 | 2.9 |  | 47 | 95.9184 |  |  | 6.96
|  |  |  | 10 | 49 | 10.8 | 4.35 | 10.44 | 49 | 100 | 10.44 | 10 | 10.44 | 10.44
A | 9030 | Terbufos | 68699 | 5 | 49 | 5.3 | 0.68 | 1.632 | 49 | 100 | 1.632 | 5 | 1.632 | 1.632
|  |  |  | 10 | 49 | 9.5 | 0.84 |  | 49 | 100 | 2.016 |  | 2.016
A | 9030 | Terbufos oxon | 68700 | 5 | 49 | 4.6 | 0.63 | 1.512 | 49 | 100 | 1.512 | 5 | 1.512 | 1.512
|  |  |  | 10 | 49 | 8.3 | 1.24 |  | 49 | 100 | 2.976 |  | 2.976
A | 9030 | Terbufos oxon sulfone | 68701 | 5 | 49 | 4.1 | 2.09 | 5.016 | 49 | 100 | 5.016 | 5 | 5.016 | 5.016
|  |  |  | 10 | 49 | 8 | 3.25 |  | 48 | 97.9592 | 7.8 |  | 7.8
A | 9030 | Terbufos oxon sulfoxide | 68702 | 5 | 49 | 4.6 | 1.03 | 2.472 | 48 | 97.9592 | 2.472 | 5 | 2.472 | 2.472
|  |  |  | 10 | 49 | 8.6 | 2.05 |  | 47 | 95.9184 | 4.92 |  | 4.92
A | 9030 | Terbufos sulfone | 68703 | 5 | 49 | 8 | 2.23 | 5.352 | 49 | 100 | 5.352 | 5 | 5.352 | 5.352
|  |  |  | 10 | 49 | 12.9 | 3.13 |  | 49 | 100 | 7.512 |  | 7.512
A | 9030 | Terbufos sulfoxide | 68704 | 5 | 49 | 4.4 | 1.01 | 2.424 | 47 | 95.9184 | 2.424 | 5 | 2.424 | 2.424
|  |  |  | 10 | 49 | 9 | 0.84 | 2.016 | 49 | 100 | 2.016 | 10 | 2.016 | 2.016
A | 9030 | Terbuthylazine | 66651 | 5 | 49 | 5.4 | 0.43 | 1.032 | 49 | 100 | 1.032 | 5 | 1.032 | 1.032
|  |  |  | 10 | 49 | 9.4 | 0.6 |  | 49 | 100 | 1.44 |  | 1.44
A | 9030 | Tetraconazole | 66654 | 5 | 49 | 4.9 | 1.22 | 2.928 | 49 | 100 | 2.928 | 5 | 2.928 | 2.928
|  |  |  | 10 | 49 | 9.1 | 3.47 |  | 49 | 100 | 8.328 |  | 8.328
A | 9030 | Thiobencarb | 65107 | 5 | 49 | 4.8 | 0.32 | 0.768 | 49 | 100 | 0.768 | 5 | 0.768 | 0.768
|  |  |  | 10 | 49 | 9.1 | 0.46 |  | 49 | 100 | 1.104 |  | 1.104
A | 9030 | Triallate | 68710 | 5 | 49 | 4.5 | 2.33 |  | 48 | 97.9592 |  |  | 5.592
|  |  |  | 10 | 49 | 8.7 | 3.13 | 7.512 | 49 | 100 | 7.512 | 10 | 7.512 | 7.512
A | 9030 | Tribufos | 68711 | 5 | 49 | 4.9 | 0.34 | 0.816 | 49 | 100 | 0.816 | 5 | 0.816 | 0.816
|  |  |  | 10 | 49 | 8.9 | 0.67 |  | 49 | 100 | 1.608 |  | 1.608
A | 9030 | Trifloxystrobin | 66660 | 5 | 49 | 4.6 | 0.35 | 0.84 | 49 | 100 | 0.84 | 5 | 0.84 | 0.84
|  |  |  | 10 | 49 | 8.6 | 0.53 |  | 49 | 100 | 1.272 |  | 1.272
[8 empty] | 0 |  |  |  | 0 | 0 | 0
Pesticide compounds deleted(D) or moved from ESI+ mode analytical method to ESI- mode method (lab code 9031) | [7 empty] | 0 |  |  |  | 0 | 0 | 0
D | 9030 | 3-Ketocarbofuran | 68509 | 5 | 49 | 6.3 | 3.41 |  | 47 | 95.9184 |  |  | 8.184
|  |  |  | 10 | 49 | 9.6 | 4.73 |  | 47 | 95.9184 |  |  | 11.352
D | 9030 | 7-Hydroxycarbofuran | 68518 | 5 | 49 | 96.8 | 102.68 |  | 33 | 67.3469 |  |  | 246.432
|  |  |  | 10 | 49 | 85.2 | 117.06 |  | 28 | 57.1429 |  |  | 280.944
D | 9030 | Aminopyralid | 68534 | 5 | 49 | 6.2 | 10.86 |  | 16 | 32.6531 |  |  | 26.064
|  |  |  | 10 | 49 | 12.4 | 19.02 |  | 25 | 51.0204 |  |  | 45.648
D | 9030 | Bifenazate | 68540 | 5 | 49 | 10.6 | 2 | 4.8 | 49 | 100 | 4.8 | 5 | 4.8 | 4.8
|  |  |  | 10 | 49 | 11.8 | 2.53 |  | 48 | 97.9592 | 6.072 |  | 6.072
D | 9030 | Bifenazate diazene | 68541 | 5 | 49 | 4.8 | 0.65 | 1.56 | 49 | 100 | 1.56 | 5 | 1.56 | 1.56
|  |  |  | 10 | 49 | 8.5 | 1.82 |  | 49 | 100 | 4.368 |  | 4.368
9031 | 9031 | Diflufenzopyr | 68577 | 5 | 49 | 10.9 | 17.14 |  | 21 | 42.8571 |  |  | 41.136
|  |  |  | 10 | 49 | 15 | 16.58 |  | 26 | 53.0612 |  |  | 39.792
9031 | 9031 | Diketonitrile isoxaflutole | 68578 | 5 | 49 | 69.3 | 96.69 |  | 19 | 38.7755 |  |  | 232.056
|  |  |  | 10 | 49 | 89.3 | 99.91 |  | 26 | 53.0612 |  |  | 239.784
9031 | 9031 | Dimethenamid OA | 68581 | 5 | 49 | 34.4 | 33.31 |  | 49 | 100 |  |  | 79.944
|  |  |  | 10 | 49 | 37 | 33.98 |  | 49 | 100 |  |  | 81.552
9031 | 9031 | Dimethenamid SA | 68582 | 5 | 49 | 29.3 | 26.91 |  | 36 | 73.4694 |  |  | 64.584
|  |  |  | 10 | 49 | 30.5 | 25.32 |  | 36 | 73.4694 |  |  | 60.768
9031 | 9031 | Dimethenamid SAA | 68583 | 5 | 49 | 25.2 | 20.67 |  | 38 | 77.551 |  |  | 49.608
|  |  |  | 10 | 49 | 25.2 | 22.94 |  | 36 | 73.4694 |  |  | 55.056
D | 9030 | Diquat | 68585 | 5 | 49 | 112.1 | 160.34 |  | 28 | 57.1429 |  |  | 384.816
|  |  |  | 10 | 49 | 114 | 159.7 |  | 32 | 65.3061 |  |  | 383.28
9031 | 9031 | Fipronil | 66604 | 5 | 49 | 8.4 | 12.82 |  | 31 | 63.2653 |  |  | 30.768
|  |  |  | 10 | 49 | 12.6 | 16.44 |  | 34 | 69.3878 |  |  | 39.456
9031 | 9031 | Fipronil amide | 68604 | 5 | 49 | 4.9 | 6.54 |  | 27 | 55.102 |  |  | 15.696
|  |  |  | 10 | 49 | 8.4 | 10.11 |  | 30 | 61.2245 |  |  | 24.264
9031 | 9031 | Fipronil sulfide | 66610 | 5 | 49 | 10.3 | 10.26 |  | 35 | 71.4286 |  |  | 24.624
|  |  |  | 10 | 49 | 13.3 | 14.8 |  | 35 | 71.4286 |  |  | 35.52
D | 9030 | Flumiclorac-pentyl | 68607 | 5 | 49 | 4.9 | 0.84 | 2.016 | 49 | 100 | 2.016 | 5 | 2.016 | 2.016
|  |  |  | 10 | 49 | 8.6 | 2.02 |  | 49 | 100 | 4.848 |  | 4.848
D | 9030 | Formetanate | 68609 | 5 | 49 | 4.7 | 0.55 | 1.32 | 49 | 100 | 1.32 | 5 | 1.32 | 1.32
|  |  |  | 10 | 49 | 8.5 | 1.07 |  | 49 | 100 | 2.568 |  | 2.568
D | 9030 | Kresoxim-methyl BF490-1 | 68635 | 5 | 49 | 30.8 | 40.99 |  | 25 | 51.0204 |  |  | 98.376
|  |  |  | 10 | 49 | 34.9 | 46.76 |  | 22 | 44.898 |  |  | 112.224
D | 9030 | Kresoxim-methyl BF490-2 | 68636 | 5 | 49 | 1597.2 | 3068.09 |  | 35 | 71.4286 |  |  | 7363.416
|  |  |  | 10 | 49 | 4323.1 | 5998.9 |  | 35 | 71.4286 |  |  | 14397.36
D | 9030 | Kresoxim-methyl BF490-9 | 68637 | 5 | 49 | 11.3 | 28.44 |  | 15 | 30.6122 |  |  | 68.256
|  |  |  | 10 | 49 | 34 | 110.71 |  | 16 | 32.6531 |  |  | 265.704
9031 | 9031 | Metribuzin DK | 68653 | 5 | 49 | 13.4 | 16.26 |  | 29 | 59.1837 |  |  | 39.024
|  |  |  | 10 | 49 | 15.9 | 15.6 |  | 35 | 71.4286 |  |  | 37.44
D | 9030 | Paraquat | 68667 | 5 | 49 | 38.3 | 30.18 |  | 48 | 97.9592 |  |  | 72.432
|  |  |  | 10 | 49 | 70.2 | 179.74 |  | 46 | 93.8776 |  |  | 431.376
D | 9030 | Phosmet oxon | 68674 | 5 | 49 | 6.1 | 2.99 |  | 49 | 100 |  |  | 7.176
|  |  |  | 10 | 49 | 8.2 | 4.31 | 10.344 | 48 | 97.9592 | 10.344 | 10 | 10.344 | 10.344
9031 | 9031 | Sulfentrazone | 68687 | 5 | 49 | 17.9 | 6.95 |  | 45 | 91.8367 |  |  | 16.68
|  |  |  | 10 | 49 | 23.9 | 6.37 |  | 49 | 100 |  |  | 15.288
Note: aldicarb sulfone, cis-permethrin, trans-permethrin, phosmet and prosulfuron had incorrect MRMs used in this analysis and their MDLs were detemined in table 19.
|  | 180 | [5 empty] | -20
A | 9030 | trans-Permethrin | 68708 | 5 | 49 | 40 | 68.56 | 164.544 | 22 | 44.898
|  |  |  | 10 | 49 | 56.3 | 113.44 | 272.256 | 25 | 51.0204
A | 9030 | cis-Permethrin | 68769 | 5 | 49 | 39.8 | 68.54 | 164.496 | 22 | 44.898
|  |  |  | 10 | 49 | 56.4 | 113.76 | 273.024 | 23 | 46.9388
A | 9030 | Aldicarb sulfone | 68529 | 5 | 49 | na | na | na | na | na
|  |  |  | 10 | 49 | na | na | na | na | na

## T19_esipos_mdl_high_summary
Table 19. Estimated method detection levels for pesticide compounds determined by direct aqueous-injection liquid chromatography-tandem mass spectrometry in positive electrospray ionization mode calculated according to U.S. Environmental Protection Agency procedure from analysis of eight replicate 5, 10, 25, 50, and 100 nanograms per liter standards determined in one analytical batch [Method detection level (MDL) is calculated as standard deviation times t, where t is Student's t value for eight replicates (t= 2.998); valid estimated MDLs must be greater than nominal spike concentration and have a detection frequency greater than 50 percent; MDL selected was the lowest valid MDL determined from the nominal concentration levels; pesticide compounds shown in bold have invalid MDLs using 5 and 10 ng/L concentration spike samples (see table 18). Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; USGS, U.S. Geological Survey; N, number of replicates; Std. Dev., standard deviation; RSD, relative standard deviation; min, minimum] | [12 empty] | These columns used for selecting valid MDL
[5 empty] | Concentration (ng/L) |  |  |  | Detections
Validation quality code | Pesticide compound | USGS Parameter code | Nominal concentration (ng/L) | N | Mean (ng/L) | Std Dev (ng/L) | RSD (ng/L) | MDL (ng/L) | Number of detections | Detection frequency (%) | LKK notes | Table 17 | MDL (ng/L) | Valid MDL (ng/L)
E | 1H-1,2,4-Triazole | 68498 | 10 | 8 | 10.973 | 3.7217 | 33.9171 | 11.1577 | 8 | 100 | 11.0926 | 11 | 11.1577
|  |  | 25 | 8 | 23.4105 | 3.1252 | 13.3496 | 9.3694 | 8 | 100 |  |  | 9.3694 | 9.3694 | 9.3694
|  |  | 50 | 8 | 53.7034 | 7.9868 | 14.872 |  | 8 | 100 |  |  | 23.9444 | 23.9444
|  |  | 100 | 8 | 94.9773 | 19.868 | 20.9187 |  | 8 | 100 |  |  | 59.5642 | 59.5642
E | 2-(1-hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 100 | 8 | 116.0575 | 15.6937 | 13.5223 |  | 8 | 100 | 47.0686 | 47 | 47.0497 | 47.0497
A | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 5 | 8 | 3.0286 | 0.69299 | 22.8816 | 2.0776 | 8 | 100 | 2.0986 | 2.448 | 2.0776 | 2.0776 | 2.0776
|  |  | 10 | 8 | 8.5229 | 1.8688 | 21.9264 |  | 8 | 100 |  |  | 5.6026 | 5.6026
|  |  | 25 | 8 | 24.3508 | 3.9203 | 16.0991 |  | 8 | 100 |  |  | 11.7529 | 11.7529
|  |  | 50 | 8 | 43.6082 | 3.8498 | 8.8283 |  | 8 | 100 |  |  | 11.5418 | 11.5418
|  |  | 100 | 8 | 114.3746 | 18.3599 | 16.0524 |  | 8 | 100 |  |  | 55.043 | 55.043
A | 2-Amino-N-isopropylbenzamide | 68503 | 5 | 8 | 6.8642 | 0.61222 | 8.9191 | 1.8354 | 8 | 100 | 1.7988 | 1.872 | 1.8354 | 1.8354 | 1.8354
|  |  | 10 | 8 | 10.9462 | 0.6166 | 5.633 |  | 8 | 100 |  |  | 1.8486 | 1.8486
|  |  | 25 | 8 | 26.4697 | 2.7905 | 10.5421 |  | 8 | 100 |  |  | 8.3658 | 8.3658
|  |  | 50 | 8 | 46.7296 | 4.5926 | 9.828 |  | 8 | 100 |  |  | 13.7685 | 13.7685
|  |  | 100 | 8 | 99.6502 | 11.1473 | 11.1865 |  | 8 | 100 |  |  | 33.4197 | 33.4197
A | 2-Aminobenzimidazole | 68502 | 5 | 8 | 4.4432 | 0.88935 | 20.016 | 2.6663 | 8 | 100 | 2.6982 | 4.08 | 2.6663 | 2.6663 | 2.6663
|  |  | 10 | 8 | 8.7502 | 0.93228 | 10.6545 |  | 8 | 100 |  |  | 2.795 | 2.795
|  |  | 25 | 8 | 24.3233 | 1.8179 | 7.4739 |  | 8 | 100 |  |  | 5.4501 | 5.4501
|  |  | 50 | 8 | 43.5063 | 1.3141 | 3.0204 |  | 8 | 100 |  |  | 3.9395 | 3.9395
|  |  | 100 | 8 | 91.1424 | 3.5715 | 3.9186 |  | 8 | 100 |  |  | 10.7073 | 10.7073
A | 2-Chloro-2,6-diethylacetanilide | 68525 | 5 | 8 | 3.8244 | 1.3585 | 35.523 | 4.0729 | 8 | 100 | 4.1972 | 2.256 | 4.0729 | 4.0729 | 4.0729
|  |  | 10 | 8 | 10.3003 | 2.6517 | 25.7443 |  | 8 | 100 |  |  | 7.9499 | 7.9499
|  |  | 25 | 8 | 25.2873 | 5.1916 | 20.5306 |  | 8 | 100 |  |  | 15.5645 | 15.5645
|  |  | 50 | 8 | 47.9253 | 4.6589 | 9.7211 |  | 8 | 100 |  |  | 13.9673 | 13.9673
|  |  | 100 | 8 | 113.1568 | 15.1556 | 13.3935 |  | 8 | 100 |  |  | 45.4365 | 45.4365
A | 2-Chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 5 | 8 | 2.9364 | 1.9603 | 66.7605 |  | 8 | 100 | 5.996 | 2.232 | 5.8771
|  |  | 10 | 8 | 8.082 | 1.6834 | 20.8286 | 5.0467 | 8 | 100 |  |  | 5.0467 | 5.0467 | 5.0467
|  |  | 25 | 8 | 22.8296 | 3.5202 | 15.4194 |  | 8 | 100 |  |  | 10.5535 | 10.5535
|  |  | 50 | 8 | 47.8032 | 5.7596 | 12.0486 |  | 8 | 100 |  |  | 17.2673 | 17.2673
|  |  | 100 | 8 | 112.7222 | 19.2939 | 17.1163 |  | 8 | 100 |  |  | 57.843 | 57.843
A | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 5 | 8 | 3.5858 | 0.37638 | 10.4964 | 1.1284 | 8 | 100 | 1.1992 | 1.8 | 1.1284 | 1.1284 | 1.1284
|  |  | 10 | 8 | 8.5662 | 1.2383 | 14.4553 |  | 8 | 100 |  |  | 3.7123 | 3.7123
|  |  | 25 | 8 | 25.0303 | 2.9087 | 11.6206 |  | 8 | 100 |  |  | 8.7202 | 8.7202
|  |  | 50 | 8 | 45.492 | 3.9009 | 8.5749 |  | 8 | 100 |  |  | 11.6948 | 11.6948
|  |  | 100 | 8 | 97.5528 | 13.03 | 13.3568 |  | 8 | 100 |  |  | 39.0639 | 39.0639
A | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 5 | 8 | 2.8427 | 1.1292 | 39.7229 | 3.3854 | 8 | 100 | 3.2978 | 2.808 | 3.3854 | 3.3854 | 3.3854
|  |  | 10 | 8 | 6.4985 | 1.1123 | 17.1167 |  | 8 | 100 |  |  | 3.3347 | 3.3347
|  |  | 25 | 8 | 23.8127 | 2.5895 | 10.8745 |  | 8 | 100 |  |  | 7.7634 | 7.7634
|  |  | 50 | 8 | 41.3422 | 3.5881 | 8.6791 |  | 8 | 100 |  |  | 10.7572 | 10.7572
|  |  | 100 | 8 | 98.0901 | 12.5831 | 12.8281 |  | 8 | 100 |  |  | 37.724 | 37.724
A | 2-Hydroxyatrazine | 68660 | 5 | 8 | 8.2304 | 1.1635 | 14.1362 | 3.488 | 8 | 100 | 3.5976 | 4 | 3.488 | 3.488 | 3.488
|  |  | 10 | 8 | 12.4619 | 1.5062 | 12.0863 |  | 8 | 100 |  |  | 4.5156 | 4.5156
|  |  | 25 | 8 | 27.8109 | 2.3787 | 8.5531 |  | 8 | 100 |  |  | 7.1313 | 7.1313
|  |  | 50 | 8 | 47.9357 | 4.6121 | 9.6214 |  | 8 | 100 |  |  | 13.8271 | 13.8271
|  |  | 100 | 8 | 100.2437 | 13.7229 | 13.6895 |  | 8 | 100 |  |  | 41.1411 | 41.1411
E | 3-Hydroxycarbofuran | 68508 | 5 | 8 | 2.5087 | 1.8415 | 73.4064 |  | 7 | 87.5 | 20.0866 | 7.848 | 5.5209
|  |  | 10 | 8 | 9.2204 | 4.4219 | 47.9577 |  | 8 | 100 |  |  | 13.2568
|  |  | 25 | 8 | 25.0859 | 6.7364 | 26.8535 | 20.1958 | 8 | 100 |  |  | 20.1958 | 20.1958 | 20.1958
|  |  | 50 | 8 | 46.4664 | 10.2472 | 22.053 |  | 8 | 100 |  |  | 30.7212 | 30.7212
|  |  | 100 | 8 | 101.2642 | 7.3129 | 7.2216 |  | 8 | 100 |  |  | 21.9241 | 21.9241
A | 3,4-Dichlorophenylurea | 68226 | 25 | 8 | 27.2888 | 19.3457 | 70.8925 |  | 8 | 100 | 71.952 | 72 | 57.9983
|  |  | 50 | 8 | 44.2063 | 23.7572 | 53.7417 |  | 8 | 100 |  |  | 71.2242
|  |  | 100 | 8 | 90.7627 | 23.9647 | 26.4037 | 71.8461 | 8 | 100 |  |  | 71.8461 | 71.8461 | 71.8461
A | 4-Chlorobenzylmethyl sulfoxide | 68514 | 5 | 8 | 5.2297 | 2.3728 | 45.3713 |  | 7 | 87.5 | 7.495 | 2 | 7.1135
|  |  | 10 | 8 | 9.1705 | 4.2421 | 46.2585 |  | 7 | 87.5 |  |  | 12.7179
|  |  | 25 | 8 | 27.4944 | 2.5372 | 9.2281 | 7.6065 | 8 | 100 |  |  | 7.6065 | 7.6065 | 7.6065
|  |  | 50 | 8 | 52.6718 | 3.1404 | 5.9621 |  | 8 | 100 |  |  | 9.4148 | 9.4148
|  |  | 100 | 8 | 106.8051 | 6.1576 | 5.7653 |  | 8 | 100 |  |  | 18.4604 | 18.4604
A | 4-Hydroxy molinate | 68515 | 5 | 8 | 4.7466 | 1.9144 | 40.3317 |  | 8 | 100 | 5.6962 | 3.144 | 5.7394
|  |  | 10 | 8 | 8.7068 | 2.3689 | 27.2075 | 7.102 | 8 | 100 |  |  | 7.102 | 7.102 | 7.102
|  |  | 25 | 8 | 23.9198 | 4.6339 | 19.3725 |  | 8 | 100 |  |  | 13.8924 | 13.8924
|  |  | 50 | 8 | 40.4438 | 3.8599 | 9.5439 |  | 8 | 100 |  |  | 11.572 | 11.572
|  |  | 100 | 8 | 98.4082 | 7.4744 | 7.5953 |  | 8 | 100 |  |  | 22.4081 | 22.4081
A | 4-Hydroxy-tert-fluometuron | 68619 | 25 | 8 | 13.8582 | 6.5606 | 47.3409 |  | 8 | 100 | 19.7868 | 9.936 | 19.6687 | 19.6687
|  |  | 50 | 8 | 29.6414 | 7.3957 | 24.9507 |  | 8 | 100 |  |  | 22.1724 | 22.1724
|  |  | 100 | 8 | 88.9545 | 20.7427 | 23.3183 |  | 8 | 100 |  |  | 62.1866 | 62.1866
A | 4-Hydroxyhexazinone A | 68517 | 5 | 8 | 5.3735 | 0.63691 | 11.8527 | 1.9094 | 8 | 100 | 1.7988 | 1.296 | 1.9094 | 1.9094 | 1.9094
|  |  | 10 | 8 | 10.0443 | 0.87212 | 8.6828 |  | 8 | 100 |  |  | 2.6146 | 2.6146
|  |  | 25 | 8 | 25.3944 | 2.3284 | 9.1691 |  | 8 | 100 |  |  | 6.9806 | 6.9806
|  |  | 50 | 8 | 45.2937 | 3.6262 | 8.006 |  | 8 | 100 |  |  | 10.8714 | 10.8714
|  |  | 100 | 8 | 99.3104 | 10.2138 | 10.2847 |  | 8 | 100 |  |  | 30.621 | 30.621
A | Acephate | 68519 | 5 | 8 | 4.613 | 0.59078 | 12.8068 | 1.7712 | 8 | 100 | 1.7988 | 5 | 1.7712 | 1.7712 | 1.7712
|  |  | 10 | 8 | 9.1655 | 0.64367 | 7.0227 |  | 8 | 100 |  |  | 1.9297 | 1.9297
|  |  | 25 | 8 | 25.5323 | 1.4444 | 5.6572 |  | 8 | 100 |  |  | 4.3304 | 4.3304
|  |  | 50 | 8 | 46.7329 | 1.3954 | 2.986 |  | 8 | 100 |  |  | 4.1835 | 4.1835
|  |  | 100 | 8 | 103.7869 | 3.2758 | 3.1563 |  | 8 | 100 |  |  | 9.8208 | 9.8208
A | Acetochlor | 68520 | 5 | 8 | 3.0325 | 1.6532 | 54.5151 | 4.9562 | 8 | 100 | 5.0966 | 4.728 | 4.9562 | 4.9562 | 4.9562
|  |  | 10 | 8 | 8.8916 | 1.7166 | 19.3063 |  | 8 | 100 |  |  | 5.1465 | 5.1465
|  |  | 25 | 8 | 25.6347 | 5.9358 | 23.1552 |  | 8 | 100 |  |  | 17.7954 | 17.7954
|  |  | 50 | 8 | 52.8594 | 8.8475 | 16.7378 |  | 8 | 100 |  |  | 26.5248 | 26.5248
|  |  | 100 | 8 | 114.0221 | 18.4129 | 16.1485 |  | 8 | 100 |  |  | 55.2017 | 55.2017
A | Alachlor | 65064 | 5 | 8 | 6.8357 | 1.0811 | 15.8152 | 3.2411 | 8 | 100 | 3.2978 | 4.368 | 3.2411 | 3.2411 | 3.2411
|  |  | 10 | 8 | 12.7059 | 0.8359 | 6.5788 |  | 8 | 100 |  |  | 2.506 | 2.506
|  |  | 25 | 8 | 26.2036 | 2.0172 | 7.6982 |  | 8 | 100 |  |  | 6.0476 | 6.0476
|  |  | 50 | 8 | 48.4391 | 1.6766 | 3.4613 |  | 8 | 100 |  |  | 5.0266 | 5.0266
|  |  | 100 | 8 | 101.9183 | 6.2696 | 6.1516 |  | 8 | 100 |  |  | 18.7963 | 18.7963
A | Aldicarb | 68528 | 5 | 8 | 3.889 | 1.7713 | 45.5452 | 5.3103 | 8 | 100 | 5.3964 | 3.936 | 5.3103 | 5.3103 | 5.3103
|  |  | 10 | 8 | 7.06 | 2.1781 | 30.8512 |  | 8 | 100 |  |  | 6.5299 | 6.5299
|  |  | 25 | 8 | 21.1921 | 2.8265 | 13.3373 |  | 8 | 100 |  |  | 8.4737 | 8.4737
|  |  | 50 | 8 | 37.8515 | 5.0054 | 13.2237 |  | 8 | 100 |  |  | 15.0061 | 15.0061
|  |  | 100 | 8 | 94.9489 | 8.5069 | 8.9594 |  | 8 | 100 |  |  | 25.5036 | 25.5036
A | Aldicarb sulfone | 68529 | 5 | 8 | 1.5596 | 2.234 | 143.2445 |  | 3 | 37.5 | 9.8934 | 10 | 6.6976
|  |  | 10 | 8 | 8.7569 | 3.2977 | 37.6579 | 9.8864 | 8 | 100 |  |  | 9.8864 | 9.8864 | 9.8864
|  |  | 25 | 8 | 24.4547 | 7.3197 | 29.9318 |  | 8 | 100 |  |  | 21.9446 | 21.9446
|  |  | 50 | 8 | 44.7751 | 11.3712 | 25.3963 |  | 8 | 100 |  |  | 34.0909 | 34.0909
|  |  | 100 | 8 | 111.6865 | 10.2351 | 9.1642 |  | 8 | 100 |  |  | 30.685 | 30.685
A | Aldicarb sulfoxide | 68530 | 5 | 8 | 4.3241 | 0.66526 | 15.3849 | 1.9944 | 8 | 100 | 2.0986 | 1.128 | 1.9944 | 1.9944 | 1.9944
|  |  | 10 | 8 | 9.5838 | 0.99635 | 10.3961 |  | 8 | 100 |  |  | 2.987 | 2.987
|  |  | 25 | 8 | 24.4511 | 1.3732 | 5.6161 |  | 8 | 100 |  |  | 4.1168 | 4.1168
|  |  | 50 | 8 | 44.7289 | 2.5198 | 5.6334 |  | 8 | 100 |  |  | 7.5543 | 7.5543
|  |  | 100 | 8 | 100.8629 | 5.8131 | 5.7633 |  | 8 | 100 |  |  | 17.4276 | 17.4276
A | Ametryn | 68533 | 5 | 8 | 5.0247 | 0.59392 | 11.8201 | 1.7806 | 8 | 100 | 1.7988 | 1.32 | 1.7806 | 1.7806 | 1.7806
|  |  | 10 | 8 | 9.7575 | 1.0067 | 10.3167 |  | 8 | 100 |  |  | 3.0179 | 3.0179
|  |  | 25 | 8 | 24.8894 | 2.0123 | 8.085 |  | 8 | 100 |  |  | 6.0329 | 6.0329
|  |  | 50 | 8 | 43.6656 | 3.0375 | 6.9562 |  | 8 | 100 |  |  | 9.1063 | 9.1063
|  |  | 100 | 8 | 97.6376 | 7.0095 | 7.1791 |  | 8 | 100 |  |  | 21.0145 | 21.0145
E | Asulam | 68536 | 25 | 8 | 22.2416 | 4.6534 | 20.922 |  | 8 | 100 | 14.0906 | 14 | 13.9509 | 13.9509
|  |  | 50 | 8 | 42.2419 | 8.7636 | 20.7463 |  | 8 | 100 |  |  | 26.2733 | 26.2733
|  |  | 100 | 8 | 109.4321 | 12.0173 | 10.9815 |  | 8 | 100 |  |  | 36.0277 | 36.0277
A | Atrazine | 65065 | 5 | 8 | 3.6911 | 1.6787 | 45.48 | 5.0327 | 7 | 87.5 | 5.6962 | 2.472 | 5.0327 | 5.0327 | 5.0327
|  |  | 10 | 8 | 9.2383 | 1.9158 | 20.7374 |  | 8 | 100 |  |  | 5.7435 | 5.7435
|  |  | 25 | 8 | 26.2976 | 2.6547 | 10.0948 |  | 8 | 100 |  |  | 7.9588 | 7.9588
|  |  | 50 | 8 | 41.7553 | 2.637 | 6.3154 |  | 8 | 100 |  |  | 7.9058 | 7.9058
|  |  | 100 | 8 | 98.938 | 10.4098 | 10.5216 |  | 8 | 100 |  |  | 31.2086 | 31.2086
A | Azinphos-methyl | 65066 | 5 | 8 | 4.9235 | 1.0025 | 20.3608 | 3.0054 | 8 | 100 | 2.998 | 3.624 | 3.0054 | 3.0054 | 3.0054
|  |  | 10 | 8 | 9.3755 | 1.7009 | 18.1423 |  | 8 | 100 |  |  | 5.0994 | 5.0994
|  |  | 25 | 8 | 24.206 | 2.5258 | 10.4345 |  | 8 | 100 |  |  | 7.5723 | 7.5723
|  |  | 50 | 8 | 44.0751 | 2.9592 | 6.7139 |  | 8 | 100 |  |  | 8.8716 | 8.8716
|  |  | 100 | 8 | 97.1038 | 5.1417 | 5.2951 |  | 8 | 100 |  |  | 15.415 | 15.415
A | Azinphos-methyl oxon | 68211 | 5 | 8 | 4.3533 | 1.2624 | 28.9982 | 3.7846 | 8 | 100 | 3.8974 | 7.104 | 3.7846 | 3.7846 | 3.7846
|  |  | 10 | 8 | 8.3184 | 2.7805 | 33.4253 |  | 8 | 100 |  |  | 8.3358 | 8.3358
|  |  | 25 | 8 | 22.1576 | 7.1413 | 32.2295 |  | 8 | 100 |  |  | 21.4096 | 21.4096
|  |  | 50 | 8 | 45.9726 | 7.9436 | 17.2789 |  | 8 | 100 |  |  | 23.8148 | 23.8148
|  |  | 100 | 8 | 95.7235 | 7.7442 | 8.0901 |  | 8 | 100 |  |  | 23.217 | 23.217
A | Azoxystrobin | 66589 | 5 | 8 | 4.4602 | 0.31094 | 6.9714 | 0.93221 | 8 | 100 | 0.8994 | 1.392 | 0.93221 | 0.93221 | 0.93221
|  |  | 10 | 8 | 9.0723 | 0.81073 | 8.9363 |  | 8 | 100 |  |  | 2.4306 | 2.4306
|  |  | 25 | 8 | 24.1373 | 1.4547 | 6.0269 |  | 8 | 100 |  |  | 4.3613 | 4.3613
|  |  | 50 | 8 | 44.2656 | 2.1781 | 4.9204 |  | 8 | 100 |  |  | 6.5298 | 6.5298
|  |  | 100 | 8 | 96.8609 | 3.5793 | 3.6953 |  | 8 | 100 |  |  | 10.7308 | 10.7308
E | Bifenthrin | 65067 | 5 | 8 | 4.5532 | 0.97614 | 21.4385 | 2.9265 | 8 | 100 | 2.998 | 9.072 | 2.9265 | 2.9265 | 2.9265
|  |  | 10 | 8 | 7.8143 | 1.4458 | 18.5016 |  | 8 | 100 |  |  | 4.3344 | 4.3344
|  |  | 25 | 8 | 20.9019 | 3.5564 | 17.0148 |  | 8 | 100 |  |  | 10.6621 | 10.6621
|  |  | 50 | 8 | 38.7354 | 11.514 | 29.7247 |  | 8 | 100 |  |  | 34.5188 | 34.5188
|  |  | 100 | 8 | 91.735 | 17.0028 | 18.5346 |  | 8 | 100 |  |  | 50.9743 | 50.9743
A | Bromacil | 68542 | 5 | 8 | 4.1125 | 2.1188 | 51.5203 |  | 8 | 100 | 4.497 | 2.904 | 6.352
|  |  | 10 | 8 | 7.5301 | 1.5055 | 19.993 | 4.5135 | 8 | 100 |  |  | 4.5135 | 4.5135 | 4.5135
|  |  | 25 | 8 | 27.0616 | 4.555 | 16.8319 |  | 8 | 100 |  |  | 13.6558 | 13.6558
|  |  | 50 | 8 | 43.2952 | 6.4315 | 14.8549 |  | 8 | 100 |  |  | 19.2816 | 19.2816
|  |  | 100 | 8 | 94.4803 | 10.0792 | 10.6681 |  | 8 | 100 |  |  | 30.2176 | 30.2176
E | Butralin | 68545 | 5 | 8 | 5.1106 | 0.52027 | 10.1801 | 1.5598 | 8 | 100 | 1.499 | 2.088 | 1.5598 | 1.5598 | 1.5598
|  |  | 10 | 8 | 8.3104 | 1.1254 | 13.5423 |  | 8 | 100 |  |  | 3.374 | 3.374
|  |  | 25 | 8 | 22.6119 | 1.9885 | 8.7941 |  | 8 | 100 |  |  | 5.9616 | 5.9616
|  |  | 50 | 8 | 41.5798 | 2.6732 | 6.4291 |  | 8 | 100 |  |  | 8.0143 | 8.0143
|  |  | 100 | 8 | 93.8592 | 4.2336 | 4.5106 |  | 8 | 100 |  |  | 12.6923 | 12.6923
A | Butylate | 65068 | 5 | 8 | 6.5686 | 1.1093 | 16.8879 | 3.3257 | 8 | 100 | 3.2978 | 2.736 | 3.3257 | 3.3257 | 3.3257
|  |  | 10 | 8 | 10.3044 | 4.6103 | 44.7412 |  | 7 | 87.5 |  |  | 13.8217
|  |  | 25 | 8 | 23.2168 | 2.8783 | 12.3974 | 8.6291 | 8 | 100 |  |  | 8.6291 | 8.6291 | 8.6291
|  |  | 50 | 8 | 46.8177 | 6.0389 | 12.8988 |  | 8 | 100 |  |  | 18.1046 | 18.1046
|  |  | 100 | 8 | 97.8381 | 5.5313 | 5.6535 |  | 8 | 100 |  |  | 16.5827 | 16.5827
A | Carbaryl | 65069 | 5 | 8 | 5.6362 | 1.3706 | 24.3184 | 4.1091 | 8 | 100 | 4.1972 | 3 | 4.1091 | 4.1091 | 4.1091
|  |  | 10 | 8 | 10.2328 | 2.0414 | 19.9494 |  | 8 | 100 |  |  | 6.1201 | 6.1201
|  |  | 25 | 8 | 22.7425 | 2.0842 | 9.1643 |  | 8 | 100 |  |  | 6.2484 | 6.2484
|  |  | 50 | 8 | 39.2025 | 5.6786 | 14.4853 |  | 8 | 100 |  |  | 17.0244 | 17.0244
|  |  | 100 | 8 | 81.0189 | 8.9444 | 11.0399 |  | 8 | 100 |  |  | 26.8154 | 26.8154
E | Carbendazim | 68548 | 5 | 8 | 4.4616 | 0.66811 | 14.9747 | 2.003 | 8 | 100 | 2.0986 | 1.248 | 2.003 | 2.003 | 2.003
|  |  | 10 | 8 | 9.5035 | 1.0501 | 11.0495 |  | 8 | 100 |  |  | 3.1482 | 3.1482
|  |  | 25 | 8 | 25.0084 | 1.0217 | 4.0855 |  | 8 | 100 |  |  | 3.0631 | 3.0631
|  |  | 50 | 8 | 44.8807 | 1.7883 | 3.9845 |  | 8 | 100 |  |  | 5.3613 | 5.3613
|  |  | 100 | 8 | 101.2602 | 5.2618 | 5.1964 |  | 8 | 100 |  |  | 15.775 | 15.775
A | Carbofuran | 65070 | 5 | 8 | 4.3728 | 0.58383 | 13.3512 | 1.7503 | 8 | 100 | 1.7988 | 1.632 | 1.7503 | 1.7503 | 1.7503
|  |  | 10 | 8 | 8.5328 | 1.1991 | 14.0527 |  | 8 | 100 |  |  | 3.5949 | 3.5949
|  |  | 25 | 8 | 24.5651 | 2.2337 | 9.0931 |  | 8 | 100 |  |  | 6.6967 | 6.6967
|  |  | 50 | 8 | 44.2614 | 3.5477 | 8.0153 |  | 8 | 100 |  |  | 10.6359 | 10.6359
|  |  | 100 | 8 | 100.8491 | 6.1004 | 6.0491 |  | 8 | 100 |  |  | 18.2891 | 18.2891
A | Carboxy molinate | 68549 | 10 | 8 | 11.432 | 3.634 | 31.7878 |  | 8 | 100 | 10.7928 | 11 | 10.8947 | 10.8947
|  |  | 25 | 8 | 26.7882 | 4.6064 | 17.1956 |  | 8 | 100 |  |  | 13.8099 | 13.8099
|  |  | 50 | 8 | 45.4311 | 7.9804 | 17.566 |  | 8 | 100 |  |  | 23.9253 | 23.9253
|  |  | 100 | 8 | 98.7894 | 11.4423 | 11.5825 |  | 8 | 100 |  |  | 34.304 | 34.304
A | Chlorimuron-ethyl | 68872 | 5 | 8 | 4.1218 | 1.1649 | 28.2616 | 3.4923 | 8 | 100 |  | 4 | 3.4923 | 3.4923 | 3.4923
|  |  | 10 | 8 | 8.8251 | 2.8151 | 31.8988 |  | 8 | 100 |  |  | 8.4397 | 8.4397
|  |  | 25 | 8 | 23.1558 | 4.4816 | 19.3539 |  | 8 | 100 |  |  | 13.4357 | 13.4357
|  |  | 50 | 8 | 45.0911 | 4.4055 | 9.7702 |  | 8 | 100 |  |  | 13.2076 | 13.2076
|  |  | 100 | 8 | 94.3788 | 5.8517 | 6.2003 |  | 8 | 100 |  |  | 17.5435 | 17.5435
A | Chlorpyrifos | 65072 | 5 | 8 | 4.523 | 0.70491 | 15.5851 | 2.1133 | 8 | 100 | 2.0986 | 1.512 | 2.1133 | 2.1133 | 2.1133
|  |  | 10 | 8 | 8.756 | 1.8191 | 20.7749 |  | 8 | 100 |  |  | 5.4535 | 5.4535
|  |  | 25 | 8 | 24.3779 | 1.764 | 7.2361 |  | 8 | 100 |  |  | 5.2885 | 5.2885
|  |  | 50 | 8 | 43.9722 | 4.2999 | 9.7786 |  | 8 | 100 |  |  | 12.891 | 12.891
|  |  | 100 | 8 | 97.3192 | 7.948 | 8.1669 |  | 8 | 100 |  |  | 23.828 | 23.828
A | Chlorpyrifos oxon | 68216 | 5 | 8 | 4.4178 | 0.75462 | 17.0811 | 2.2623 | 8 | 100 |  | 1.512 | 2.2623 | 2.2623 | 2.2623
|  |  | 10 | 8 | 9.0348 | 0.8313 | 9.2011 |  | 8 | 100 |  |  | 2.4922 | 2.4922
|  |  | 25 | 8 | 23.0657 | 3.3057 | 14.3317 |  | 8 | 100 |  |  | 9.9105 | 9.9105
|  |  | 50 | 8 | 44.5877 | 4.4212 | 9.9158 |  | 8 | 100 |  |  | 13.2549 | 13.2549
|  |  | 100 | 8 | 93.7141 | 3.767 | 4.0197 |  | 8 | 100 |  |  | 11.2936 | 11.2936
A | Chlorsulfuron | 61678 | 5 | 8 | 0 | 0 | 0 |  | 0 | 0 | 28.481 | 25 | 0
|  |  | 10 | 8 | 8.8733 | 5.7993 | 65.3571 |  | 8 | 100 |  |  | 17.3864
|  |  | 25 | 8 | 23.9379 | 9.5374 | 39.8423 |  | 8 | 100 |  |  | 28.5932
|  |  | 50 | 8 | 40.9866 | 8.5674 | 20.9031 | 25.6852 | 8 | 100 |  |  | 25.6852 | 25.6852 | 25.6852
|  |  | 100 | 8 | 103.7018 | 15.9532 | 15.3838 |  | 8 | 100 |  |  | 47.8278 | 47.8278
A | cis-Permethrin | 68769 | 5 | 8 | 4.5126 | 0.50588 | 11.2103 | 1.5166 | 8 | 100 | 1.499 | 1 | 1.5166 | 1.5166 | 1.5166
|  |  | 10 | 8 | 8.8039 | 0.53611 | 6.0895 |  | 8 | 100 |  |  | 1.6073 | 1.6073
|  |  | 25 | 8 | 24.2003 | 0.87175 | 3.6023 |  | 8 | 100 |  |  | 2.6135 | 2.6135
|  |  | 50 | 8 | 45.1406 | 1.4332 | 3.1749 |  | 8 | 100 |  |  | 4.2966 | 4.2966
|  |  | 100 | 8 | 95.4526 | 1.6059 | 1.6824 |  | 8 | 100 |  |  | 4.8145 | 4.8145
A | Cyanazine | 66592 | 25 | 8 | 30.4758 | 19.2644 | 63.2122 |  | 7 | 87.5 | 46.7688 | 7.728 | 57.7548
|  |  | 50 | 8 | 51.6491 | 15.613 | 30.229 | 46.8078 | 8 | 100 |  |  | 46.8078 | 46.8078 | 46.8078
|  |  | 100 | 8 | 97.2255 | 29.3071 | 30.1435 |  | 8 | 100 |  |  | 87.8628 | 87.8628
A | Dechlorometolachlor | 68562 | 5 | 8 | 4.4444 | 0.42559 | 9.5758 | 1.2759 | 8 | 100 | 1.1992 | 0.552 | 1.2759 | 1.2759 | 1.2759
|  |  | 10 | 8 | 8.7192 | 0.50311 | 5.7702 |  | 8 | 100 |  |  | 1.5083 | 1.5083
|  |  | 25 | 8 | 22.4252 | 1.1671 | 5.2046 |  | 8 | 100 |  |  | 3.4991 | 3.4991
|  |  | 50 | 8 | 42.3169 | 2.0085 | 4.7463 |  | 8 | 100 |  |  | 6.0214 | 6.0214
|  |  | 100 | 8 | 96.2761 | 5.0409 | 5.2359 |  | 8 | 100 |  |  | 15.1126 | 15.1126
A | Deethylatrazine | 68552 | 5 | 8 | 4.6501 | 1.1429 | 24.578 | 3.4265 | 8 | 100 | 3.2978 | 5.208 | 3.4265 | 3.4265 | 3.4265
|  |  | 10 | 8 | 8.074 | 3.0211 | 37.4176 |  | 8 | 100 |  |  | 9.0573 | 9.0573
|  |  | 25 | 8 | 26.8052 | 4.9259 | 18.3765 |  | 8 | 100 |  |  | 14.7677 | 14.7677
|  |  | 50 | 8 | 38.7058 | 3.2589 | 8.4196 |  | 8 | 100 |  |  | 9.7701 | 9.7701
|  |  | 100 | 8 | 97.9702 | 10.5244 | 10.7425 |  | 8 | 100 |  |  | 31.5523 | 31.5523
A | Desiodo flubendiamide | 68563 | 25 | 8 | 22.465 | 2.9624 | 13.1866 |  | 8 | 100 | 8.994 | 1.848 | 8.8812 | 8.8812
|  |  | 50 | 8 | 44.9419 | 5.7171 | 12.7212 |  | 8 | 100 |  |  | 17.14 | 17.14
|  |  | 100 | 8 | 109.1392 | 19.2193 | 17.6099 |  | 8 | 100 |  |  | 57.6194 | 57.6194
A | Deisopropyl prometryn | 68564 | 5 | 8 | 4.3786 | 0.54878 | 12.5332 | 1.6452 | 8 | 100 | 1.499 | 1.248 | 1.6452 | 1.6452 | 1.6452
|  |  | 10 | 8 | 8.8217 | 0.4969 | 5.6326 |  | 8 | 100 |  |  | 1.4897 | 1.4897
|  |  | 25 | 8 | 25.5153 | 0.97858 | 3.8353 |  | 8 | 100 |  |  | 2.9338 | 2.9338
|  |  | 50 | 8 | 46.3171 | 3.0487 | 6.5823 |  | 8 | 100 |  |  | 9.1401 | 9.1401
|  |  | 100 | 8 | 102.2705 | 7.3382 | 7.1753 |  | 8 | 100 |  |  | 21.9999 | 21.9999
A | Deisopropylatrazine | 68550 | 5 | 8 | 3.9855 | 1.7034 | 42.7388 | 5.1067 | 8 | 100 | 5.0966 | 10 | 5.1067 | 5.1067 | 5.1067
|  |  | 10 | 8 | 11.2963 | 3.1006 | 27.4479 |  | 8 | 100 |  |  | 9.2956 | 9.2956
|  |  | 25 | 8 | 27.3114 | 6.0219 | 22.049 |  | 8 | 100 |  |  | 18.0536 | 18.0536
|  |  | 50 | 8 | 49.0979 | 9.33 | 19.0028 |  | 8 | 100 |  |  | 27.9713 | 27.9713
|  |  | 100 | 8 | 104.7607 | 13.1017 | 12.5063 |  | 8 | 100 |  |  | 39.2789 | 39.2789
A | Demethyl fluometuron | 68591 | 5 | 8 | 4.9317 | 1.4073 | 28.5364 | 4.2192 | 8 | 100 | 4.1972 | 0.816 | 4.2192 | 4.2192 | 4.2192
|  |  | 10 | 8 | 10.3252 | 1.0867 | 10.5249 |  | 8 | 100 |  |  | 3.258 | 3.258
|  |  | 25 | 8 | 24.6509 | 1.7232 | 6.9903 |  | 8 | 100 |  |  | 5.1661 | 5.1661
|  |  | 50 | 8 | 48.504 | 3.2339 | 6.6672 |  | 8 | 100 |  |  | 9.6951 | 9.6951
|  |  | 100 | 8 | 104.6063 | 4.5791 | 4.3775 |  | 8 | 100 |  |  | 13.7282 | 13.7282
A | Demethyl hexazinone B | 68566 | 5 | 8 | 4.3251 | 0.59993 | 13.8708 | 1.7986 | 8 | 100 | 1.7988 | 1.224 | 1.7986 | 1.7986 | 1.7986
|  |  | 10 | 8 | 8.8449 | 3.6352 | 41.0994 |  | 7 | 87.5 |  |  | 10.8983 | 10.8983
|  |  | 25 | 8 | 26.6209 | 2.5148 | 9.4468 |  | 8 | 100 |  |  | 7.5394 | 7.5394
|  |  | 50 | 8 | 48.9847 | 3.5351 | 7.2167 |  | 8 | 100 |  |  | 10.5982 | 10.5982
|  |  | 100 | 8 | 106.5693 | 8.1788 | 7.6747 |  | 8 | 100 |  |  | 24.5201 | 24.5201
A | Demethyl norflurazon | 68567 | 5 | 8 | 4.756 | 1.9411 | 40.8138 |  | 7 | 87.5 | 10.493 | 1.944 | 5.8194
|  |  | 10 | 8 | 9.4 | 0.28008 | 2.9796 | 0.83969 | 8 | 100 |  |  | 0.83969 | 0.83969 | 0.83969
|  |  | 25 | 8 | 24.0612 | 1.2102 | 5.0297 |  | 8 | 100 |  |  | 3.6282 | 3.6282
|  |  | 50 | 8 | 44.3627 | 3.533 | 7.964 |  | 8 | 100 |  |  | 10.592 | 10.592
|  |  | 100 | 8 | 95.1228 | 8.3867 | 8.8167 |  | 8 | 100 |  |  | 25.1432 | 25.1432
A | Diazinon | 65078 | 5 | 8 | 4.7376 | 1.0135 | 21.3932 | 3.0385 | 8 | 100 | 2.998 | 1.2 | 3.0385 | 3.0385 | 3.0385
|  |  | 10 | 8 | 9.3391 | 0.80977 | 8.6707 |  | 8 | 100 |  |  | 2.4277 | 2.4277
|  |  | 25 | 8 | 23.9003 | 2.0716 | 8.6675 |  | 8 | 100 |  |  | 6.2106 | 6.2106
|  |  | 50 | 8 | 43.9613 | 3.4653 | 7.8827 |  | 8 | 100 |  |  | 10.389 | 10.389
|  |  | 100 | 8 | 98.8324 | 5.9923 | 6.0631 |  | 8 | 100 |  |  | 17.965 | 17.965
E | Diazinon oxon | 68236 | 5 | 8 | 4.5577 | 0.56325 | 12.3582 | 1.6886 | 8 | 100 | 1.7988 | 2 | 1.6886 | 1.6886 | 1.6886
|  |  | 10 | 8 | 8.3896 | 0.96668 | 11.5224 |  | 8 | 100 |  |  | 2.8981 | 2.8981
|  |  | 25 | 8 | 22.5256 | 2.0028 | 8.8914 |  | 8 | 100 |  |  | 6.0045 | 6.0045
|  |  | 50 | 8 | 41.3922 | 1.6604 | 4.0113 |  | 8 | 100 |  |  | 4.9778 | 4.9778
|  |  | 100 | 8 | 92.3629 | 4.8741 | 5.2771 |  | 8 | 100 |  |  | 14.6125 | 14.6125
A | Dichlorvos | 68572 | 10 | 8 | 7.1328 | 4.0243 | 56.4205 |  | 7 | 87.5 | 25.7828 | 26 | 12.065
|  |  | 25 | 8 | 26.966 | 8.5921 | 31.8628 | 25.7592 | 8 | 100 |  |  | 25.7592 | 25.7592 | 25.7592
|  |  | 50 | 8 | 50.3623 | 11.0755 | 21.9917 |  | 8 | 100 |  |  | 33.2045 | 33.2045
|  |  | 100 | 8 | 108.0744 | 11.2098 | 10.3723 |  | 8 | 100 |  |  | 33.607 | 33.607
A | Dicrotophos | 68573 | 5 | 8 | 4.4877 | 0.67817 | 15.1117 | 2.0332 | 8 | 100 | 2.0986 | 1.896 | 2.0332 | 2.0332 | 2.0332
|  |  | 10 | 8 | 8.6975 | 1.1622 | 13.3622 |  | 8 | 100 |  |  | 3.4842 | 3.4842
|  |  | 25 | 8 | 24.4597 | 1.2565 | 5.1371 |  | 8 | 100 |  |  | 3.767 | 3.767
|  |  | 50 | 8 | 43.9582 | 2.4172 | 5.4988 |  | 8 | 100 |  |  | 7.2466 | 7.2466
|  |  | 100 | 8 | 97.3592 | 5.7716 | 5.9281 |  | 8 | 100 |  |  | 17.3031 | 17.3031
A | Didealkylatrazine | 68547 | 10 | 8 | 7.0913 | 4.0427 | 57.0094 |  | 8 | 100 | 11.992 | 12 | 12.12
|  |  | 25 | 8 | 24.3542 | 6.4775 | 26.5972 | 19.4197 | 8 | 100 |  |  | 19.4197 | 19.4197 | 19.4197
|  |  | 50 | 8 | 47.9266 | 9.8713 | 20.5966 |  | 8 | 100 |  |  | 29.5941 | 29.5941
|  |  | 100 | 8 | 108.5337 | 14.9574 | 13.7814 |  | 8 | 100 |  |  | 44.8424 | 44.8424
A | Diflubenzuron | 68576 | 5 | 8 | 3.5954 | 1.285 | 35.739 | 3.8523 | 8 | 100 | 3.8974 | 2.568 | 3.8523 | 3.8523 | 3.8523
|  |  | 10 | 8 | 7.6001 | 2.0744 | 27.295 |  | 8 | 100 |  |  | 6.2192 | 6.2192
|  |  | 25 | 8 | 21.3699 | 3.4882 | 16.3231 |  | 8 | 100 |  |  | 10.4577 | 10.4577
|  |  | 50 | 8 | 42.4973 | 2.5153 | 5.9188 |  | 8 | 100 |  |  | 7.541 | 7.541
|  |  | 100 | 8 | 86.2731 | 5.224 | 6.0552 |  | 8 | 100 |  |  | 15.6615 | 15.6615
A | Dimethenamid | 68580 | 5 | 8 | 3.6117 | 0.5189 | 14.3671 | 1.5557 | 8 | 100 | 1.499 | 1.224 | 1.5557 | 1.5557 | 1.5557
|  |  | 10 | 8 | 7.5144 | 1.2677 | 16.8698 |  | 8 | 100 |  |  | 3.8005 | 3.8005
|  |  | 25 | 8 | 23.7614 | 3.5817 | 15.0735 |  | 8 | 100 |  |  | 10.7379 | 10.7379
|  |  | 50 | 8 | 46.4875 | 6.3589 | 13.6788 |  | 8 | 100 |  |  | 19.0641 | 19.0641
|  |  | 100 | 8 | 113.2355 | 15.5413 | 13.7248 |  | 8 | 100 |  |  | 46.5928 | 46.5928
A | Dimethoate | 66596 | 5 | 8 | 4.3679 | 0.39563 | 9.0577 | 1.1861 | 8 | 100 | 1.1992 | 1.44 | 1.1861 | 1.1861 | 1.1861
|  |  | 10 | 8 | 8.8527 | 0.47288 | 5.3417 |  | 8 | 100 |  |  | 1.4177 | 1.4177
|  |  | 25 | 8 | 24.8565 | 1.6433 | 6.6113 |  | 8 | 100 |  |  | 4.9267 | 4.9267
|  |  | 50 | 8 | 43.3084 | 4.0484 | 9.3479 |  | 8 | 100 |  |  | 12.1371 | 12.1371
|  |  | 100 | 8 | 102.545 | 2.8272 | 2.757 |  | 8 | 100 |  |  | 8.476 | 8.476
A | Dimethoate oxon | 68661 | 5 | 8 | 5.0903 | 0.32776 | 6.439 | 0.98262 | 8 | 100 | 0.8994 | 0.864 | 0.98262 | 0.98262 | 0.98262
|  |  | 10 | 8 | 9.8992 | 0.64585 | 6.5242 |  | 8 | 100 |  |  | 1.9362 | 1.9362
|  |  | 25 | 8 | 24.9718 | 0.98927 | 3.9615 |  | 8 | 100 |  |  | 2.9658 | 2.9658
|  |  | 50 | 8 | 45.809 | 2.0879 | 4.5579 |  | 8 | 100 |  |  | 6.2596 | 6.2596
|  |  | 100 | 8 | 97.4484 | 6.6452 | 6.8192 |  | 8 | 100 |  |  | 19.9223 | 19.9223
A | Disulfoton | 67595 | 5 | 8 | 4.3016 | 0.86049 | 20.0037 | 2.5797 | 8 | 100 | 2.6982 | 6.24 | 2.5797 | 2.5797 | 2.5797
|  |  | 10 | 8 | 7.9425 | 1.5657 | 19.7124 |  | 8 | 100 |  |  | 4.6938 | 4.6938
|  |  | 25 | 8 | 21.0982 | 3.0108 | 14.2705 |  | 8 | 100 |  |  | 9.0264 | 9.0264
|  |  | 50 | 8 | 39.3837 | 5.8715 | 14.9084 |  | 8 | 100 |  |  | 17.6027 | 17.6027
|  |  | 100 | 8 | 87.7508 | 7.3003 | 8.3193 |  | 8 | 100 |  |  | 21.8862 | 21.8862
A | Disulfoton oxon | 68586 | 5 | 8 | 4.5956 | 0.38657 | 8.4116 | 1.1589 | 8 | 100 | 1.1992 | 1.224 | 1.1589 | 1.1589 | 1.1589
|  |  | 10 | 8 | 8.6758 | 0.76109 | 8.7726 |  | 8 | 100 |  |  | 2.2817 | 2.2817
|  |  | 25 | 8 | 24.536 | 0.7173 | 2.9234 |  | 8 | 100 |  |  | 2.1505 | 2.1505
|  |  | 50 | 8 | 45.8295 | 2.2033 | 4.8075 |  | 8 | 100 |  |  | 6.6053 | 6.6053
|  |  | 100 | 8 | 99.779 | 4.0358 | 4.0447 |  | 8 | 100 |  |  | 12.0993 | 12.0993
A | Disulfoton oxon sulfone | 68588 | 5 | 8 | 4.1489 | 1.4374 | 34.6443 | 4.3092 | 8 | 100 | 4.1972 | 3.312 | 4.3092 | 4.3092 | 4.3092
|  |  | 10 | 8 | 9.2855 | 2.046 | 22.0348 |  | 8 | 100 |  |  | 6.134 | 6.134
|  |  | 25 | 8 | 27.1486 | 1.6137 | 5.9439 |  | 8 | 100 |  |  | 4.8379 | 4.8379
|  |  | 50 | 8 | 43.4415 | 4.2333 | 9.7447 |  | 8 | 100 |  |  | 12.6913 | 12.6913
|  |  | 100 | 8 | 101.1107 | 5.7738 | 5.7104 |  | 8 | 100 |  |  | 17.3098 | 17.3098
A | Disulfoton oxon sulfoxide | 68587 | 5 | 8 | 3.4853 | 0.60386 | 17.326 | 1.8104 | 8 | 100 | 1.7988 | 3.144 | 1.8104 | 1.8104 | 1.8104
|  |  | 10 | 8 | 8.3376 | 1.4006 | 16.7983 |  | 8 | 100 |  |  | 4.1989 | 4.1989
|  |  | 25 | 8 | 25.6153 | 2.6558 | 10.368 |  | 8 | 100 |  |  | 7.9621 | 7.9621
|  |  | 50 | 8 | 43.0002 | 3.2564 | 7.5729 |  | 8 | 100 |  |  | 9.7625 | 9.7625
|  |  | 100 | 8 | 101.2074 | 5.2577 | 5.1949 |  | 8 | 100 |  |  | 15.7624 | 15.7624
A | Disulfoton sulfone | 68589 | 10 | 8 | 4.8256 | 2.4865 | 51.5277 |  | 8 | 100 | 7.495 | 4.272 | 7.4545 | 7.4545
|  |  | 25 | 8 | 25.4607 | 7.8863 | 30.9743 |  | 8 | 100 |  |  | 23.6431 | 23.6431
|  |  | 50 | 8 | 38.1223 | 10.8781 | 28.5347 |  | 8 | 100 |  |  | 32.6125 | 32.6125
|  |  | 100 | 8 | 99.5149 | 18.2536 | 18.3426 |  | 8 | 100 |  |  | 54.7242 | 54.7242
A | Disulfoton sulfoxide | 68590 | 5 | 8 | 4.2273 | 0.92062 | 21.7782 | 2.76 | 8 | 100 | 2.6982 | 1.512 | 2.76 | 2.76 | 2.76
|  |  | 10 | 8 | 9.6471 | 0.92827 | 9.6223 |  | 8 | 100 |  |  | 2.7829 | 2.7829
|  |  | 25 | 8 | 24.9489 | 1.9958 | 7.9995 |  | 8 | 100 |  |  | 5.9834 | 5.9834
|  |  | 50 | 8 | 46.3914 | 4.0789 | 8.7923 |  | 8 | 100 |  |  | 12.2284 | 12.2284
|  |  | 100 | 8 | 103.4498 | 5.2283 | 5.054 |  | 8 | 100 |  |  | 15.6745 | 15.6745
A | Diuron | 66598 | 5 | 8 | 5.9841 | 1.1055 | 18.4745 | 3.3144 | 8 | 100 | 3.2978 | 2.304 | 3.3144 | 3.3144 | 3.3144
|  |  | 10 | 8 | 11.7043 | 1.3931 | 11.9028 |  | 8 | 100 |  |  | 4.1766 | 4.1766
|  |  | 25 | 8 | 26.2266 | 2.4172 | 9.2167 |  | 8 | 100 |  |  | 7.2468 | 7.2468
|  |  | 50 | 8 | 46.2646 | 1.4283 | 3.0873 |  | 8 | 100 |  |  | 4.2821 | 4.2821
|  |  | 100 | 8 | 103.8257 | 8.7416 | 8.4195 |  | 8 | 100 |  |  | 26.2072 | 26.2072
A | EPTC | 65080 | 50 | 8 | 44.0207 | 26.6594 | 60.561 |  | 8 | 100 | 102.8314 | 103 | 79.9249
|  |  | 100 | 8 | 113.6786 | 34.2663 | 30.1431 | 102.7302 | 8 | 100 |  |  | 102.7302 | 102.7302 | 102.7302
A | EPTC R248722 | 68594 | 5 | 8 | 4.4575 | 1.9026 | 42.6831 |  | 7 | 87.5 | 2.3984 | 1.704 | 5.704
|  |  | 10 | 8 | 8.7425 | 0.79803 | 9.1281 | 2.3925 | 8 | 100 |  |  | 2.3925 | 2.3925 | 2.3925
|  |  | 25 | 8 | 24.7246 | 2.9808 | 12.0559 |  | 8 | 100 |  |  | 8.9363 | 8.9363
|  |  | 50 | 8 | 43.6682 | 4.6469 | 10.6414 |  | 8 | 100 |  |  | 13.9315 | 13.9315
|  |  | 100 | 8 | 101.2126 | 5.3416 | 5.2776 |  | 8 | 100 |  |  | 16.014 | 16.014
A | Ethoprop | 68596 | 5 | 8 | 4.6067 | 0.81464 | 17.6838 | 2.4423 | 8 | 100 | 2.3984 | 2.064 | 2.4423 | 2.4423 | 2.4423
|  |  | 10 | 8 | 9.1751 | 1.0712 | 11.675 |  | 8 | 100 |  |  | 3.2115 | 3.2115
|  |  | 25 | 8 | 23.4114 | 2.07 | 8.8419 |  | 8 | 100 |  |  | 6.2059 | 6.2059
|  |  | 50 | 8 | 46.6494 | 5.3841 | 11.5415 |  | 8 | 100 |  |  | 16.1414 | 16.1414
|  |  | 100 | 8 | 100.0719 | 5.6866 | 5.6825 |  | 8 | 100 |  |  | 17.0485 | 17.0485
A | Etoxazole | 68598 | 5 | 8 | 4.178 | 0.22687 | 5.4301 | 0.68015 | 8 | 100 | 1 | 0.768 | 0.68015 | 0.68015 | 0.68015
|  |  | 10 | 8 | 8.0919 | 0.35152 | 4.3441 |  | 8 | 100 |  |  | 1.0538 | 1.0538
|  |  | 25 | 8 | 22.9732 | 0.76333 | 3.3227 |  | 8 | 100 |  |  | 2.2885 | 2.2885
|  |  | 50 | 8 | 42.2417 | 1.6422 | 3.8876 |  | 8 | 100 |  |  | 4.9233 | 4.9233
|  |  | 100 | 8 | 96.2831 | 2.9626 | 3.077 |  | 8 | 100 |  |  | 8.8819 | 8.8819
A | Fenamiphos | 68599 | 5 | 8 | 4.4294 | 0.47981 | 10.8324 | 1.4385 | 8 | 100 | 1.499 | 0.984 | 1.4385 | 1.4385 | 1.4385
|  |  | 10 | 8 | 9.2335 | 0.65676 | 7.1128 |  | 8 | 100 |  |  | 1.969 | 1.969
|  |  | 25 | 8 | 24.8107 | 1.1637 | 4.6901 |  | 8 | 100 |  |  | 3.4886 | 3.4886
|  |  | 50 | 8 | 44.5802 | 2.948 | 6.6127 |  | 8 | 100 |  |  | 8.838 | 8.838
|  |  | 100 | 8 | 99.2297 | 5.665 | 5.709 |  | 8 | 100 |  |  | 16.9837 | 16.9837
A | Fenamiphos sulfone | 68600 | 5 | 8 | 4.2841 | 1.0029 | 23.4103 | 3.0067 | 8 | 100 | 2.998 | 2.064 | 3.0067 | 3.0067 | 3.0067
|  |  | 10 | 8 | 8.8719 | 2.3155 | 26.0989 |  | 8 | 100 |  |  | 6.9418 | 6.9418
|  |  | 25 | 8 | 23.6298 | 2.8074 | 11.881 |  | 8 | 100 |  |  | 8.4167 | 8.4167
|  |  | 50 | 8 | 45.4573 | 4.3207 | 9.505 |  | 8 | 100 |  |  | 12.9534 | 12.9534
|  |  | 100 | 8 | 96.955 | 6.8946 | 7.1111 |  | 8 | 100 |  |  | 20.67 | 20.67
A | Fenamiphos sulfoxide | 68601 | 5 | 8 | 4.0393 | 0.90952 | 22.5168 | 2.7267 | 8 | 100 | 2.6982 | 2.472 | 2.7267 | 2.7267 | 2.7267
|  |  | 10 | 8 | 8.2319 | 2.688 | 32.6541 |  | 8 | 100 |  |  | 8.0588 | 8.0588
|  |  | 25 | 8 | 25.1336 | 1.9322 | 7.6877 |  | 8 | 100 |  |  | 5.7928 | 5.7928
|  |  | 50 | 8 | 41.6479 | 3.5591 | 8.5457 |  | 8 | 100 |  |  | 10.6702 | 10.6702
|  |  | 100 | 8 | 98.5797 | 5.8392 | 5.9233 |  | 8 | 100 |  |  | 17.5058 | 17.5058
E | Fenbutatin oxide | 68602 | 5 | 8 | 5.5318 | 5.268 | 95.2305 |  | 8 | 100 | 10.493 | 10 | 15.7934
|  |  | 10 | 8 | 7.0362 | 3.4864 | 49.5492 | 10.4522 | 8 | 100 |  |  | 10.4522 | 10.4522 | 10.4522
|  |  | 25 | 8 | 22.5557 | 3.1751 | 14.0767 |  | 8 | 100 |  |  | 9.519 | 9.519
|  |  | 50 | 8 | 43.1585 | 6.2167 | 14.4044 |  | 8 | 100 |  |  | 18.6377 | 18.6377
|  |  | 100 | 8 | 82.2629 | 12.6216 | 15.343 |  | 8 | 100 |  |  | 37.8396 | 37.8396
A | Fentin | 68603 | 5 | 8 | 2.8836 | 1.8681 | 64.783 |  | 8 | 100 | 5.6962 | 7.848 | 5.6006
|  |  | 10 | 8 | 7.6204 | 2.7316 | 35.8462 | 8.1894 | 8 | 100 |  |  | 8.1894 | 8.1894 | 8.1894
|  |  | 25 | 8 | 23.2945 | 2.9247 | 12.5555 |  | 8 | 100 |  |  | 8.7684 | 8.7684
|  |  | 50 | 8 | 42.0199 | 6.5581 | 15.607 |  | 8 | 100 |  |  | 19.6611 | 19.6611
|  |  | 100 | 8 | 91.4251 | 8.7287 | 9.5474 |  | 8 | 100 |  |  | 26.1687 | 26.1687
E | Flumetsulam | 61679 | 10 | 8 | 7.4183 | 3.937 | 53.0713 |  | 8 | 100 | 11.6922 | 8.112 | 11.803
|  |  | 25 | 8 | 23.1605 | 4.0437 | 17.4593 | 12.1229 | 8 | 100 |  |  | 12.1229 | 12.1229 | 12.1229
|  |  | 50 | 8 | 40.7478 | 12.7736 | 31.3478 |  | 8 | 100 |  |  | 38.2951 | 38.2951
|  |  | 100 | 8 | 102.7816 | 13.6476 | 13.2783 |  | 8 | 100 |  |  | 40.9155 | 40.9155
D | flumiclorac-pentyl | 68607 | 5 | 8 | 4.1731 | 0.90636 | 21.7192 | 2.7173 | 8 | 100 | 2.6982 |  | 2.7173 | 2.7173 | 2.7173
|  |  | 10 | 8 | 8.411 | 1.34 | 15.9313 |  | 8 | 100 |  |  | 4.0173 | 4.0173
|  |  | 25 | 8 | 18.8835 | 4.7079 | 24.9311 |  | 8 | 100 |  |  | 14.1142 | 14.1142
|  |  | 50 | 8 | 36.7739 | 2.6665 | 7.251 |  | 8 | 100 |  |  | 7.9941 | 7.9941
|  |  | 100 | 8 | 67.2097 | 23.7083 | 35.2752 |  | 8 | 100 |  |  | 71.0776 | 71.0776
A | Fluometuron | 68608 | 5 | 8 | 4.0523 | 3.189 | 78.6961 |  | 8 | 100 | 7.1952 | 1.368 | 9.5607
|  |  | 10 | 8 | 15.1542 | 2.3828 | 15.7238 | 7.1437 | 8 | 100 |  |  | 7.1437 | 7.1437 | 7.1437
|  |  | 25 | 8 | 23.5703 | 2.9123 | 12.3558 |  | 8 | 100 |  |  | 8.7311 | 8.7311
|  |  | 50 | 8 | 49.5553 | 9.1572 | 18.4788 |  | 8 | 100 |  |  | 27.4534 | 27.4534
|  |  | 100 | 8 | 111.7071 | 13.623 | 12.1953 |  | 8 | 100 |  |  | 40.8418 | 40.8418
A | Fonofos | 65084 | 5 | 8 | 4.884 | 1.6 | 32.7611 | 4.7969 | 8 | 100 | 4.7968 | 5.208 | 4.7969 | 4.7969 | 4.7969
|  |  | 10 | 8 | 8.9516 | 2.8656 | 32.0119 |  | 8 | 100 |  |  | 8.591 | 8.591
|  |  | 25 | 8 | 24.2588 | 3.0107 | 12.4108 |  | 8 | 100 |  |  | 9.0261 | 9.0261
|  |  | 50 | 8 | 48.3854 | 5.0287 | 10.393 |  | 8 | 100 |  |  | 15.0761 | 15.0761
|  |  | 100 | 8 | 104.1001 | 7.3451 | 7.0558 |  | 8 | 100 |  |  | 22.0205 | 22.0205
A | Halosulfuron-methyl | 61680 | 5 | 8 | 4.2573 | 2.4483 | 57.5085 |  | 7 | 87.5 | 4.7968 | 10.704 | 7.34
|  |  | 10 | 8 | 9.8467 | 1.6117 | 16.368 | 4.8319 | 8 | 100 |  |  | 4.8319 | 4.8319 | 4.8319
|  |  | 25 | 8 | 22.6206 | 3.5018 | 15.4807 |  | 8 | 100 |  |  | 10.4984 | 10.4984
|  |  | 50 | 8 | 44.2177 | 3.5933 | 8.1263 |  | 8 | 100 |  |  | 10.7726 | 10.7726
|  |  | 100 | 8 | 95.0431 | 9.4039 | 9.8943 |  | 8 | 100 |  |  | 28.1929 | 28.1929
A | Hexazinone | 65085 | 5 | 8 | 4.3471 | 0.33362 | 7.6746 | 1.0002 | 8 | 100 | 0.8994 | 1.32 | 1.0002 | 1.0002 | 1.0002
|  |  | 10 | 8 | 10.1051 | 0.90724 | 8.978 |  | 8 | 100 |  |  | 2.7199 | 2.7199
|  |  | 25 | 8 | 26.7961 | 1.896 | 7.0756 |  | 8 | 100 |  |  | 5.6842 | 5.6842
|  |  | 50 | 8 | 47.2362 | 3.6286 | 7.6819 |  | 8 | 100 |  |  | 10.8786 | 10.8786
|  |  | 100 | 8 | 106.3863 | 6.4 | 6.0158 |  | 8 | 100 |  |  | 19.1871 | 19.1871
A | Hexazinone TP C | 68612 | 5 | 8 | 4.7287 | 0.49711 | 10.5125 | 1.4903 | 8 | 100 | 1.499 | 0.912 | 1.4903 | 1.4903 | 1.4903
|  |  | 10 | 8 | 9.1717 | 0.4393 | 4.7898 |  | 8 | 100 |  |  | 1.317 | 1.317
|  |  | 25 | 8 | 25.1743 | 2.4397 | 9.691 |  | 8 | 100 |  |  | 7.3141 | 7.3141
|  |  | 50 | 8 | 44.2137 | 3.3163 | 7.5006 |  | 8 | 100 |  |  | 9.9422 | 9.9422
|  |  | 100 | 8 | 97.3149 | 11.6887 | 12.0112 |  | 8 | 100 |  |  | 35.0428 | 35.0428
A | Hexazinone TP F | 68574 | 10 | 8 | 11.2576 | 4.9079 | 43.5967 |  | 7 | 87.5 | 19.7868 | 4.92 | 14.714
|  |  | 25 | 8 | 20.791 | 9.1845 | 44.1753 |  | 7 | 87.5 |  |  | 27.5351
|  |  | 50 | 8 | 47.0037 | 6.6153 | 14.0739 | 19.8326 | 8 | 100 |  |  | 19.8326 | 19.8326 | 19.8326
|  |  | 100 | 8 | 96.5754 | 6.1778 | 6.3969 |  | 8 | 100 |  |  | 18.5211 | 18.5211
A | Hexazinone TP G | 68713 | 10 | 8 | 8.9449 | 2.5634 | 28.6572 |  | 8 | 100 | 7.7948 | 10.776 | 7.6849 | 7.6849
|  |  | 25 | 8 | 23.4141 | 8.8977 | 38.0013 |  | 8 | 100 |  |  | 26.6753 | 26.6753
|  |  | 50 | 8 | 43.5446 | 13.6236 | 31.2866 |  | 8 | 100 |  |  | 40.8436 | 40.8436
|  |  | 100 | 8 | 99.2486 | 11.7998 | 11.8892 |  | 8 | 100 |  |  | 35.3759 | 35.3759
A | Hydroxy mono demethyl fluometuron | 68617 | 5 | 8 | 5.2014 | 1.9091 | 36.703 |  | 8 | 100 | 5.6962 | 4.968 | 5.7234
|  |  | 10 | 8 | 11.67 | 3.3065 | 28.3336 | 9.913 | 8 | 100 |  |  | 9.913 | 9.913 | 9.913
|  |  | 25 | 8 | 22.5786 | 4.4857 | 19.867 |  | 8 | 100 |  |  | 13.4481 | 13.4481
|  |  | 50 | 8 | 46.0273 | 5.1083 | 11.0985 |  | 8 | 100 |  |  | 15.3148 | 15.3148
|  |  | 100 | 8 | 97.2163 | 12.3493 | 12.7029 |  | 8 | 100 |  |  | 37.0231 | 37.0231
A | Hydroxyacetochlor | 68615 | 5 | 8 | 5.6114 | 2.6797 | 47.7556 |  | 8 | 100 | 17.3884 | 4.344 | 8.0339
|  |  | 10 | 8 | 9.9597 | 6.7165 | 67.4368 |  | 7 | 87.5 |  |  | 20.136
|  |  | 25 | 8 | 26.4076 | 5.7691 | 21.8462 | 17.2956 | 8 | 100 |  |  | 17.2956 | 17.2956 | 17.2956
|  |  | 50 | 8 | 51.1339 | 12.2638 | 23.9837 |  | 8 | 100 |  |  | 36.7669 | 36.7669
|  |  | 100 | 8 | 110.1541 | 21.7959 | 19.7867 |  | 8 | 100 |  |  | 65.3441 | 65.3441
A | Hydroxyalachlor | 68616 | 5 | 8 | 1.2096 | 1.6161 | 133.6072 | 4.8452 | 5 | 62.5 | 7.1952 | 3.768 | 4.8452 | 4.8452 | 4.8452
|  |  | 10 | 8 | 6.2151 | 2.4462 | 39.3595 |  | 8 | 100 |  |  | 7.3339 | 7.3339
|  |  | 25 | 8 | 21.2527 | 5.0941 | 23.9691 |  | 8 | 100 |  |  | 15.2721 | 15.2721
|  |  | 50 | 8 | 48.6144 | 8.5689 | 17.6263 |  | 8 | 100 |  |  | 25.6897 | 25.6897
|  |  | 100 | 8 | 111.9578 | 19.1474 | 17.1023 |  | 8 | 100 |  |  | 57.4038 | 57.4038
A | Hydroxydiazinon | 68618 | 5 | 8 | 4.2925 | 0.47499 | 11.0658 | 1.424 | 8 | 100 | 1.499 | 5.472 | 1.424 | 1.424 | 1.424
|  |  | 10 | 8 | 8.9322 | 0.55852 | 6.2529 |  | 8 | 100 |  |  | 1.6744 | 1.6744
|  |  | 25 | 8 | 24.7648 | 1.151 | 4.6478 |  | 8 | 100 |  |  | 3.4508 | 3.4508
|  |  | 50 | 8 | 45.0707 | 2.0635 | 4.5783 |  | 8 | 100 |  |  | 6.1863 | 6.1863
|  |  | 100 | 8 | 103.9829 | 7.2908 | 7.0115 |  | 8 | 100 |  |  | 21.8577 | 21.8577
A | Hydroxyfluometuron | 68620 | 5 | 8 | 1.7549 | 1.5142 | 86.2793 | 4.5394 | 5 | 62.5 | 3.8974 | 3.672 | 4.5394 | 4.5394 | 4.5394
|  |  | 10 | 8 | 8.3757 | 1.3378 | 15.9729 |  | 8 | 100 |  |  | 4.0108 | 4.0108
|  |  | 25 | 8 | 22.1648 | 1.2317 | 5.5569 |  | 8 | 100 |  |  | 3.6926 | 3.6926
|  |  | 50 | 8 | 46.2482 | 3.1521 | 6.8156 |  | 8 | 100 |  |  | 9.4499 | 9.4499
|  |  | 100 | 8 | 101.5079 | 5.4052 | 5.3249 |  | 8 | 100 |  |  | 16.2047 | 16.2047
A | Hydroxymetolachlor | 68622 | 5 | 8 | 4.7568 | 0.95732 | 20.1251 | 2.87 | 8 | 100 | 2.998 | 0.888 | 2.87 | 2.87 | 2.87
|  |  | 10 | 8 | 9.8255 | 0.75197 | 7.6532 |  | 8 | 100 |  |  | 2.2544 | 2.2544
|  |  | 25 | 8 | 24.9786 | 3.2744 | 13.1087 |  | 8 | 100 |  |  | 9.8165 | 9.8165
|  |  | 50 | 8 | 43.743 | 3.5484 | 8.1118 |  | 8 | 100 |  |  | 10.638 | 10.638
|  |  | 100 | 8 | 95.9485 | 6.5712 | 6.8487 |  | 8 | 100 |  |  | 19.7006 | 19.7006
A | Hydroxyphthalazinone | 68623 | 50 | 8 | 23.6875 | 7.6808 | 32.4256 |  | 8 | 100 | 23.0846 | 23 | 23.0271 | 23.0271
|  |  | 100 | 8 | 79.4337 | 8.714 | 10.9702 |  | 8 | 100 |  |  | 26.1246 | 26.1246
A | Hydroxysimazine | 68624 | 5 | 8 | 4.0738 | 0.68123 | 16.7223 | 2.0423 | 8 | 100 | 2.0986 | 1.248 | 2.0423 | 2.0423 | 2.0423
|  |  | 10 | 8 | 8.8673 | 0.80158 | 9.0398 |  | 8 | 100 |  |  | 2.4031 | 2.4031
|  |  | 25 | 8 | 25.2136 | 1.8092 | 7.1755 |  | 8 | 100 |  |  | 5.4239 | 5.4239
|  |  | 50 | 8 | 44.8084 | 4.834 | 10.7882 |  | 8 | 100 |  |  | 14.4924 | 14.4924
|  |  | 100 | 8 | 96.5502 | 12.1216 | 12.5547 |  | 8 | 100 |  |  | 36.3404 | 36.3404
A | Hydroxytebuthurion | 68621 | 5 | 8 | 1.8247 | 1.4347 | 78.627 | 4.3012 | 6 | 75 | 5.6962 | 5.352 | 4.3012 | 4.3012 | 4.3012
|  |  | 10 | 8 | 7.6719 | 1.9458 | 25.3634 |  | 8 | 100 |  |  | 5.8336 | 5.8336
|  |  | 25 | 8 | 22.3636 | 5.6344 | 25.1944 |  | 8 | 100 |  |  | 16.8918 | 16.8918
|  |  | 50 | 8 | 43.2782 | 9.7029 | 22.4199 |  | 8 | 100 |  |  | 29.0894 | 29.0894
|  |  | 100 | 8 | 97.9274 | 10.1917 | 10.4074 |  | 8 | 100 |  |  | 30.5547 | 30.5547
A | Imazamox | 68625 | 5 | 8 | 0 | 0 | 0 |  | 0 | 0 | 13.7908 | 14 | 0
|  |  | 10 | 8 | 11.4705 | 5.1175 | 44.6146 |  | 8 | 100 |  |  | 15.3424
|  |  | 25 | 8 | 30.8708 | 4.6116 | 14.9385 | 13.8256 | 8 | 100 |  |  | 13.8256 | 13.8256 | 13.8256
|  |  | 50 | 8 | 50.0978 | 6.7326 | 13.439 |  | 8 | 100 |  |  | 20.1844 | 20.1844
|  |  | 100 | 8 | 105.2687 | 11.4429 | 10.8702 |  | 8 | 100 |  |  | 34.3059 | 34.3059
A | Imazaquin | 61682 | 5 | 8 | 4.3559 | 1.9044 | 43.7211 |  | 7 | 87.5 | 11.3924 | 9.072 | 5.7095
|  |  | 10 | 8 | 10.2116 | 5.695 | 55.7693 |  | 8 | 100 |  |  | 17.0735
|  |  | 25 | 8 | 31.0059 | 3.8109 | 12.2909 | 11.4251 | 8 | 100 |  |  | 11.4251 | 11.4251 | 11.4251
|  |  | 50 | 8 | 46.7502 | 5.3665 | 11.479 |  | 8 | 100 |  |  | 16.0887 | 16.0887
|  |  | 100 | 8 | 97.3551 | 9.6001 | 9.8609 |  | 8 | 100 |  |  | 28.781 | 28.781
A | Imazethapyr | 61683 | 5 | 8 | 9.9618 | 3.0492 | 30.6084 |  | 8 | 100 | 8.6942 | 9.528 | 9.1414
|  |  | 10 | 8 | 12.991 | 2.9346 | 22.5897 | 8.798 | 8 | 100 |  |  | 8.798 | 8.798 | 8.798
|  |  | 25 | 8 | 29.8294 | 5.3446 | 17.9171 |  | 8 | 100 |  |  | 16.023 | 16.023
|  |  | 50 | 8 | 53.2774 | 12.0281 | 22.5764 |  | 8 | 100 |  |  | 36.0602 | 36.0602
|  |  | 100 | 8 | 103.32 | 19.1066 | 18.4926 |  | 8 | 100 |  |  | 57.2815 | 57.2815
A | Imidacloprid | 68426 | 5 | 8 | 3.543 | 1.7576 | 49.6067 | 5.2692 | 7 | 87.5 | 9.5936 | 5.4 | 5.2692 | 5.2692 | 5.2692
|  |  | 10 | 8 | 10.1856 | 3.2065 | 31.4808 |  | 8 | 100 |  |  | 9.6131 | 9.6131
|  |  | 25 | 8 | 22.7134 | 3.2866 | 14.4698 |  | 8 | 100 |  |  | 9.8532 | 9.8532
|  |  | 50 | 8 | 46.3218 | 5.0194 | 10.8359 |  | 8 | 100 |  |  | 15.0482 | 15.0482
|  |  | 100 | 8 | 97.0579 | 8.2281 | 8.4775 |  | 8 | 100 |  |  | 24.6679 | 24.6679
A | Indoxacarb | 68627 | 5 | 8 | 3.3708 | 1.5961 | 47.3507 | 4.7852 | 8 | 100 | 4.7968 | 3.816 | 4.7852 | 4.7852 | 4.7852
|  |  | 10 | 8 | 5.8537 | 2.5212 | 43.0694 |  | 8 | 100 |  |  | 7.5584 | 7.5584
|  |  | 25 | 8 | 23.0451 | 4.1307 | 17.9242 |  | 8 | 100 |  |  | 12.3837 | 12.3837
|  |  | 50 | 8 | 40.4831 | 6.6169 | 16.3449 |  | 8 | 100 |  |  | 19.8375 | 19.8375
|  |  | 100 | 8 | 95.5813 | 5.7751 | 6.0421 |  | 8 | 100 |  |  | 17.3137 | 17.3137
A | Isoxaflutole | 68632 | 5 | 8 | 9.5816 | 1.616 | 16.8659 | 4.8448 | 8 | 100 | 19.7868 | 6.36 | 4.8448 | 4.8448 | 4.8448
|  |  | 10 | 8 | 17.0318 | 5.039 | 29.586 |  | 8 | 100 |  |  | 15.107
|  |  | 25 | 8 | 29.0013 | 6.6328 | 22.8706 | 19.885 | 8 | 100 |  |  | 19.885 | 19.885 | 19.885
|  |  | 50 | 8 | 52.7364 | 11.0452 | 20.9443 |  | 8 | 100 |  |  | 33.1137 | 33.1137
|  |  | 100 | 8 | 99.6899 | 11.8926 | 11.9295 |  | 8 | 100 |  |  | 35.6539 | 35.6539
A | Kresoxim-methyl | 67670 | 5 | 8 | 4.1334 | 0.45223 | 10.9408 | 1.3558 | 8 | 100 | 1.499 | 2.28 | 1.3558 | 1.3558 | 1.3558
|  |  | 10 | 8 | 8.8954 | 1.0939 | 12.2969 |  | 8 | 100 |  |  | 3.2794 | 3.2794
|  |  | 25 | 8 | 23.208 | 2.5024 | 10.7825 |  | 8 | 100 |  |  | 7.5022 | 7.5022
|  |  | 50 | 8 | 41.8413 | 1.9331 | 4.62 |  | 8 | 100 |  |  | 5.7954 | 5.7954
|  |  | 100 | 8 | 94.9841 | 2.2257 | 2.3432 |  | 8 | 100 |  |  | 6.6727 | 6.6727
A | Lactofen | 68638 | 5 | 8 | 4.8413 | 0.55084 | 11.378 | 1.6514 | 8 | 100 | 1.7988 | 1.44 | 1.6514 | 1.6514 | 1.6514
|  |  | 10 | 8 | 8.1202 | 0.59727 | 7.3554 |  | 8 | 100 |  |  | 1.7906 | 1.7906
|  |  | 25 | 8 | 21.7982 | 1.3593 | 6.2357 |  | 8 | 100 |  |  | 4.0751 | 4.0751
|  |  | 50 | 8 | 40.3971 | 2.7309 | 6.7601 |  | 8 | 100 |  |  | 8.1871 | 8.1871
|  |  | 100 | 8 | 95.5239 | 7.5922 | 7.9479 |  | 8 | 100 |  |  | 22.7613 | 22.7613
A | Linuron | 68639 | 5 | 8 | 6.8985 | 2.2515 | 32.6373 |  | 8 | 100 | 5.996 | 2.568 | 6.7499
|  |  | 10 | 8 | 9.4083 | 2.0432 | 21.7175 | 6.1256 | 8 | 100 |  |  | 6.1256 | 6.1256 | 6.1256
|  |  | 25 | 8 | 24.8349 | 3.704 | 14.9144 |  | 8 | 100 |  |  | 11.1045 | 11.1045
|  |  | 50 | 8 | 45.9021 | 8.9191 | 19.4308 |  | 8 | 100 |  |  | 26.7396 | 26.7396
|  |  | 100 | 8 | 114.5689 | 15.8518 | 13.8361 |  | 8 | 100 |  |  | 47.5237 | 47.5237
A | Malaoxon | 68240 | 5 | 8 | 5.1332 | 0.67987 | 13.2447 | 2.0383 | 8 | 100 | 2.0986 | 1.224 | 2.0383 | 2.0383 | 2.0383
|  |  | 10 | 8 | 8.6148 | 1.0104 | 11.729 |  | 8 | 100 |  |  | 3.0293 | 3.0293
|  |  | 25 | 8 | 23.8659 | 1.3485 | 5.6501 |  | 8 | 100 |  |  | 4.0427 | 4.0427
|  |  | 50 | 8 | 45.6968 | 2.4554 | 5.3732 |  | 8 | 100 |  |  | 7.3612 | 7.3612
|  |  | 100 | 8 | 98.2229 | 2.6217 | 2.6691 |  | 8 | 100 |  |  | 7.8598 | 7.8598
A | Malathion | 65087 | 5 | 8 | 6.2055 | 1.2949 | 20.8661 | 3.882 | 8 | 100 | 3.8974 | 1.896 | 3.882 | 3.882 | 3.882
|  |  | 10 | 8 | 10.9065 | 1.9504 | 17.8825 |  | 8 | 100 |  |  | 5.8472 | 5.8472
|  |  | 25 | 8 | 25.9703 | 3.8496 | 14.8232 |  | 8 | 100 |  |  | 11.5412 | 11.5412
|  |  | 50 | 8 | 49.4698 | 3.1337 | 6.3346 |  | 8 | 100 |  |  | 9.3948 | 9.3948
|  |  | 100 | 8 | 99.0837 | 7.276 | 7.3432 |  | 8 | 100 |  |  | 21.8133 | 21.8133
A | Metalaxyl | 68437 | 5 | 8 | 6.08 | 1.2481 | 20.5281 | 3.7418 | 8 | 100 | 3.5976 | 2.76 | 3.7418 | 3.7418 | 3.7418
|  |  | 10 | 8 | 9.54 | 1.2477 | 13.0788 |  | 8 | 100 |  |  | 3.7406 | 3.7406
|  |  | 25 | 8 | 25.4248 | 2.593 | 10.1987 |  | 8 | 100 |  |  | 7.7738 | 7.7738
|  |  | 50 | 8 | 45.2574 | 2.9657 | 6.5529 |  | 8 | 100 |  |  | 8.8911 | 8.8911
|  |  | 100 | 8 | 101.6059 | 5.2411 | 5.1583 |  | 8 | 100 |  |  | 15.7129 | 15.7129
A | Metconazole | 66620 | 5 | 8 | 4.3873 | 0.78639 | 17.9242 | 2.3576 | 8 | 100 | 2.3984 | 2.16 | 2.3576 | 2.3576 | 2.3576
|  |  | 10 | 8 | 8.6011 | 0.8959 | 10.4161 |  | 8 | 100 |  |  | 2.6859 | 2.6859
|  |  | 25 | 8 | 23.4445 | 1.559 | 6.6496 |  | 8 | 100 |  |  | 4.6738 | 4.6738
|  |  | 50 | 8 | 42.6578 | 2.4172 | 5.6664 |  | 8 | 100 |  |  | 7.2467 | 7.2467
|  |  | 100 | 8 | 97.9375 | 5.1958 | 5.3053 |  | 8 | 100 |  |  | 15.5771 | 15.5771
A | Methamidophos | 68644 | 5 | 8 | 5.4759 | 2.2416 | 40.9356 |  | 7 | 87.5 | 2.3984 | 1.272 | 6.7203
|  |  | 10 | 8 | 10.9229 | 0.82476 | 7.5507 | 2.4726 | 8 | 100 |  |  | 2.4726 | 2.4726 | 2.4726
|  |  | 25 | 8 | 26.3954 | 1.1305 | 4.2831 |  | 8 | 100 |  |  | 3.3894 | 3.3894
|  |  | 50 | 8 | 46.9235 | 2.7718 | 5.907 |  | 8 | 100 |  |  | 8.3097 | 8.3097
|  |  | 100 | 8 | 102.6619 | 3.6846 | 3.589 |  | 8 | 100 |  |  | 11.0464 | 11.0464
A | Methidathion | 65088 | 5 | 8 | 4.0004 | 0.64578 | 16.1431 | 1.9361 | 8 | 100 | 1.7988 | 5 | 1.9361 | 1.9361 | 1.9361
|  |  | 10 | 8 | 8.3735 | 0.90397 | 10.7957 |  | 8 | 100 |  |  | 2.7101 | 2.7101
|  |  | 25 | 8 | 23.7335 | 2.7018 | 11.3838 |  | 8 | 100 |  |  | 8.0999 | 8.0999
|  |  | 50 | 8 | 43.995 | 5.422 | 12.3241 |  | 8 | 100 |  |  | 16.2551 | 16.2551
|  |  | 100 | 8 | 97.0607 | 9.0437 | 9.3176 |  | 8 | 100 |  |  | 27.113 | 27.113
A | Methomyl | 68645 | 5 | 8 | 4.4769 | 0.40701 | 9.0914 | 1.2202 | 8 | 100 | 1.1992 | 1.08 | 1.2202 | 1.2202 | 1.2202
|  |  | 10 | 8 | 9.148 | 0.96259 | 10.5224 |  | 8 | 100 |  |  | 2.8858 | 2.8858
|  |  | 25 | 8 | 23.7205 | 1.5981 | 6.7371 |  | 8 | 100 |  |  | 4.791 | 4.791
|  |  | 50 | 8 | 41.1315 | 2.7491 | 6.6836 |  | 8 | 100 |  |  | 8.2417 | 8.2417
|  |  | 100 | 8 | 93.9463 | 5.5137 | 5.869 |  | 8 | 100 |  |  | 16.5301 | 16.5301
A | Methomyl oxime | 68646 | 5 | 8 | 0 | 0 | 0 |  | 0 | 0 | 38.6742 |  | 0
|  |  | 10 | 8 | 0 | 0 | 0 |  | 0 | 0 |  |  | 0
|  |  | 25 | 8 | 0 | 0 | 0 |  | 0 | 0 |  |  | 0
|  |  | 50 | 7 | 58.5954 | 22.2896 | 38.0398 |  | 7 | 100 |  |  | 66.8243
|  |  | 100 | 8 | 105.1836 | 12.9181 | 12.2814 | 38.7283 | 8 | 100 |  |  | 38.7283 | 38.7283 | 38.7283
A | Methoxyfenozide | 68647 | 5 | 8 | 3.8585 | 0.41914 | 10.8627 | 1.2566 | 8 | 100 | 1.1992 | 0.96 | 1.2566 | 1.2566 | 1.2566
|  |  | 10 | 8 | 8.4503 | 0.94771 | 11.2151 |  | 8 | 100 |  |  | 2.8412 | 2.8412
|  |  | 25 | 8 | 25.0986 | 3.1485 | 12.5444 |  | 8 | 100 |  |  | 9.4391 | 9.4391
|  |  | 50 | 8 | 47.1764 | 6.605 | 14.0007 |  | 8 | 100 |  |  | 19.8019 | 19.8019
|  |  | 100 | 8 | 111.4325 | 15.2223 | 13.6606 |  | 8 | 100 |  |  | 45.6366 | 45.6366
A | Metolachlor | 65090 | 5 | 8 | 4.5431 | 0.53909 | 11.866 | 1.6162 | 8 | 100 | 1.499 | 4.2 | 1.6162 | 1.6162 | 1.6162
|  |  | 10 | 8 | 8.6741 | 0.56005 | 6.4566 |  | 8 | 100 |  |  | 1.679 | 1.679
|  |  | 25 | 8 | 23.6529 | 1.2225 | 5.1686 |  | 8 | 100 |  |  | 3.6652 | 3.6652
|  |  | 50 | 8 | 43.7168 | 1.7257 | 3.9475 |  | 8 | 100 |  |  | 5.1737 | 5.1737
|  |  | 100 | 8 | 98.3906 | 4.1194 | 4.1868 |  | 8 | 100 |  |  | 12.3499 | 12.3499
A | Metolachlor hydroxy morpholinone | 68649 | 5 | 8 | 2.1451 | 1.3073 | 60.9433 | 3.9193 | 7 | 87.5 | 5.3964 | 3.048 | 3.9193 | 3.9193 | 3.9193
|  |  | 10 | 8 | 7.9036 | 1.8278 | 23.1259 |  | 8 | 100 |  |  | 5.4797 | 5.4797
|  |  | 25 | 8 | 24.5378 | 2.836 | 11.5575 |  | 8 | 100 |  |  | 8.5022 | 8.5022
|  |  | 50 | 8 | 46.7729 | 7.9796 | 17.0603 |  | 8 | 100 |  |  | 23.9228 | 23.9228
|  |  | 100 | 8 | 112.0677 | 17.0194 | 15.1867 |  | 8 | 100 |  |  | 51.0241 | 51.0241
A | Metribuzin | 68652 | 5 | 8 | 0 | 0 | #DIV/0! |  | 0 | 0 | 20.3864 | 9.744 | 0
|  |  | 10 | 8 | 12.0552 | 5.994 | 49.7215 |  | 7 | 87.5 |  |  | 17.9701
|  |  | 25 | 8 | 29.553 | 6.8438 | 23.1576 | 20.5177 | 8 | 100 |  |  | 20.5177 | 20.5177 | 20.5177
|  |  | 50 | 8 | 49.1181 | 8.4267 | 17.1559 |  | 8 | 100 |  |  | 25.2631 | 25.2631
|  |  | 100 | 8 | 102.0534 | 12.3321 | 12.0839 |  | 8 | 100 |  |  | 36.9715 | 36.9715
A | Metribuzin DA | 68568 | 5 | 8 | 3.8455 | 2.464 | 64.0764 |  | 7 | 87.5 | 11.6922 | 4.392 | 7.3872
|  |  | 10 | 8 | 9.9434 | 4.6795 | 47.0612 |  | 8 | 100 |  |  | 14.0291
|  |  | 25 | 8 | 28.6857 | 3.8992 | 13.5927 | 11.6897 | 8 | 100 |  |  | 11.6897 | 11.6897 | 11.6897
|  |  | 50 | 8 | 50.1063 | 3.4391 | 6.8637 |  | 8 | 100 |  |  | 10.3105 | 10.3105
|  |  | 100 | 8 | 103.9935 | 7.8813 | 7.5786 |  | 8 | 100 |  |  | 23.628 | 23.628
A | Molinate | 65091 | 10 | 8 | 5.9906 | 3.1955 | 53.3418 |  | 8 | 100 | 9.5936 | 8.544 | 9.58 | 9.58
|  |  | 25 | 8 | 20.8058 | 9.8522 | 47.3532 |  | 8 | 100 |  |  | 29.537
|  |  | 50 | 8 | 40.2778 | 11.2159 | 27.8464 | 33.6254 | 8 | 100 |  |  | 33.6254 | 33.6254 | 33.6254
|  |  | 100 | 8 | 107.8994 | 14.7868 | 13.7043 |  | 8 | 100 |  |  | 44.3308 | 44.3308
A | Myclobutanil | 66632 | 5 | 8 | 5.3437 | 1.1282 | 21.1124 | 3.3823 | 8 | 100 | 3.2978 | 3.384 | 3.3823 | 3.3823 | 3.3823
|  |  | 10 | 8 | 8.4793 | 2.8283 | 33.3552 |  | 8 | 100 |  |  | 8.4792 | 8.4792
|  |  | 25 | 8 | 24.8664 | 2.9169 | 11.7304 |  | 8 | 100 |  |  | 8.7449 | 8.7449
|  |  | 50 | 8 | 43.0375 | 3.1714 | 7.3689 |  | 8 | 100 |  |  | 9.5078 | 9.5078
|  |  | 100 | 8 | 101.6185 | 6.9295 | 6.8191 |  | 8 | 100 |  |  | 20.7746 | 20.7746
A | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 5 | 8 | 6.1931 | 2.2833 | 36.8681 |  | 8 | 100 | 8.6942 | 2.496 | 6.8453
|  |  | 10 | 8 | 12.9306 | 2.9095 | 22.5011 | 8.7228 | 8 | 100 |  |  | 8.7228 | 8.7228 | 8.7228
|  |  | 25 | 8 | 26.1185 | 2.0492 | 7.846 |  | 8 | 100 |  |  | 6.1436 | 6.1436
|  |  | 50 | 8 | 51.3353 | 5.4438 | 10.6044 |  | 8 | 100 |  |  | 16.3205 | 16.3205
|  |  | 100 | 8 | 108.4995 | 11.4887 | 10.5887 |  | 8 | 100 |  |  | 34.4432 | 34.4432
E | Naled | 68654 | 5 | 8 | 2.488 | 1.7769 | 71.4206 | 5.3273 | 7 | 87.5 | 14.6902 | 15 | 5.3273 | 5.3273 | 5.3273
|  |  | 10 | 8 | 4.3688 | 2.9436 | 67.3782 |  | 7 | 87.5 |  |  | 8.825 | 8.825
|  |  | 25 | 8 | 23.4322 | 4.8548 | 20.7183 |  | 8 | 100 |  |  | 14.5546 | 14.5546
|  |  | 50 | 8 | 42.5334 | 8.0813 | 18.9999 |  | 8 | 100 |  |  | 24.2277 | 24.2277
|  |  | 100 | 8 | 92.655 | 9.4962 | 10.249 |  | 8 | 100 |  |  | 28.4695 | 28.4695
A | Nicosulfuron | 61685 | 5 | 8 | 13.3732 | 2.8231 | 21.11 |  | 8 | 100 | 6.8954 | 7 | 8.4636
|  |  | 10 | 8 | 20.6075 | 2.328 | 11.2969 | 6.9794 | 8 | 100 |  |  | 6.9794 | 6.9794 | 6.9794
|  |  | 25 | 8 | 36.0364 | 2.9014 | 8.0514 |  | 8 | 100 |  |  | 8.6985 | 8.6985
|  |  | 50 | 8 | 55.783 | 6.9821 | 12.5166 |  | 8 | 100 |  |  | 20.9325 | 20.9325
|  |  | 100 | 8 | 109.751 | 10.0269 | 9.136 |  | 8 | 100 |  |  | 30.0605 | 30.0605
A | Norflurazon | 67685 | 5 | 8 | 3.7134 | 0.56819 | 15.301 | 1.7034 | 8 | 100 | 1.7988 | 1.512 | 1.7034 | 1.7034 | 1.7034
|  |  | 10 | 8 | 8.6988 | 1.1312 | 13.0039 |  | 8 | 100 |  |  | 3.3913 | 3.3913
|  |  | 25 | 8 | 24.1016 | 1.5051 | 6.2448 |  | 8 | 100 |  |  | 4.5123 | 4.5123
|  |  | 50 | 8 | 44.7282 | 3.7103 | 8.2952 |  | 8 | 100 |  |  | 11.1234 | 11.1234
|  |  | 100 | 8 | 98.6708 | 4.6309 | 4.6933 |  | 8 | 100 |  |  | 13.8833 | 13.8833
E | Novaluron | 68655 | 5 | 8 | 4.0551 | 1.3236 | 32.6396 | 3.9681 | 8 | 100 | 3.8974 | 4.032 | 3.9681 | 3.9681 | 3.9681
|  |  | 10 | 8 | 7.7185 | 2.2061 | 28.5823 |  | 8 | 100 |  |  | 6.6139 | 6.6139
|  |  | 25 | 8 | 20.2117 | 4.0035 | 19.8078 |  | 8 | 100 |  |  | 12.0025 | 12.0025
|  |  | 50 | 8 | 41.2483 | 5.3131 | 12.8807 |  | 8 | 100 |  |  | 15.9286 | 15.9286
|  |  | 100 | 8 | 88.0525 | 3.5889 | 4.0758 |  | 8 | 100 |  |  | 10.7595 | 10.7595
A | O-Ethyl-O-methyl-S-propyl phosphorothioate | 68597 | 5 | 8 | 4.0031 | 0.96314 | 24.0597 | 2.8875 | 8 | 100 | 2.998 | 2.28 | 2.8875 | 2.8875 | 2.8875
|  |  | 10 | 8 | 7.7831 | 0.91557 | 11.7636 |  | 8 | 100 |  |  | 2.7449 | 2.7449
|  |  | 25 | 8 | 22.9257 | 2.62 | 11.4284 |  | 8 | 100 |  |  | 7.8549 | 7.8549
|  |  | 50 | 8 | 44.1869 | 3.7361 | 8.4552 |  | 8 | 100 |  |  | 11.2008 | 11.2008
|  |  | 100 | 8 | 101.6837 | 4.3033 | 4.2321 |  | 8 | 100 |  |  | 12.9013 | 12.9013
A | O-Ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 5 | 8 | 4.7124 | 0.72077 | 15.2952 | 2.1609 | 8 | 100 | 2.0986 | 2.04 | 2.1609 | 2.1609 | 2.1609
|  |  | 10 | 8 | 8.3859 | 1.5383 | 18.3436 |  | 8 | 100 |  |  | 4.6117 | 4.6117
|  |  | 25 | 8 | 23.9115 | 1.7317 | 7.2421 |  | 8 | 100 |  |  | 5.1916 | 5.1916
|  |  | 50 | 8 | 46.1797 | 4.1024 | 8.8836 |  | 8 | 100 |  |  | 12.299 | 12.299
|  |  | 100 | 8 | 102.0509 | 5.613 | 5.5002 |  | 8 | 100 |  |  | 16.8278 | 16.8278
A | O-Ethyl-S-propyl phosphorothioate | 68658 | 25 | 8 | 12.844 | 8.7119 | 67.8283 |  | 7 | 87.5 | 31.7788 | 32 | 26.1181 | 26.1181
|  |  | 50 | 8 | 42.9785 | 10.5578 | 24.5652 |  | 8 | 100 |  |  | 31.6522 | 31.6522
|  |  | 100 | 8 | 101.1678 | 18.3737 | 18.1617 |  | 8 | 100 |  |  | 55.0845 | 55.0845
A | Orthosulfamuron | 68662 | 5 | 8 | 5.7501 | 0.9546 | 16.6015 | 2.8619 | 8 | 100 | 2.998 | 3 | 2.8619 | 2.8619 | 2.8619
|  |  | 10 | 8 | 10.5549 | 1.8212 | 17.255 |  | 8 | 100 |  |  | 5.4601 | 5.4601
|  |  | 25 | 8 | 23.1306 | 3.1864 | 13.7759 |  | 8 | 100 |  |  | 9.5529 | 9.5529
|  |  | 50 | 8 | 40.1946 | 3.5189 | 8.7547 |  | 8 | 100 |  |  | 10.5497 | 10.5497
|  |  | 100 | 8 | 91.3784 | 6.5902 | 7.212 |  | 8 | 100 |  |  | 19.7575 | 19.7575
A | Oxamyl | 68664 | 5 | 8 | 4.2017 | 1.1285 | 26.8569 | 3.3831 | 8 | 100 | 3.2978 | 0.816 | 3.3831 | 3.3831 | 3.3831
|  |  | 10 | 8 | 10.2836 | 1.0017 | 9.7412 |  | 8 | 100 |  |  | 3.0032 | 3.0032
|  |  | 25 | 8 | 24.8529 | 1.9674 | 7.9161 |  | 8 | 100 |  |  | 5.8982 | 5.8982
|  |  | 50 | 8 | 42.2998 | 2.6308 | 6.2195 |  | 8 | 100 |  |  | 7.8872 | 7.8872
|  |  | 100 | 8 | 100.26 | 5.5036 | 5.4894 |  | 8 | 100 |  |  | 16.4999 | 16.4999
A | Oxamyl oxime | 68665 | 5 | 8 | 4.4271 | 0.1813 | 4.0951 | 0.54352 | 8 | 100 |  | 0.696 | 0.54352 | 0.54352 | 0.54352
|  |  | 10 | 8 | 8.8247 | 0.2861 | 3.242 |  | 8 | 100 |  |  | 0.85772 | 0.85772
|  |  | 25 | 8 | 23.6237 | 0.85924 | 3.6372 |  | 8 | 100 |  |  | 2.576 | 2.576
|  |  | 50 | 8 | 41.0913 | 1.3775 | 3.3524 |  | 8 | 100 |  |  | 4.1298 | 4.1298
|  |  | 100 | 8 | 94.4655 | 5.3661 | 5.6805 |  | 8 | 100 |  |  | 16.0875 | 16.0875
A | Paraoxon | 68666 | 5 | 8 | 4.6935 | 1.4317 | 30.5032 | 4.2921 | 8 | 100 | 4.1972 | 1.536 | 4.2921 | 4.2921 | 4.2921
|  |  | 10 | 8 | 8.5431 | 1.4867 | 17.402 |  | 8 | 100 |  |  | 4.457 | 4.457
|  |  | 25 | 8 | 25.3291 | 2.3187 | 9.1542 |  | 8 | 100 |  |  | 6.9514 | 6.9514
|  |  | 50 | 8 | 46.0055 | 3.3784 | 7.3436 |  | 8 | 100 |  |  | 10.1286 | 10.1286
|  |  | 100 | 8 | 105.7602 | 4.7948 | 4.5336 |  | 8 | 100 |  |  | 14.3748 | 14.3748
A | Paraoxon-methyl | 68648 | 10 | 7 | 6.0634 | 4.7689 | 78.6504 |  | 7 | 100 | 46.7688 | 9.312 | 14.2971
|  |  | 25 | 8 | 22.9804 | 12.0006 | 52.2209 |  | 8 | 100 |  |  | 35.9777
|  |  | 50 | 8 | 39.6355 | 15.5683 | 39.2788 | 46.6739 | 8 | 100 |  |  | 46.6739 | 46.6739 | 46.6739
|  |  | 100 | 8 | 103.0039 | 19.4534 | 18.8861 |  | 8 | 100 |  |  | 58.3214 | 58.3214
A | Pendimethalin | 65098 | 5 | 8 | 4.7286 | 0.85123 | 18.0019 | 2.552 | 8 | 100 | 2.6982 | 1.416 | 2.552 | 2.552 | 2.552
|  |  | 10 | 8 | 8.8409 | 0.49291 | 5.5754 |  | 8 | 100 |  |  | 1.4777 | 1.4777
|  |  | 25 | 8 | 23.0305 | 2.0232 | 8.7848 |  | 8 | 100 |  |  | 6.0655 | 6.0655
|  |  | 50 | 8 | 42.5124 | 2.6906 | 6.3289 |  | 8 | 100 |  |  | 8.0663 | 8.0663
|  |  | 100 | 8 | 92.4765 | 4.6438 | 5.0216 |  | 8 | 100 |  |  | 13.9222 | 13.9222
A | Phorate | 68668 | 5 | 8 | 3.0094 | 1.0615 | 35.2736 | 3.1824 | 8 | 100 | 3.2978 | 3 | 3.1824 | 3.1824 | 3.1824
|  |  | 10 | 8 | 6.7953 | 2.031 | 29.8884 |  | 8 | 100 |  |  | 6.0889 | 6.0889
|  |  | 25 | 8 | 18.7128 | 3.6039 | 19.2592 |  | 8 | 100 |  |  | 10.8046 | 10.8046
|  |  | 50 | 8 | 36.8003 | 6.1588 | 16.7358 |  | 8 | 100 |  |  | 18.4642 | 18.4642
|  |  | 100 | 8 | 82.7641 | 4.7467 | 5.7352 |  | 8 | 100 |  |  | 14.2307 | 14.2307
A | Phorate oxon | 68669 | 5 | 8 | 4.2014 | 0.68002 | 16.1857 | 2.0387 | 8 | 100 | 2.0986 | 2 | 2.0387 | 2.0387 | 2.0387
|  |  | 10 | 8 | 8.3882 | 1.0842 | 12.9255 |  | 8 | 100 |  |  | 3.2505 | 3.2505
|  |  | 25 | 8 | 23.3517 | 2.0615 | 8.828 |  | 8 | 100 |  |  | 6.1803 | 6.1803
|  |  | 50 | 8 | 45.9361 | 2.3708 | 5.1612 |  | 8 | 100 |  |  | 7.1078 | 7.1078
|  |  | 100 | 8 | 98.7837 | 4.712 | 4.77 |  | 8 | 100 |  |  | 14.1266 | 14.1266
A | Phorate oxon sulfone | 68670 | 5 | 8 | 3.4856 | 1.8821 | 53.9965 |  | 8 | 100 | 5.6962 | 9.72 | 5.6425
|  |  | 10 | 8 | 8.16 | 3.3661 | 41.251 | 10.0915 | 8 | 100 |  |  | 10.0915 | 10.0915 | 10.0915
|  |  | 25 | 8 | 24.5061 | 4.6637 | 19.0309 |  | 8 | 100 |  |  | 13.9818 | 13.9818
|  |  | 50 | 8 | 45.8858 | 6.9618 | 15.172 |  | 8 | 100 |  |  | 20.8714 | 20.8714
|  |  | 100 | 8 | 97.9072 | 10.3502 | 10.5714 |  | 8 | 100 |  |  | 31.0299 | 31.0299
A | Phorate oxon sulfoxide | 68671 | 5 | 8 | 2.9712 | 0.89603 | 30.1571 | 2.6863 | 8 | 100 | 2.6982 | 3.432 | 2.6863 | 2.6863 | 2.6863
|  |  | 10 | 8 | 7.1951 | 1.3938 | 19.3717 |  | 8 | 100 |  |  | 4.1786 | 4.1786
|  |  | 25 | 8 | 22.1298 | 2.8404 | 12.8352 |  | 8 | 100 |  |  | 8.5155 | 8.5155
|  |  | 50 | 8 | 43.4665 | 3.1219 | 7.1823 |  | 8 | 100 |  |  | 9.3595 | 9.3595
|  |  | 100 | 8 | 96.0835 | 4.8545 | 5.0523 |  | 8 | 100 |  |  | 14.5537 | 14.5537
A | Phorate sulfone | 68672 | 10 | 8 | 9.7062 | 4.3308 | 44.6189 |  | 8 | 100 | 29.3804 | 4.44 | 12.9838
|  |  | 25 | 8 | 31.6267 | 12.3415 | 39.0223 |  | 8 | 100 |  |  | 36.9997
|  |  | 50 | 8 | 51.1027 | 9.8325 | 19.2406 | 29.4777 | 8 | 100 |  |  | 29.4777 | 29.4777 | 29.4777
|  |  | 100 | 8 | 100.2281 | 13.5358 | 13.505 |  | 8 | 100 |  |  | 40.5805 | 40.5805
A | Phorate sulfoxide | 68673 | 5 | 8 | 4.6007 | 1.2225 | 26.572 | 3.665 | 8 | 100 | 3.5976 | 2.424 | 3.665 | 3.665 | 3.665
|  |  | 10 | 8 | 9.4472 | 1.4722 | 15.5835 |  | 8 | 100 |  |  | 4.4137 | 4.4137
|  |  | 25 | 8 | 24.6658 | 2.5127 | 10.187 |  | 8 | 100 |  |  | 7.5331 | 7.5331
|  |  | 50 | 8 | 44.2276 | 3.4782 | 7.8643 |  | 8 | 100 |  |  | 10.4277 | 10.4277
|  |  | 100 | 8 | 97.42 | 4.0092 | 4.1154 |  | 8 | 100 |  |  | 12.0196 | 12.0196
E | Phosmet | 65101 | 5 | 8 | 3.4538 | 1.3625 | 39.4487 | 4.0848 | 8 | 100 | 4.1972 | 10 | 4.0848 | 4.0848 | 4.0848
|  |  | 10 | 8 | 9.5914 | 1.9562 | 20.3953 |  | 8 | 100 |  |  | 5.8647 | 5.8647
|  |  | 25 | 8 | 11.106 | 6.1172 | 55.0808 |  | 8 | 100 |  |  | 18.3395 | 18.3395
|  |  | 50 | 8 | 38.5161 | 8.6311 | 22.4091 |  | 8 | 100 |  |  | 25.876 | 25.876
|  |  | 100 | 8 | 58.5653 | 38.9623 | 66.5279 |  | 8 | 100 |  |  | 116.8089
A | Phthalazinone | 68675 | 10 | 8 | 12.4034 | 2.5443 | 20.5133 | 7.6279 | 8 | 100 | 7.495 | 2.712 | 7.6279 | 7.6279 | 7.6279
|  |  | 25 | 8 | 26.5949 | 2.4207 | 9.1022 |  | 8 | 100 |  |  | 7.2573 | 7.2573
|  |  | 50 | 8 | 53.3297 | 8.4285 | 15.8045 |  | 8 | 100 |  |  | 25.2686 | 25.2686
|  |  | 100 | 8 | 102.4438 | 14.7223 | 14.3711 |  | 8 | 100 |  |  | 44.1373 | 44.1373
A | Piperonyl butoxide | 65102 | 5 | 8 | 4.1667 | 0.32183 | 7.724 | 0.96485 | 8 | 100 | 0.8994 | 1.752 | 0.96485 | 0.96485 | 0.96485
|  |  | 10 | 8 | 8.5512 | 0.29513 | 3.4514 |  | 8 | 100 |  |  | 0.88481 | 0.88481
|  |  | 25 | 8 | 23.7843 | 0.66704 | 2.8045 |  | 8 | 100 |  |  | 1.9998 | 1.9998
|  |  | 50 | 8 | 43.6052 | 1.3035 | 2.9894 |  | 8 | 100 |  |  | 3.908 | 3.908
|  |  | 100 | 8 | 98.4742 | 4.1923 | 4.2572 |  | 8 | 100 |  |  | 12.5684 | 12.5684
A | Profenofos | 68676 | 5 | 8 | 4.405 | 0.94469 | 21.4459 | 2.8322 | 8 | 100 | 2.6982 | 1.248 | 2.8322 | 2.8322 | 2.8322
|  |  | 10 | 8 | 9.2651 | 1.0058 | 10.8557 |  | 8 | 100 |  |  | 3.0154 | 3.0154
|  |  | 25 | 8 | 25.3324 | 1.6861 | 6.656 |  | 8 | 100 |  |  | 5.055 | 5.055
|  |  | 50 | 8 | 46.6763 | 2.8028 | 6.0048 |  | 8 | 100 |  |  | 8.4029 | 8.4029
|  |  | 100 | 8 | 100.1442 | 5.2976 | 5.2899 |  | 8 | 100 |  |  | 15.8821 | 15.8821
A | Prometon | 67702 | 5 | 8 | 5.1477 | 0.55119 | 10.7076 | 1.6525 | 8 | 100 | 1.7988 | 1.56 | 1.6525 | 1.6525 | 1.6525
|  |  | 10 | 8 | 9.3952 | 1.2705 | 13.5228 |  | 8 | 100 |  |  | 3.8089 | 3.8089
|  |  | 25 | 8 | 24.8809 | 2.5231 | 10.1408 |  | 8 | 100 |  |  | 7.5644 | 7.5644
|  |  | 50 | 8 | 42.7922 | 2.463 | 5.7556 |  | 8 | 100 |  |  | 7.384 | 7.384
|  |  | 100 | 8 | 94.4012 | 5.6538 | 5.9891 |  | 8 | 100 |  |  | 16.95 | 16.95
A | Prometryn | 65103 | 5 | 8 | 4.6884 | 0.44295 | 9.4478 | 1.328 | 8 | 100 | 1.1992 | 0.96 | 1.328 | 1.328 | 1.328
|  |  | 10 | 8 | 8.8767 | 0.72369 | 8.1527 |  | 8 | 100 |  |  | 2.1696 | 2.1696
|  |  | 25 | 8 | 24.5996 | 1.5149 | 6.158 |  | 8 | 100 |  |  | 4.5415 | 4.5415
|  |  | 50 | 8 | 43.8178 | 3.0692 | 7.0044 |  | 8 | 100 |  |  | 9.2015 | 9.2015
|  |  | 100 | 8 | 97.4229 | 4.6686 | 4.7921 |  | 8 | 100 |  |  | 13.9966 | 13.9966
A | Pronamide | 67706 | 5 | 8 | 4.0011 | 0.99067 | 24.7602 | 2.97 | 8 | 100 | 2.998 | 1.92 | 2.97 | 2.97 | 2.97
|  |  | 10 | 8 | 9.3343 | 2.6601 | 28.498 |  | 8 | 100 |  |  | 7.9749 | 7.9749
|  |  | 25 | 8 | 25.6417 | 3.2699 | 12.7522 |  | 8 | 100 |  |  | 9.8031 | 9.8031
|  |  | 50 | 8 | 49.2633 | 4.0872 | 8.2967 |  | 8 | 100 |  |  | 12.2535 | 12.2535
|  |  | 100 | 8 | 113.1063 | 13.7625 | 12.1677 |  | 8 | 100 |  |  | 41.2599 | 41.2599
A | Propanil | 66641 | 5 | 8 | 2.3792 | 2.3928 | 100.5706 |  | 7 | 87.5 | 11.6922 | 5.352 | 7.1735
|  |  | 10 | 8 | 7.3332 | 3.9031 | 53.2258 |  | 8 | 100 |  |  | 11.7016
|  |  | 25 | 8 | 20.7799 | 3.8014 | 18.2935 | 11.3965 | 8 | 100 |  |  | 11.3965 | 11.3965 | 11.3965
|  |  | 50 | 8 | 47.1671 | 8.5715 | 18.1726 |  | 8 | 100 |  |  | 25.6973 | 25.6973
|  |  | 100 | 8 | 99.7382 | 14.1046 | 14.1417 |  | 8 | 100 |  |  | 42.2857 | 42.2857
A | Propargite | 68677 | 5 | 8 | 4.3812 | 0.34003 | 7.7612 | 1.0194 | 8 | 100 | 0.8994 | 0.792 | 1.0194 | 1.0194 | 1.0194
|  |  | 10 | 8 | 8.2054 | 0.65529 | 7.9861 |  | 8 | 100 |  |  | 1.9646 | 1.9646
|  |  | 25 | 8 | 23.3313 | 0.88137 | 3.7776 |  | 8 | 100 |  |  | 2.6423 | 2.6423
|  |  | 50 | 8 | 41.8558 | 2.4529 | 5.8603 |  | 8 | 100 |  |  | 7.3537 | 7.3537
|  |  | 100 | 8 | 97.1502 | 4.1115 | 4.2322 |  | 8 | 100 |  |  | 12.3264 | 12.3264
A | Propazine | 68678 | 5 | 8 | 4.1093 | 0.78664 | 19.1428 | 2.3584 | 8 | 100 | 2.3984 | 1.632 | 2.3584 | 2.3584 | 2.3584
|  |  | 10 | 8 | 8.6898 | 1.2158 | 13.9912 |  | 8 | 100 |  |  | 3.645 | 3.645
|  |  | 25 | 8 | 24.3537 | 2.4489 | 10.0554 |  | 8 | 100 |  |  | 7.3417 | 7.3417
|  |  | 50 | 8 | 42.0497 | 3.5295 | 8.3937 |  | 8 | 100 |  |  | 10.5815 | 10.5815
|  |  | 100 | 8 | 94.6487 | 6.4398 | 6.8039 |  | 8 | 100 |  |  | 19.3064 | 19.3064
A | Propiconazole | 66643 | 5 | 8 | 4.0951 | 1.2015 | 29.3402 | 3.6021 | 8 | 100 | 3.5976 | 2.856 | 3.6021 | 3.6021 | 3.6021
|  |  | 10 | 8 | 8.0406 | 1.3425 | 16.6962 |  | 8 | 100 |  |  | 4.0248 | 4.0248
|  |  | 25 | 8 | 23.2323 | 2.7775 | 11.9555 |  | 8 | 100 |  |  | 8.327 | 8.327
|  |  | 50 | 8 | 44.204 | 5.9741 | 13.5149 |  | 8 | 100 |  |  | 17.9105 | 17.9105
|  |  | 100 | 8 | 89.246 | 5.5107 | 6.1748 |  | 8 | 100 |  |  | 16.5212 | 16.5212
A | Propoxur | 68679 | 5 | 8 | 4.0583 | 0.863 | 21.265 | 2.5873 | 8 | 100 | 2.6982 | 1.656 | 2.5873 | 2.5873 | 2.5873
|  |  | 10 | 8 | 7.7424 | 0.85316 | 11.0193 |  | 8 | 100 |  |  | 2.5578 | 2.5578
|  |  | 25 | 8 | 21.8227 | 2.2865 | 10.4775 |  | 8 | 100 |  |  | 6.8548 | 6.8548
|  |  | 50 | 8 | 37.6263 | 4.0115 | 10.6613 |  | 8 | 100 |  |  | 12.0264 | 12.0264
|  |  | 100 | 8 | 85.3928 | 8.4616 | 9.909 |  | 8 | 100 |  |  | 25.3679 | 25.3679
9031 | Prosulfuron | 61687 | 5 | 8 | 5.179 | 1.3801 | 26.6469 | 4.1374 | 8 | 100 | 4.1972 | 5 | 4.1374 | 4.1374 | 4.1374
|  |  | 10 | 8 | 9.2349 | 4.7955 | 51.9281 |  | 7 | 87.5 |  |  | 14.3769
|  |  | 25 | 8 | 28.2371 | 3.8317 | 13.5698 | 11.4875 | 8 | 100 |  |  | 11.4875 | 11.4875 | 11.4875
|  |  | 50 | 8 | 49.1408 | 5.5163 | 11.2255 |  | 8 | 100 |  |  | 16.5379 | 16.5379
|  |  | 100 | 8 | 93.8189 | 12.9247 | 13.7762 |  | 8 | 100 |  |  | 38.7482 | 38.7482
A | Pymetrozine | 68680 | 5 | 8 | 3.5137 | 0.60252 | 17.1478 | 1.8064 | 8 | 100 | 1.7988 | 0.936 | 1.8064 | 1.8064 | 1.8064
|  |  | 10 | 8 | 8.8516 | 0.7593 | 8.5781 |  | 8 | 100 |  |  | 2.2764 | 2.2764
|  |  | 25 | 8 | 24.7155 | 1.8095 | 7.3213 |  | 8 | 100 |  |  | 5.4249 | 5.4249
|  |  | 50 | 8 | 44.5788 | 2.1219 | 4.7599 |  | 8 | 100 |  |  | 6.3615 | 6.3615
|  |  | 100 | 8 | 99.196 | 4.3712 | 4.4066 |  | 8 | 100 |  |  | 13.1049 | 13.1049
A | Pyraclostrobin | 66646 | 5 | 8 | 4.5997 | 0.62967 | 13.6895 | 1.8878 | 8 | 100 | 1.7988 | 1.128 | 1.8878 | 1.8878 | 1.8878
|  |  | 10 | 8 | 8.714 | 1.0426 | 11.965 |  | 8 | 100 |  |  | 3.1258 | 3.1258
|  |  | 25 | 8 | 23.3022 | 1.2835 | 5.5078 |  | 8 | 100 |  |  | 3.8478 | 3.8478
|  |  | 50 | 8 | 42.1455 | 2.7627 | 6.5551 |  | 8 | 100 |  |  | 8.2826 | 8.2826
|  |  | 100 | 8 | 92.8511 | 5.8827 | 6.3356 |  | 8 | 100 |  |  | 17.6362 | 17.6362
A | Pyridaben | 68682 | 5 | 8 | 4.3089 | 0.41839 | 9.7099 | 1.2543 | 8 | 100 | 1.1992 | 1.08 | 1.2543 | 1.2543 | 1.2543
|  |  | 10 | 8 | 8.2682 | 0.51988 | 6.2878 |  | 8 | 100 |  |  | 1.5586 | 1.5586
|  |  | 25 | 8 | 22.8347 | 0.72106 | 3.1577 |  | 8 | 100 |  |  | 2.1617 | 2.1617
|  |  | 50 | 8 | 43.4575 | 1.7913 | 4.122 |  | 8 | 100 |  |  | 5.3704 | 5.3704
|  |  | 100 | 8 | 98.9335 | 2.6819 | 2.7108 |  | 8 | 100 |  |  | 8.0404 | 8.0404
A | Pyrimidinol | 68505 | 5 | 8 | 4.9129 | 1.5553 | 31.6583 | 4.6629 | 8 | 100 | 4.7968 | 9.816 | 4.6629 | 4.6629 | 4.6629
|  |  | 10 | 8 | 8.8685 | 1.7507 | 19.7411 |  | 8 | 100 |  |  | 5.2487 | 5.2487
|  |  | 25 | 8 | 26.5006 | 2.5515 | 9.6281 |  | 8 | 100 |  |  | 7.6494 | 7.6494
|  |  | 50 | 8 | 46.1373 | 3.3574 | 7.277 |  | 8 | 100 |  |  | 10.0655 | 10.0655
|  |  | 100 | 8 | 102.2127 | 6.4236 | 6.2846 |  | 8 | 100 |  |  | 19.2581 | 19.2581
A | Pyriproxyfen | 68683 | 5 | 8 | 4.5124 | 0.5744 | 12.7294 | 1.722 | 8 | 100 | 1.7988 | 1.248 | 1.722 | 1.722 | 1.722
|  |  | 10 | 8 | 8.6613 | 0.51478 | 5.9434 |  | 8 | 100 |  |  | 1.5433 | 1.5433
|  |  | 25 | 8 | 23.8954 | 0.90855 | 3.8022 |  | 8 | 100 |  |  | 2.7238 | 2.7238
|  |  | 50 | 8 | 44.0704 | 2.0217 | 4.5875 |  | 8 | 100 |  |  | 6.0611 | 6.0611
|  |  | 100 | 8 | 99.5226 | 4.1976 | 4.2177 |  | 8 | 100 |  |  | 12.5844 | 12.5844
A | Siduron | 68686 | 5 | 8 | 7.1014 | 1.5263 | 21.4926 | 4.5758 | 8 | 100 | 4.497 | 2.28 | 4.5758 | 4.5758 | 4.5758
|  |  | 10 | 8 | 12.5201 | 1.3224 | 10.5618 |  | 8 | 100 |  |  | 3.9644 | 3.9644
|  |  | 25 | 8 | 26.4104 | 3.2088 | 12.1499 |  | 8 | 100 |  |  | 9.6201 | 9.6201
|  |  | 50 | 8 | 44.5762 | 4.1706 | 9.3561 |  | 8 | 100 |  |  | 12.5035 | 12.5035
|  |  | 100 | 8 | 102.7357 | 7.5426 | 7.3418 |  | 8 | 100 |  |  | 22.6128 | 22.6128
A | Simazine | 65105 | 5 | 8 | 6.9562 | 2.9279 | 42.0896 |  | 8 | 100 | 10.7928 | 4.632 | 8.7777
|  |  | 10 | 8 | 11.3392 | 3.5644 | 31.4344 | 10.6861 | 8 | 100 |  |  | 10.6861 | 10.6861 | 10.6861
|  |  | 25 | 8 | 29.487 | 7.2803 | 24.6898 |  | 8 | 100 |  |  | 21.8263 | 21.8263
|  |  | 50 | 8 | 50.2071 | 9.4318 | 18.7858 |  | 8 | 100 |  |  | 28.2765 | 28.2765
|  |  | 100 | 8 | 104.6794 | 12.382 | 11.8285 |  | 8 | 100 |  |  | 37.1213 | 37.1213
A | Sulfometuron-methyl | 68688 | 5 | 8 | 3.9732 | 0.61568 | 15.4956 | 1.8458 | 8 | 100 | 1.7988 | 1.824 | 1.8458 | 1.8458 | 1.8458
|  |  | 10 | 8 | 9.3052 | 0.64013 | 6.8793 |  | 8 | 100 |  |  | 1.9191 | 1.9191
|  |  | 25 | 8 | 24.4627 | 1.9467 | 7.9577 |  | 8 | 100 |  |  | 5.8361 | 5.8361
|  |  | 50 | 8 | 43.8674 | 2.6295 | 5.9943 |  | 8 | 100 |  |  | 7.8833 | 7.8833
|  |  | 100 | 8 | 100.1452 | 5.9876 | 5.9789 |  | 8 | 100 |  |  | 17.9507 | 17.9507
A | Sulfosulfuron | 68689 | 5 | 8 | 5.321 | 1.9292 | 36.2554 |  | 8 | 100 | 5.6962 | 4.392 | 5.7836
|  |  | 10 | 8 | 11.5587 | 2.7825 | 24.0732 | 8.342 | 8 | 100 |  |  | 8.342 | 8.342 | 8.342
|  |  | 25 | 8 | 24.8667 | 4.5723 | 18.3872 |  | 8 | 100 |  |  | 13.7078 | 13.7078
|  |  | 50 | 8 | 43.986 | 5.4778 | 12.4536 |  | 8 | 100 |  |  | 16.4226 | 16.4226
|  |  | 100 | 8 | 95.7537 | 10.9602 | 11.4463 |  | 8 | 100 |  |  | 32.8588 | 32.8588
A | Sulfosulfuron ethyl sulfone | 68690 | 5 | 8 | 4.6915 | 0.69334 | 14.7785 | 2.0786 | 8 | 100 | 2.0986 | 1.08 | 2.0786 | 2.0786 | 2.0786
|  |  | 10 | 8 | 9.9105 | 0.61192 | 6.1745 |  | 8 | 100 |  |  | 1.8345 | 1.8345
|  |  | 25 | 8 | 24.5144 | 1.644 | 6.7064 |  | 8 | 100 |  |  | 4.9288 | 4.9288
|  |  | 50 | 8 | 46.897 | 2.7457 | 5.8547 |  | 8 | 100 |  |  | 8.2316 | 8.2316
|  |  | 100 | 8 | 103.8131 | 7.9275 | 7.6364 |  | 8 | 100 |  |  | 23.7668 | 23.7668
A | tebuconazole | 66649 | 5 | 8 | 4.5505 | 0.69406 | 15.2524 | 2.0808 | 8 | 100 | 2.0986 | 2.472 | 2.0808 | 2.0808 | 2.0808
|  |  | 10 | 8 | 9.1573 | 0.86863 | 9.4857 |  | 8 | 100 |  |  | 2.6042 | 2.6042
|  |  | 25 | 8 | 23.6662 | 2.0529 | 8.6743 |  | 8 | 100 |  |  | 6.1545 | 6.1545
|  |  | 50 | 8 | 43.9596 | 4.1985 | 9.5509 |  | 8 | 100 |  |  | 12.5872 | 12.5872
|  |  | 100 | 8 | 96.9276 | 3.5033 | 3.6144 |  | 8 | 100 |  |  | 10.503 | 10.503
A | Tebufenozide | 68692 | 5 | 8 | 4.5518 | 0.2822 | 6.1998 | 0.84604 | 8 | 100 | 0.8994 | 0.792 | 0.84604 | 0.84604 | 0.84604
|  |  | 10 | 8 | 8.684 | 0.42875 | 4.9373 |  | 8 | 100 |  |  | 1.2854 | 1.2854
|  |  | 25 | 8 | 23.6281 | 1.0751 | 4.5503 |  | 8 | 100 |  |  | 3.2233 | 3.2233
|  |  | 50 | 8 | 43.5618 | 2.0847 | 4.7856 |  | 8 | 100 |  |  | 6.25 | 6.25
|  |  | 100 | 8 | 94.1245 | 4.4036 | 4.6785 |  | 8 | 100 |  |  | 13.2019 | 13.2019
A | Tebupirimfos | 68693 | 5 | 8 | 4.5642 | 0.3068 | 6.7218 | 0.91977 | 8 | 100 | 0.8994 | 0.84 | 0.91977 | 0.91977 | 0.91977
|  |  | 10 | 8 | 9.0019 | 0.51166 | 5.6839 |  | 8 | 100 |  |  | 1.534 | 1.534
|  |  | 25 | 8 | 25.1117 | 1.4939 | 5.9489 |  | 8 | 100 |  |  | 4.4786 | 4.4786
|  |  | 50 | 8 | 46.0408 | 2.5712 | 5.5846 |  | 8 | 100 |  |  | 7.7084 | 7.7084
|  |  | 100 | 8 | 101.974 | 5.4979 | 5.3915 |  | 8 | 100 |  |  | 16.4828 | 16.4828
A | Tebupirimfos oxon | 68694 | 5 | 8 | 4.3571 | 0.31248 | 7.1718 | 0.93682 | 8 | 100 | 0.8994 | 0.768 | 0.93682 | 0.93682 | 0.93682
|  |  | 10 | 8 | 9.4222 | 0.48295 | 5.1256 |  | 8 | 100 |  |  | 1.4479 | 1.4479
|  |  | 25 | 8 | 24.9958 | 1.884 | 7.5371 |  | 8 | 100 |  |  | 5.6482 | 5.6482
|  |  | 50 | 8 | 46.9778 | 2.5129 | 5.349 |  | 8 | 100 |  |  | 7.5336 | 7.5336
|  |  | 100 | 8 | 102.6359 | 6.593 | 6.4237 |  | 8 | 100 |  |  | 19.7659 | 19.7659
A | Tebuthiuron | 68695 | 5 | 8 | 3.7171 | 0.86712 | 23.3276 | 2.5996 | 8 | 100 | 2.6982 | 1.44 | 2.5996 | 2.5996 | 2.5996
|  |  | 10 | 8 | 9.2737 | 0.89424 | 9.6427 |  | 8 | 100 |  |  | 2.6809 | 2.6809
|  |  | 25 | 8 | 23.4926 | 2.2838 | 9.7215 |  | 8 | 100 |  |  | 6.8469 | 6.8469
|  |  | 50 | 8 | 47.5164 | 5.3713 | 11.3041 |  | 8 | 100 |  |  | 16.1032 | 16.1032
|  |  | 100 | 8 | 103.6592 | 7.2812 | 7.0241 |  | 8 | 100 |  |  | 21.8289 | 21.8289
A | Tebuthiuron TP 104 | 68575 | 5 | 8 | 3.9521 | 0.51021 | 12.9098 | 1.5296 | 8 | 100 | 1.499 | 2.736 | 1.5296 | 1.5296 | 1.5296
|  |  | 10 | 8 | 9.4228 | 2.1547 | 22.8668 |  | 8 | 100 |  |  | 6.4598 | 6.4598
|  |  | 25 | 8 | 23.6861 | 2.5829 | 10.9046 |  | 8 | 100 |  |  | 7.7435 | 7.7435
|  |  | 50 | 8 | 47.7505 | 5.8212 | 12.1909 |  | 8 | 100 |  |  | 17.452 | 17.452
|  |  | 100 | 8 | 95.1049 | 7.0524 | 7.4154 |  | 8 | 100 |  |  | 21.1432 | 21.1432
A | Tebuthiuron TP 106 | 68714 | 5 | 7 | 4.709 | 5.9059 | 125.4175 |  | 3 | 42.8571 | 37.7748 | 38 | 17.7059
|  |  | 10 | 8 | 11.0274 | 6.2649 | 56.8121 |  | 7 | 87.5 |  |  | 18.7822
|  |  | 25 | 8 | 33.1672 | 11.0643 | 33.3591 |  | 8 | 100 |  |  | 33.1707
|  |  | 50 | 8 | 54.9978 | 12.5713 | 22.8579 | 37.6889 | 8 | 100 |  |  | 37.6889 | 37.6889 | 37.6889
|  |  | 100 | 8 | 104.6819 | 16.9001 | 16.1442 |  | 8 | 100 |  |  | 50.6665 | 50.6665
A | Tebuthiuron TP 108 | 68696 | 5 | 8 | 4.9025 | 1.0703 | 21.8321 | 3.2088 | 8 | 100 | 3.2978 | 4.536 | 3.2088 | 3.2088 | 3.2088
|  |  | 10 | 8 | 9.2987 | 1.4168 | 15.2367 |  | 8 | 100 |  |  | 4.2476 | 4.2476
|  |  | 25 | 8 | 25.437 | 2.658 | 10.4495 |  | 8 | 100 |  |  | 7.9688 | 7.9688
|  |  | 50 | 8 | 46.0187 | 6.1554 | 13.376 |  | 8 | 100 |  |  | 18.454 | 18.454
|  |  | 100 | 8 | 98.0612 | 8.6824 | 8.854 |  | 8 | 100 |  |  | 26.0297 | 26.0297
A | Tebuthiuron TP 109 (OH) | 68697 | 5 | 7 | 1.8918 | 1.2541 | 66.2926 | 3.7599 | 6 | 85.7143 | 18.5876 | 8.328 | 3.7599 | 3.7599 | 3.7599
|  |  | 10 | 8 | 8.1053 | 4.8781 | 60.1844 |  | 8 | 100 |  |  | 14.6245
|  |  | 25 | 8 | 21.9087 | 6.2437 | 28.4988 | 18.7187 | 8 | 100 |  |  | 18.7187 | 18.7187 | 18.7187
|  |  | 50 | 8 | 41.4774 | 9.7311 | 23.4612 |  | 8 | 100 |  |  | 29.1738 | 29.1738
|  |  | 100 | 8 | 97.095 | 21.4302 | 22.0713 |  | 8 | 100 |  |  | 64.2476 | 64.2476
A | Terbacil | 68698 | 25 | 8 | 32.8218 | 10.8278 | 32.9896 |  | 8 | 100 | 15.8894 | 10.44 | 32.4617
|  |  | 50 | 8 | 50.1103 | 5.2528 | 10.4825 | 15.7479 | 8 | 100 |  |  | 15.7479 | 15.7479 | 15.7479
|  |  | 100 | 8 | 105.12 | 16.6321 | 15.822 |  | 8 | 100 |  |  | 49.8629 | 49.8629
A | Terbufos | 68699 | 5 | 8 | 4.1763 | 0.60407 | 14.4642 | 1.811 | 8 | 100 | 1.7988 | 1.632 | 1.811 | 1.811 | 1.811
|  |  | 10 | 8 | 8.5473 | 0.93855 | 10.9807 |  | 8 | 100 |  |  | 2.8138 | 2.8138
|  |  | 25 | 8 | 21.5978 | 1.1495 | 5.3223 |  | 8 | 100 |  |  | 3.4462 | 3.4462
|  |  | 50 | 8 | 39.6667 | 3.9359 | 9.9225 |  | 8 | 100 |  |  | 11.7999 | 11.7999
|  |  | 100 | 8 | 91.1834 | 8.66 | 9.4974 |  | 8 | 100 |  |  | 25.9628 | 25.9628
A | Terbufos oxon | 68700 | 5 | 8 | 4.3548 | 0.32647 | 7.4968 | 0.97875 | 8 | 100 | 0.8994 | 1.512 | 0.97875 | 0.97875 | 0.97875
|  |  | 10 | 8 | 8.4221 | 0.81033 | 9.6215 |  | 8 | 100 |  |  | 2.4294 | 2.4294
|  |  | 25 | 8 | 22.5745 | 2.1112 | 9.3523 |  | 8 | 100 |  |  | 6.3295 | 6.3295
|  |  | 50 | 8 | 42.2054 | 3.6674 | 8.6893 |  | 8 | 100 |  |  | 10.9948 | 10.9948
|  |  | 100 | 8 | 93.943 | 7.7664 | 8.2671 |  | 8 | 100 |  |  | 23.2836 | 23.2836
A | Terbufos oxon sulfone | 68701 | 5 | 8 | 5.0442 | 1.6379 | 32.4722 | 4.9106 | 8 | 100 | 4.7968 | 5.016 | 4.9106 | 4.9106 | 4.9106
|  |  | 10 | 8 | 10.1208 | 2.6107 | 25.7957 |  | 8 | 100 |  |  | 7.827 | 7.827
|  |  | 25 | 8 | 25.649 | 3.8544 | 15.0274 |  | 8 | 100 |  |  | 11.5555 | 11.5555
|  |  | 50 | 8 | 45.3722 | 6.6275 | 14.607 |  | 8 | 100 |  |  | 19.8694 | 19.8694
|  |  | 100 | 8 | 99.3371 | 10.156 | 10.2238 |  | 8 | 100 |  |  | 30.4478 | 30.4478
A | Terbufos oxon sulfoxide | 68702 | 5 | 8 | 4.5357 | 1.0232 | 22.5587 | 3.0675 | 8 | 100 | 2.998 | 1.872 | 3.0675 | 3.0675 | 3.0675
|  |  | 10 | 8 | 9.356 | 1.1546 | 12.3404 |  | 8 | 100 |  |  | 3.4614 | 3.4614
|  |  | 25 | 8 | 25.1845 | 1.795 | 7.1275 |  | 8 | 100 |  |  | 5.3815 | 5.3815
|  |  | 50 | 8 | 48.4993 | 3.6852 | 7.5984 |  | 8 | 100 |  |  | 11.0481 | 11.0481
|  |  | 100 | 8 | 101.7219 | 6.2788 | 6.1725 |  | 8 | 100 |  |  | 18.8237 | 18.8237
A | Terbufos sulfone | 68703 | 5 | 8 | 4.4541 | 3.2728 | 73.478 |  | 7 | 87.5 | 7.1952 | 5.352 | 9.8118
|  |  | 10 | 8 | 9.7136 | 2.4192 | 24.9048 | 7.2526 | 8 | 100 |  |  | 7.2526 | 7.2526 | 7.2526
|  |  | 25 | 8 | 29.5112 | 6.4708 | 21.9266 |  | 8 | 100 |  |  | 19.3994 | 19.3994
|  |  | 50 | 8 | 49.1587 | 7.1789 | 14.6034 |  | 8 | 100 |  |  | 21.5222 | 21.5222
|  |  | 100 | 8 | 105.3432 | 17.4584 | 16.5729 |  | 8 | 100 |  |  | 52.3404 | 52.3404
A | Terbufos sulfoxide | 68704 | 5 | 8 | 4.5615 | 0.57734 | 12.6567 | 1.7309 | 8 | 100 | 1.7988 | 1.032 | 1.7309 | 1.7309 | 1.7309
|  |  | 10 | 8 | 8.7093 | 0.41051 | 4.7135 |  | 8 | 100 |  |  | 1.2307 | 1.2307
|  |  | 25 | 8 | 25.5998 | 1.6053 | 6.2706 |  | 8 | 100 |  |  | 4.8126 | 4.8126
|  |  | 50 | 8 | 47.314 | 2.5713 | 5.4346 |  | 8 | 100 |  |  | 7.7089 | 7.7089
|  |  | 100 | 8 | 102.8815 | 4.8204 | 4.6854 |  | 8 | 100 |  |  | 14.4516 | 14.4516
A | Terbuthylazine | 66651 | 5 | 8 | 5.4166 | 0.49049 | 9.0553 | 1.4705 | 8 | 100 | 1.499 | 1.032 | 1.4705 | 1.4705 | 1.4705
|  |  | 10 | 8 | 9.4857 | 1.1865 | 12.5084 |  | 8 | 100 |  |  | 3.5572 | 3.5572
|  |  | 25 | 8 | 24.5705 | 1.6277 | 6.6244 |  | 8 | 100 |  |  | 4.8797 | 4.8797
|  |  | 50 | 8 | 43.9952 | 3.7742 | 8.5787 |  | 8 | 100 |  |  | 11.3151 | 11.3151
|  |  | 100 | 8 | 96.4485 | 6.5996 | 6.8426 |  | 8 | 100 |  |  | 19.7856 | 19.7856
A | Tetraconazole | 66654 | 5 | 8 | 4.7032 | 0.84736 | 18.0165 | 2.5404 | 8 | 100 | 2.3984 | 2.928 | 2.5404 | 2.5404 | 2.5404
|  |  | 10 | 8 | 10.516 | 2.01 | 19.114 |  | 8 | 100 |  |  | 6.0261 | 6.0261
|  |  | 25 | 8 | 23.3782 | 3.1123 | 13.3128 |  | 8 | 100 |  |  | 9.3306 | 9.3306
|  |  | 50 | 8 | 40.2384 | 3.1799 | 7.9028 |  | 8 | 100 |  |  | 9.5335 | 9.5335
|  |  | 100 | 8 | 88.7215 | 4.612 | 5.1983 |  | 8 | 100 |  |  | 13.8268 | 13.8268
A | Thiobencarb | 65107 | 5 | 8 | 4.6658 | 0.50818 | 10.8916 | 1.5235 | 8 | 100 | 1.499 | 0.768 | 1.5235 | 1.5235 | 1.5235
|  |  | 10 | 8 | 8.4717 | 0.66184 | 7.8124 |  | 8 | 100 |  |  | 1.9842 | 1.9842
|  |  | 25 | 8 | 22.9861 | 1.0702 | 4.6559 |  | 8 | 100 |  |  | 3.2085 | 3.2085
|  |  | 50 | 8 | 42.7377 | 1.9331 | 4.5231 |  | 8 | 100 |  |  | 5.7954 | 5.7954
|  |  | 100 | 8 | 96.0519 | 4.7057 | 4.8991 |  | 8 | 100 |  |  | 14.1077 | 14.1077
A | trans-Permethrin | 68708 | 5 | 8 | 4.2417 | 0.75307 | 17.7539 | 2.2577 | 8 | 100 | 2.3984 | 2 | 2.2577 | 2.2577 | 2.2577
|  |  | 10 | 8 | 8.426 | 0.62312 | 7.3952 |  | 8 | 100 |  |  | 1.8681 | 1.8681
|  |  | 25 | 8 | 24.9017 | 1.4925 | 5.9936 |  | 8 | 100 |  |  | 4.4745 | 4.4745
|  |  | 50 | 8 | 44.5577 | 2.1046 | 4.7232 |  | 8 | 100 |  |  | 6.3095 | 6.3095
|  |  | 100 | 8 | 93.713 | 2.6035 | 2.7782 |  | 8 | 100 |  |  | 7.8053 | 7.8053
A | Triallate | 68710 | 5 | 8 | 5.1497 | 0.70819 | 13.7522 | 2.1232 | 8 | 100 | 2.0986 | 7.512 | 2.1232 | 2.1232 | 2.1232
|  |  | 10 | 8 | 9.8818 | 1.9994 | 20.2336 |  | 8 | 100 |  |  | 5.9943 | 5.9943
|  |  | 25 | 8 | 23.8173 | 1.8563 | 7.7939 |  | 8 | 100 |  |  | 5.5652 | 5.5652
|  |  | 50 | 8 | 42.9766 | 3.9848 | 9.2721 |  | 8 | 100 |  |  | 11.9465 | 11.9465
|  |  | 100 | 8 | 96.0864 | 7.6685 | 7.9808 |  | 8 | 100 |  |  | 22.9902 | 22.9902
A | Tribufos | 68711 | 5 | 8 | 4.4673 | 0.36611 | 8.1954 | 1.0976 | 8 | 100 | 1.1992 | 0.816 | 1.0976 | 1.0976 | 1.0976
|  |  | 10 | 8 | 8.9425 | 0.32864 | 3.675 |  | 8 | 100 |  |  | 0.98526 | 0.98526
|  |  | 25 | 8 | 24.3144 | 0.74961 | 3.083 |  | 8 | 100 |  |  | 2.2473 | 2.2473
|  |  | 50 | 8 | 45.7403 | 2.5989 | 5.6819 |  | 8 | 100 |  |  | 7.7916 | 7.7916
|  |  | 100 | 8 | 101.4855 | 5.462 | 5.382 |  | 8 | 100 |  |  | 16.375 | 16.375
A | Trifloxystrobin | 66660 | 5 | 8 | 4.4117 | 0.28771 | 6.5216 | 0.86256 | 8 | 100 | 0.8994 | 0.84 | 0.86256 | 0.86256 | 0.86256
|  |  | 10 | 8 | 8.8152 | 0.44214 | 5.0157 |  | 8 | 100 |  |  | 1.3255 | 1.3255
|  |  | 25 | 8 | 23.8534 | 0.95666 | 4.0106 |  | 8 | 100 |  |  | 2.8681 | 2.8681
|  |  | 50 | 8 | 43.8429 | 2.3571 | 5.3763 |  | 8 | 100 |  |  | 7.0666 | 7.0666
|  |  | 100 | 8 | 96.7706 | 4.7934 | 4.9534 |  | 8 | 100 |  |  | 14.3706 | 14.3706
Pesticide compounds deleted(D) or moved from ESI+ mode analytical method to ESI- mode method (lab code 9031) | [12 empty] | 0
D | 3-Ketocarbofuran | 68509 | 100 | 8 | 76.7306 | 29.2161 | 38.0762 | 87.5899 | 8 | 100 | 87.5416 |  | 87.5899 | 87.5899 | 87.5899
D | 7-Hydroxycarbofuran | 68518 | 100 | 8 | 150.4112 | 73.2869 | 48.7244 |  | 8 | 100 |  | 250 | 219.7142
D | Aminopyralid | 68534 | 100 | 8 | 32.0923 | 22.8209 | 71.1102 | 68.4171 | 6 | 75 |  |  | 68.4171 | 68.4171 | 68.4171
D | Bifenazate | 68540 | 5 | 8 | 3.8312 | 0.76214 | 19.8931 | 2.2849 | 8 | 100 | 2.3984 |  | 2.2849 | 2.2849 | 2.2849
|  |  | 10 | 8 | 4.5469 | 1.4012 | 30.8175 |  | 8 | 100 |  |  | 4.2009 | 4.2009
|  |  | 25 | 8 | 17.3751 | 4.9067 | 28.2396 |  | 8 | 100 |  |  | 14.7102 | 14.7102
|  |  | 50 | 8 | 27.0629 | 6.3682 | 23.531 |  | 8 | 100 |  |  | 19.0918 | 19.0918
|  |  | 100 | 8 | 64.2473 | 14.0999 | 21.9463 |  | 8 | 100 |  |  | 42.2715 | 42.2715
D | Bifenazate diazene | 68541 | 5 | 8 | 4.7745 | 0.90495 | 18.9539 | 2.713 | 8 | 100 | 2.6982 |  | 2.713 | 2.713 | 2.713
|  |  | 10 | 8 | 12.3203 | 1.2394 | 10.0594 |  | 8 | 100 |  |  | 3.7156 | 3.7156
|  |  | 25 | 8 | 24.5227 | 5.8529 | 23.8671 |  | 8 | 100 |  |  | 17.5469 | 17.5469
|  |  | 50 | 8 | 54.7326 | 7.7882 | 14.2295 |  | 8 | 100 |  |  | 23.3489 | 23.3489
|  |  | 100 | 8 | 103.1302 | 37.6011 | 36.4598 |  | 8 | 100 |  |  | 112.7282
9031 | Dimethenamid SAA | 68583 | 100 | 8 | 98.4757 | 28.8326 | 29.2789 | 86.4401 | 8 | 100 | 86.3424 | 94.464 | 86.4401 | 86.4401 | 86.4401
9031 | Famoxadone | 67609 | 25 | 8 | 28.6309 | 10.0204 | 34.9985 |  | 8 | 100 | 29.98 | 30 | 30.0411
|  |  | 50 | 8 | 49.1131 | 13.2513 | 26.9812 | 39.7274 | 8 | 100 |  |  | 39.7274 | 39.7274 | 39.7274
|  |  | 100 | 8 | 100.423 | 16.3706 | 16.3017 |  | 8 | 100 |  |  | 49.0792 | 49.0792
D | Formetanate | 68609 | 5 | 8 | 3.9371 | 0.39109 | 9.9335 | 1.1725 | 8 | 100 | 1.1992 |  | 1.1725 | 1.1725 | 1.1725
|  |  | 10 | 8 | 7.9113 | 0.89736 | 11.3428 |  | 8 | 100 |  |  | 2.6903 | 2.6903
|  |  | 25 | 8 | 19.8694 | 1.8244 | 9.1819 |  | 8 | 100 |  |  | 5.4695 | 5.4695
|  |  | 50 | 8 | 39.9957 | 2.0361 | 5.0907 |  | 8 | 100 |  |  | 6.1041 | 6.1041
|  |  | 100 | 8 | 82.8161 | 10.6329 | 12.8391 |  | 8 | 100 |  |  | 31.8773 | 31.8773
9031 | Metribuzin DK | 68653 | 5 | 8 | 1812.3346 | 2270.2709 | 125.2678 |  | 8 | 100 |  | 100 | 6806.2723
|  |  | 10 | 7 | 1840.4954 | 2685.3527 | 145.9038 |  | 7 | 100 |  |  | 8050.6875
|  |  | 50 | 7 | 2132.9966 | 2389.1362 | 112.0084 |  | 7 | 100 |  |  | 7162.6304
|  |  | 100 | 7 | 1510.8447 | 2249.8393 | 148.9127 |  | 7 | 100 |  |  | 6745.0181
D | Phosmet oxon | 68674 | 5 | 8 | 3.1779 | 1.302 | 40.9704 | 3.9033 | 8 | 100 | 3.8974 |  | 3.9033 | 3.9033 | 3.9033
|  |  | 10 | 8 | 8.0033 | 1.0422 | 13.0224 |  | 8 | 100 |  |  | 3.1246 | 3.1246
|  |  | 25 | 7 | 9.7021 | 6.3081 | 65.0175 |  | 7 | 100 |  |  | 18.9116 | 18.9116
|  |  | 50 | 8 | 33.7478 | 5.7125 | 16.9269 |  | 8 | 100 |  |  | 17.126 | 17.126
|  |  | 100 | 8 | 42.7134 | 42.4348 | 99.3478 |  | 7 | 87.5 |  |  | 127.2195
E | Oxyfluorfen | 65093LCM60 | 100 | 8 | 3.7 | 10.3 | 282.8 | 30.8794
E | Parathion-methyl | 65089LCM64 | 100 | 8 | 79.8 | 115.5 | 39.3 | 346.269
Note: oxyfluorfen and parathion-methyl were not inlcuded in this summary and need to be redone.

## T20_esineg_mdl_summary
Table 20. Estimated method detection levels for pesticide compounds determined by direct aqueous-injection liquid chromatography-tandem mass spectrometry in negative electrospray ionization mode according to U.S. Environmental Protection Agency procedure from replicate analysis of 5, 10, 50, and 100 nanograms per liter standards in seven consecutive analytical batches [Method detection level (MDL) is calculated as standard deviation times t, where t is student's t value for 8 replicates (t= 2.998); valid estimated MDLs must be greater than nominal spike concentration and have a detection frequency greater than 50 percent; MDL selected was the lowest valid MDL determined from the nominal concentration levels. Validation quality code: A, no qualification; E, estimated remark code; D, deleted from method in May 2013; USGS, U.S. Geological Survey; N, number of replicates; Std. Dev., standard deviation; RSD, relative standard deviation; min, minimum] | [13 empty] | These columns used for selecting valid MDL
[5 empty] | Concentration (ng/L) |  |  |  | Detections
Validation quality code | Pesticide compound | USGS Parameter code | Nominal concentration (ng/L) | N | Mean (ng/L) | Std Dev (ng/L) | RSD (ng/L) | MDL (ng/L) | N | Sum | Detection frequency (%) | Valid MDL | Selected MDL | MDL (ng/L) | Valid MDL (ng/L)
A | 2,3,3-trichloroprop-2-ene-1-SA | 68691 | 5 | 49 | 2 | 3.05 | 152.5 |  | 49 | 18 | 36.7347 |  |  | 7.32
| Triallate-SA |  | 10 | 57 | 9.5 | 5.58 | 58.7368 |  | 57 | 49 | 85.9649 |  |  | 13.392
|  |  | 50 | 48 | 47.4 | 11.08 | 23.3755 | 26.592 | 48 | 48 | 100 | 26.592 | 26.592 | 26.592 | 26.592
|  |  | 100 | 49 | 107.9 | 19.37 | 17.9518 | 46.488 | 49 | 49 | 100 | 46.488 |  | 46.488 | 46.488
A | 2,4-D | 68500 | 5 | 49 | 3 | 5.85 | 195 |  | 49 | 17 | 34.6939 |  |  | 14.04
|  |  | 10 | 57 | 12 | 8.6 | 71.6667 |  | 57 | 45 | 78.9474 |  |  | 20.64
|  |  | 50 | 48 | 52 | 12.82 | 24.6538 | 30.768 | 48 | 48 | 100 | 30.768 | 30.768 | 30.768 | 30.768
|  |  | 100 | 49 | 111.2 | 22.21 | 19.973 | 53.304 | 49 | 49 | 100 | 53.304 |  | 53.304 | 53.304
A | 3-Phenoxybenzoic acid | 68873 | 5 | 48 | 3.4 | 7.39 | 217.3529 |  | 48 | 20 | 41.6667 |  |  | 17.736
|  |  | 10 | 56 | 9.4 | 8.97 | 95.4255 |  | 56 | 38 | 67.8571 |  |  | 21.528
|  |  | 50 | 48 | 44.6 | 12.54 | 28.1166 | 30.096 | 48 | 48 | 100 | 30.096 | 30.096 | 30.096 | 30.096
|  |  | 100 | 49 | 107.1 | 21 | 19.6078 | 50.4 | 49 | 49 | 100 | 50.4 |  | 50.4 | 50.4
A | 4-(Hydroxymethyl)pendimethalin | 68511 | 5 | 49 | 1.7 | 6.51 | 382.9412 |  | 49 | 6 | 12.2449 |  |  | 15.624
|  |  | 10 | 57 | 5.9 | 20.92 | 354.5763 |  | 57 | 7 | 12.2807 |  |  | 50.208
|  |  | 50 | 48 | 36.1 | 29.77 | 82.4654 |  | 48 | 37 | 77.0833 |  |  | 71.448
|  |  | 100 | 49 | 96.6 | 44.29 | 45.8489 | 106.296 | 49 | 45 | 91.8367 | 106.296 | 106.296 | 106.296 | 106.296
A | 4-Hydroxychlorothalonil | 68336 | 5 | 49 | 6.4 | 9.13 | 142.6562 |  | 49 | 25 | 51.0204 |  |  | 21.912
|  |  | 10 | 57 | 14.7 | 13.39 | 91.0884 |  | 57 | 37 | 64.9123 |  |  | 32.136
|  |  | 50 | 48 | 57.1 | 20.34 | 35.6217 | 48.816 | 48 | 48 | 100 | 48.816 | 48.816 | 48.816 | 48.816
|  |  | 100 | 49 | 114.6 | 30.9 | 26.9634 | 74.16 | 49 | 49 | 100 | 74.16 |  | 74.16 | 74.16
A | Acetochlor OA | 68522 | 5 | 49 | 0.6 | 2.55 | 425 |  | 49 | 4 | 8.1633 |  |  | 6.12
|  |  | 10 | 57 | 6.3 | 8.11 | 128.7302 |  | 57 | 30 | 52.6316 |  |  | 19.464
|  |  | 50 | 48 | 46.2 | 18.74 | 40.5628 | 44.976 | 48 | 48 | 100 | 44.976 | 44.976 | 44.976 | 44.976
|  |  | 100 | 49 | 107.6 | 25.11 | 23.3364 | 60.264 | 49 | 49 | 100 | 60.264 |  | 60.264 | 60.264
A | Acetochlor SA | 68523 | 5 | 49 | 5.6 | 2.9 | 51.7857 |  | 49 | 46 | 93.8776 |  |  | 6.96
|  |  | 10 | 57 | 11.8 | 5.4 | 45.7627 |  | 57 | 53 | 92.9825 |  |  | 12.96
|  |  | 50 | 48 | 54.4 | 14.88 | 27.3529 | 35.712 | 48 | 48 | 100 | 35.712 | 35.712 | 35.712 | 35.712
|  |  | 100 | 49 | 114 | 23.63 | 20.7281 | 56.712 | 49 | 49 | 100 | 56.712 |  | 56.712 | 56.712
A | Acetochlor SAA | 68524 | 5 | 49 | 3.5 | 5.56 | 158.8571 |  | 49 | 19 | 38.7755 |  |  | 13.344
|  |  | 10 | 57 | 11.4 | 8.98 | 78.7719 |  | 57 | 44 | 77.193 |  |  | 21.552
|  |  | 50 | 48 | 48.6 | 23.82 | 49.0123 |  | 48 | 48 | 100 |  |  | 57.168
|  |  | 100 | 49 | 112 | 36.56 | 32.6429 | 87.744 | 49 | 49 | 100 | 87.744 | 87.744 | 87.744 | 87.744
A | Alachlor OA | 68526 | 5 | 49 | 0.8 | 2.03 | 253.75 |  | 49 | 9 | 18.3673 |  |  | 4.872
|  |  | 10 | 57 | 8.6 | 8.29 | 96.3953 |  | 57 | 41 | 71.9298 |  |  | 19.896
|  |  | 50 | 48 | 51.3 | 17.38 | 33.8791 | 41.712 | 48 | 48 | 100 | 41.712 | 41.712 | 41.712 | 41.712
|  |  | 100 | 49 | 110.2 | 25.27 | 22.931 | 60.648 | 49 | 49 | 100 | 60.648 |  | 60.648 | 60.648
E | Alachlor SA | 68871 | 5 | 49 | 4.5 | 2.8 | 62.2222 |  | 49 | 44 | 89.7959 |  |  | 6.72
|  |  | 10 | 57 | 11.2 | 5.25 | 46.875 |  | 57 | 53 | 92.9825 |  |  | 12.6
|  |  | 50 | 48 | 53.6 | 14.79 | 27.5933 | 35.496 | 48 | 48 | 100 | 35.496 | 35.496 | 35.496 | 35.496
|  |  | 100 | 49 | 113.3 | 23.29 | 20.556 | 55.896 | 49 | 49 | 100 | 55.896 |  | 55.896 | 55.896
A | Alachlor SAA | 68527 | 5 | 49 | 3.6 | 7.63 | 211.9444 |  | 49 | 14 | 28.5714 |  |  | 18.312
|  |  | 10 | 56 | 11.2 | 12.77 | 114.0179 |  | 56 | 35 | 62.5 |  |  | 30.648
|  |  | 50 | 48 | 60 | 24.74 | 41.2333 |  | 48 | 48 | 100 |  |  | 59.376
|  |  | 100 | 49 | 120.5 | 35.17 | 29.1867 | 84.408 | 49 | 49 | 100 | 84.408 | 84.408 | 84.408 | 84.408
D | Ammelide | 68535 | 5 | 48 | 2.1 | 4.06 | 193.3333 |  | 48 | 17 | 35.4167 |  |  | 9.744
|  |  | 10 | 57 | 11.7 | 9.15 | 78.2051 |  | 57 | 45 | 78.9474 |  |  | 21.96
|  |  | 50 | 48 | 43.6 | 15.52 | 35.5963 | 37.248 | 48 | 48 | 100 | 37.248 | 37.248 | 37.248 | 37.248
|  |  | 100 | 49 | 121.5 | 39.09 | 32.1728 | 93.816 | 49 | 49 | 100 | 93.816 |  | 93.816 | 93.816
A | Bentazone | 68538 | 5 | 49 | 5.3 | 1.5 | 28.3019 | 3.6 | 49 | 49 | 100 | 3.6 | 3.6 | 3.6 | 3.6
|  |  | 10 | 57 | 11.4 | 4.61 | 40.4386 |  | 57 | 53 | 92.9825 |  |  | 11.064
|  |  | 50 | 48 | 49.2 | 11.23 | 22.8252 | 26.952 | 48 | 48 | 100 | 26.952 |  | 26.952 | 26.952
|  |  | 100 | 49 | 107.6 | 23 | 21.3755 | 55.2 | 49 | 49 | 100 | 55.2 |  | 55.2 | 55.2
A | Bromoxynil | 68543 | 5 | 49 | 3.6 | 4.02 | 111.6667 |  | 49 | 26 | 53.0612 |  |  | 9.648
|  |  | 10 | 56 | 11.5 | 10.02 | 87.1304 |  | 56 | 48 | 85.7143 |  |  | 24.048
|  |  | 50 | 48 | 49.2 | 16.27 | 33.0691 | 39.048 | 48 | 48 | 100 | 39.048 | 39.048 | 39.048 | 39.048
|  |  | 100 | 49 | 107.7 | 27.83 | 25.8403 | 66.792 | 49 | 49 | 100 | 66.792 |  | 66.792 | 66.792
9030 | Chlorimuron-ethyl | L903C | 5 | 49 | 5.2 | 2.24 | 43.0769 | 5.376 | 49 | 49 | 100 | 5.376 | 5.376 | 5.376 | 5.376
|  |  | 10 | 57 | 12.5 | 6.16 | 49.28 |  | 57 | 53 | 92.9825 |  |  | 14.784
|  |  | 50 | 48 | 52.9 | 17.43 | 32.949 | 41.832 | 48 | 48 | 100 | 41.832 |  | 41.832 | 41.832
|  |  | 100 | 49 | 112.2 | 33.23 | 29.6168 | 79.752 | 49 | 49 | 100 | 79.752 |  | 79.752 | 79.752
E | Chlorosulfonamide acid | 68551 | 5 | 49 | 3.1 | 3.69 | 119.0323 |  | 49 | 25 | 51.0204 |  |  | 8.856
|  |  | 10 | 57 | 11.4 | 6.83 | 59.9123 |  | 57 | 52 | 91.2281 |  |  | 16.392
|  |  | 50 | 48 | 55.3 | 15.49 | 28.0108 | 37.176 | 48 | 48 | 100 | 37.176 | 37.176 | 37.176 | 37.176
|  |  | 100 | 49 | 119.7 | 30.88 | 25.7978 | 74.112 | 49 | 49 | 100 | 74.112 |  | 74.112 | 74.112
9030 | Chlorsulfuron | L903R | 5 | 46 | 2.5 | 4.15 | 166 |  | 46 | 23 | 50 |  |  | 9.96
|  |  | 10 | 57 | 10 | 8.35 | 83.5 |  | 57 | 47 | 82.4561 |  |  | 20.04
|  |  | 50 | 48 | 59.2 | 20.14 | 34.0203 | 48.336 | 48 | 48 | 100 | 48.336 | 48.336 | 48.336 | 48.336
|  |  | 100 | 49 | 117.3 | 33.95 | 28.9429 | 81.48 | 49 | 49 | 100 | 81.48 |  | 81.48 | 81.48
A | cis-Cyhalothric acid | 68553 | 5 | 49 | 1.2 | 8.5 | 708.3333 |  | 49 | 1 | 2.0408 |  |  | 20.4
|  |  | 10 | 57 | 0 | 0 | #DIV/0! |  | 57 | 0 | 0 |  |  | 0
|  |  | 50 | 48 | 7 | 24.67 | 352.4286 |  | 48 | 4 | 8.3333 |  |  | 59.208
|  |  | 100 | 49 | 26.7 | 55.11 | 206.4045 |  | 49 | 11 | 22.449 |  | 250 | 132.264
E | Chlorthal-monomethyl | 68560 | 5 | 49 | 0.1 | 1.04 | 1040 |  | 49 | 1 | 2.0408 |  |  | 2.496
| Chlorthal monomethyl |  | 10 | 57 | 5.4 | 19.2 | 355.5556 |  | 57 | 5 | 8.7719 |  |  | 46.08
|  |  | 50 | 48 | 68.5 | 55.58 | 81.1387 |  | 48 | 37 | 77.0833 |  |  | 133.392
|  |  | 100 | 48 | 107.1 | 60.07 | 56.0878 |  | 48 | 46 | 95.8333 |  | 250 | 144.168
A | Dechlorofipronil | 68561 | 5 | 49 | 5.3 | 0.73 | 13.7736 | 1.752 | 49 | 49 | 100 | 1.752 |  | 1.752 | 1.752
|  |  | 10 | 57 | 12 | 4.7 | 39.1667 |  | 57 | 53 | 92.9825 |  |  | 11.28
|  |  | 50 | 48 | 50.9 | 12.32 | 24.2043 | 29.568 | 48 | 48 | 100 | 29.568 |  | 29.568 | 29.568
|  |  | 100 | 49 | 110.1 | 23.76 | 21.5804 | 57.024 | 49 | 49 | 100 | 57.024 |  | 57.024 | 57.024
D | Deisopropyliprodione | 68565 | 5 | 49 | 2.6 | 7.94 | 305.3846 |  | 49 | 6 | 12.2449 |  |  | 19.056
|  |  | 10 | 57 | 275.6 | 2066.85 | 749.9456 |  | 57 | 6 | 10.5263 |  |  | 4960.44
|  |  | 50 | 48 | 37.9 | 52.56 | 138.6807 |  | 48 | 24 | 50 |  |  | 126.144
|  |  | 100 | 49 | 154.4 | 111.09 | 71.9495 |  | 49 | 42 | 85.7143 |  | 250 | 266.616
A | Desulfinylfipronil | 66607 | 5 | 49 | 5.1 | 0.77 | 15.098 | 1.848 | 49 | 49 | 100 | 1.848 | 1.848 | 1.848 | 1.848
|  |  | 10 | 57 | 11.6 | 4.29 | 36.9828 | 10.296 | 57 | 53 | 92.9825 | 10.296 |  | 10.296 | 10.296
|  |  | 50 | 48 | 49.8 | 10.95 | 21.988 | 26.28 | 48 | 48 | 100 | 26.28 |  | 26.28 | 26.28
|  |  | 100 | 49 | 108 | 19.83 | 18.3611 | 47.592 | 49 | 49 | 100 | 47.592 |  | 47.592 | 47.592
A | Desulfinylfipronil amide | 68570 | 5 | 49 | 5.1 | 1.73 | 33.9216 | 4.152 | 49 | 49 | 100 | 4.152 | 4.152 | 4.152 | 4.152
|  |  | 10 | 57 | 12.3 | 5.18 | 42.1138 |  | 57 | 53 | 92.9825 |  |  | 12.432
|  |  | 50 | 48 | 52.3 | 12.9 | 24.6654 | 30.96 | 48 | 48 | 100 | 30.96 |  | 30.96 | 30.96
|  |  | 100 | 49 | 114.7 | 26.33 | 22.9555 | 63.192 | 49 | 49 | 100 | 63.192 |  | 63.192 | 63.192
E | Dicamba | 68571 | 5 | 49 | 0.9 | 6.15 | 683.3333 |  | 49 | 1 | 2.0408 |  |  | 14.76
|  |  | 10 | 57 | 1 | 7.64 | 764 |  | 57 | 1 | 1.7544 |  |  | 18.336
|  |  | 50 | 48 | 9.7 | 26.91 | 277.4227 |  | 48 | 9 | 18.75 |  |  | 64.584
|  |  | 100 | 49 | 58.8 | 66.08 | 112.381 |  | 49 | 25 | 51.0204 |  | 250 | 158.592
A | Diflufenzopyr | 68577 | 5 | 49 | 4 | 4.48 | 112 |  | 49 | 34 | 69.3878 |  |  | 10.752
|  |  | 10 | 57 | 11.8 | 6.88 | 58.3051 |  | 57 | 52 | 91.2281 |  |  | 16.512
|  |  | 50 | 48 | 56.7 | 18.32 | 32.3104 | 43.968 | 48 | 48 | 100 | 43.968 | 43.968 | 43.968 | 43.968
|  |  | 100 | 49 | 126.1 | 29.08 | 23.0611 | 69.792 | 49 | 49 | 100 | 69.792 |  | 69.792 | 69.792
A | Diketonitrile isoxaflutole | 68578 | 5 | 49 | 6.1 | 2.46 | 40.3279 |  | 49 | 49 | 100 |  |  | 5.904
|  |  | 10 | 57 | 11.8 | 5.22 | 44.2373 |  | 57 | 53 | 92.9825 |  |  | 12.528
|  |  | 50 | 48 | 50.4 | 12.8 | 25.3968 | 30.72 | 48 | 48 | 100 | 30.72 | 30.72 | 30.72 | 30.72
|  |  | 100 | 49 | 108.9 | 22.43 | 20.5969 | 53.832 | 49 | 49 | 100 | 53.832 |  | 53.832 | 53.832
A | Dimethenamid OA | 68581 | 5 | 49 | 2.8 | 5.01 | 178.9286 |  | 49 | 18 | 36.7347 |  |  | 12.024
|  |  | 10 | 57 | 5.9 | 8.17 | 138.4746 |  | 57 | 26 | 45.614 |  |  | 19.608
|  |  | 50 | 48 | 54.1 | 17.56 | 32.4584 | 42.144 | 48 | 47 | 97.9167 | 42.144 | 42.144 | 42.144 | 42.144
|  |  | 100 | 49 | 113.5 | 21.7 | 19.1189 | 52.08 | 49 | 49 | 100 | 52.08 |  | 52.08 | 52.08
A | Dimethenamid SA | 68582 | 5 | 49 | 4.7 | 6.02 | 128.0851 |  | 49 | 26 | 53.0612 |  |  | 14.448
|  |  | 10 | 57 | 10.4 | 10.31 | 99.1346 |  | 57 | 40 | 70.1754 |  |  | 24.744
|  |  | 50 | 48 | 53 | 16.43 | 31 | 39.432 | 48 | 48 | 100 | 39.432 | 39.432 | 39.432 | 39.432
|  |  | 100 | 49 | 112.2 | 30.3 | 27.0053 | 72.72 | 49 | 49 | 100 | 72.72 |  | 72.72 | 72.72
A | Dimethenamid SAA | 68583 | 5 | 49 | 1.9 | 7.22 | 380 |  | 49 | 4 | 8.1633 |  |  | 17.328
|  |  | 10 | 57 | 11.3 | 18.54 | 164.0708 |  | 57 | 21 | 36.8421 |  |  | 44.496
|  |  | 50 | 47 | 66.6 | 30.33 | 45.5405 |  | 47 | 48 | 102.1277 |  |  | 72.792
|  |  | 100 | 49 | 121 | 39.36 | 32.5289 | 94.464 | 49 | 49 | 100 | 94.464 | 94.464 | 94.464 | 94.464
A | Fipronil | 66604 | 5 | 49 | 5.1 | 0.83 | 16.2745 | 1.992 | 49 | 49 | 100 | 1.992 | 1.992 | 1.992 | 1.992
|  |  | 10 | 57 | 12 | 4.54 | 37.8333 | 10.896 | 57 | 53 | 92.9825 | 10.896 |  | 10.896 | 10.896
|  |  | 50 | 48 | 50.2 | 11.15 | 22.2112 | 26.76 | 48 | 48 | 100 | 26.76 |  | 26.76 | 26.76
|  |  | 100 | 49 | 108.6 | 20.64 | 19.0055 | 49.536 | 49 | 49 | 100 | 49.536 |  | 49.536 | 49.536
A | Fipronil amide | 68604 | 5 | 49 | 4.8 | 1.67 | 34.7917 | 4.008 | 49 | 49 | 100 | 4.008 | 4.008 | 4.008 | 4.008
|  |  | 10 | 57 | 12.4 | 5.38 | 43.3871 |  | 57 | 53 | 92.9825 |  |  | 12.912
|  |  | 50 | 48 | 53 | 12.83 | 24.2075 | 30.792 | 48 | 48 | 100 | 30.792 |  | 30.792 | 30.792
|  |  | 100 | 49 | 117.4 | 27.86 | 23.7308 | 66.864 | 49 | 49 | 100 | 66.864 |  | 66.864 | 66.864
A | Fipronil sulfide | 66610 | 5 | 49 | 5.4 | 0.83 | 15.3704 | 1.992 | 49 | 49 | 100 | 1.992 | 1.992 | 1.992 | 1.992
|  |  | 10 | 57 | 12 | 4.71 | 39.25 |  | 57 | 53 | 92.9825 |  |  | 11.304
|  |  | 50 | 48 | 50.2 | 11.4 | 22.7092 | 27.36 | 48 | 48 | 100 | 27.36 |  | 27.36 | 27.36
|  |  | 100 | 49 | 108.5 | 20.61 | 18.9954 | 49.464 | 49 | 49 | 100 | 49.464 |  | 49.464 | 49.464
A | Fipronil sulfonate | 68605 | 5 | 48 | 2.8 | 4 | 142.8571 |  | 48 | 20 | 41.6667 |  |  | 9.6
|  |  | 10 | 56 | 9.9 | 8.28 | 83.6364 |  | 56 | 40 | 71.4286 |  |  | 19.872
|  |  | 50 | 48 | 49.4 | 20.1 | 40.6883 | 48.24 | 48 | 46 | 95.8333 | 48.24 | 48.24 | 48.24 | 48.24
|  |  | 100 | 49 | 108.3 | 26.86 | 24.8015 | 64.464 | 49 | 49 | 100 | 64.464 |  | 64.464 | 64.464
A | Fipronil sulfone | 66613 | 5 | 49 | 5.2 | 0.77 | 14.8077 | 1.848 | 49 | 49 | 100 | 1.848 | 1.848 | 1.848 | 1.848
|  |  | 10 | 57 | 13.9 | 15.47 | 111.295 |  | 57 | 54 | 94.7368 |  |  | 37.128
|  |  | 50 | 48 | 49.1 | 11.02 | 22.444 | 26.448 | 48 | 48 | 100 | 26.448 |  | 26.448 | 26.448
|  |  | 100 | 49 | 106.2 | 19.98 | 18.8136 | 47.952 | 49 | 49 | 100 | 47.952 |  | 47.952 | 47.952
A | Flubendiamide | 68606 | 5 | 49 | 5.3 | 0.93 | 17.5472 | 2.232 | 49 | 49 | 100 | 2.232 | 2.232 | 2.232 | 2.232
|  |  | 10 | 57 | 11.9 | 4.57 | 38.4034 | 10.968 | 57 | 53 | 92.9825 | 10.968 |  | 10.968 | 10.968
|  |  | 50 | 48 | 49.6 | 12.76 | 25.7258 | 30.624 | 48 | 48 | 100 | 30.624 |  | 30.624 | 30.624
|  |  | 100 | 49 | 107.3 | 22.7 | 21.1556 | 54.48 | 49 | 49 | 100 | 54.48 |  | 54.48 | 54.48
E | Hexazinone TP D | 68613 | 5 | 49 | 2.3 | 7.14 | 310.4348 |  | 49 | 6 | 12.2449 |  |  | 17.136
|  |  | 10 | 57 | 8.5 | 17.32 | 203.7647 |  | 57 | 20 | 35.0877 |  |  | 41.568
|  |  | 50 | 48 | 64.2 | 37.38 | 58.2243 |  | 48 | 48 | 100 |  |  | 89.712
|  |  | 100 | 49 | 120.5 | 49.95 | 41.4523 |  | 49 | 49 | 100 |  | 250 | 119.88
A | Hexazinone TP E | 68614 | 5 | 49 | 0.2 | 1.11 | 555 |  | 49 | 1 | 2.0408 |  |  | 2.664
|  |  | 10 | 57 | 7.1 | 8.73 | 122.9577 |  | 57 | 27 | 47.3684 |  |  | 20.952
|  |  | 50 | 48 | 49.8 | 15.63 | 31.3855 | 37.512 | 48 | 47 | 97.9167 | 37.512 | 37.512 | 37.512 | 37.512
|  |  | 100 | 49 | 109.2 | 29.03 | 26.5842 | 69.672 | 49 | 49 | 100 | 69.672 |  | 69.672 | 69.672
D | Iprodione | 66617 | 5 | 49 | 107 | 325.11 | 303.8411 |  | 49 | 7 | 14.2857 |  |  | 780.264
|  |  | 10 | 57 | 0 | 0 | #DIV/0! |  | 57 | 0 | 0 |  |  | 0
|  |  | 50 | 48 | 556.9 | 848.73 | 152.4026 |  | 48 | 22 | 45.8333 |  |  | 2036.952
|  |  | 100 | 49 | 605.2 | 971.19 | 160.4742 |  | 49 | 25 | 51.0204 |  |  | 2330.856
A | Isoxaflutole acid RPA 203328 | 68633 | 5 | 49 | 5.1 | 1.87 | 36.6667 | 4.488 | 49 | 49 | 100 | 4.488 | 4.488 | 4.488 | 4.488
|  |  | 10 | 57 | 12 | 5.16 | 43 |  | 57 | 53 | 92.9825 |  |  | 12.384
|  |  | 50 | 48 | 53.9 | 12.05 | 22.3562 | 28.92 | 48 | 48 | 100 | 28.92 |  | 28.92 | 28.92
|  |  | 100 | 49 | 115.6 | 21.83 | 18.8841 | 52.392 | 49 | 49 | 100 | 52.392 |  | 52.392 | 52.392
A | MCPA | 68641 | 5 | 49 | 0.6 | 3.25 | 541.6667 |  | 49 | 2 | 4.0816 |  |  | 7.8
|  |  | 10 | 57 | 3.3 | 7.08 | 214.5455 |  | 57 | 12 | 21.0526 |  |  | 16.992
|  |  | 50 | 48 | 50.6 | 19.86 | 39.249 | 47.664 | 48 | 45 | 93.75 | 47.664 | 47.664 | 47.664 | 47.664
|  |  | 100 | 49 | 108.8 | 22.96 | 21.1029 | 55.104 | 49 | 49 | 100 | 55.104 |  | 55.104 | 55.104
A | Metolachlor OA | 68650 | 5 | 49 | 1.2 | 3.42 | 285 |  | 49 | 6 | 12.2449 |  |  | 8.208
|  |  | 10 | 57 | 4.7 | 7.58 | 161.2766 |  | 57 | 22 | 38.5965 |  |  | 18.192
|  |  | 50 | 47 | 46.1 | 23.03 | 49.9566 |  | 47 | 45 | 95.7447 |  |  | 55.272
|  |  | 100 | 49 | 111.9 | 30.96 | 27.6676 | 74.304 | 49 | 49 | 100 | 74.304 | 74.304 | 74.304 | 74.304
A | Metolachlor SA | 68651 | 5 | 49 | 1.9 | 3.52 | 185.2632 |  | 49 | 18 | 36.7347 |  |  | 8.448
|  |  | 10 | 56 | 10.2 | 8.06 | 79.0196 |  | 56 | 48 | 85.7143 |  |  | 19.344
|  |  | 50 | 48 | 47.4 | 14.03 | 29.5992 | 33.672 | 48 | 48 | 100 | 33.672 | 33.672 | 33.672 | 33.672
|  |  | 100 | 49 | 116.4 | 27.17 | 23.3419 | 65.208 | 49 | 49 | 100 | 65.208 |  | 65.208 | 65.208
A | Metribuzin DADK | 68569 | 5 | 49 | 1.3 | 5.82 | 447.6923 |  | 49 | 3 | 6.1224 |  |  | 13.968
|  |  | 10 | 57 | 2.3 | 6.23 | 270.8696 |  | 57 | 10 | 17.5439 |  |  | 14.952
|  |  | 50 | 48 | 39.3 | 29.09 | 74.0204 |  | 48 | 42 | 87.5 |  |  | 69.816
|  |  | 100 | 49 | 94.2 | 34.83 | 36.9745 | 83.592 | 49 | 49 | 100 | 83.592 | 83.592 | 83.592 | 83.592
A | Metribuzin DK | 68653 | 5 | 49 | 0 | 0 | #DIV/0! |  | 49 | 0 | 0 |  |  | 0
|  |  | 10 | 57 | 0.6 | 4.48 | 746.6667 |  | 57 | 1 | 1.7544 |  |  | 10.752
|  |  | 50 | 46 | 35.5 | 34.69 | 97.7183 |  | 46 | 37 | 80.4348 |  |  | 83.256
|  |  | 100 | 49 | 106.5 | 41.36 | 38.8357 | 99.264 | 49 | 49 | 100 | 99.264 | 99.264 | 99.264 | 99.264
9030 | Novaluron | L903J | 5 | 49 | 5.3 | 1.31 | 24.717 | 3.144 | 49 | 49 | 100 | 3.144 | 3.144 | 3.144 | 3.144
|  |  | 10 | 57 | 12.2 | 4.71 | 38.6066 |  | 57 | 53 | 92.9825 |  |  | 11.304
|  |  | 50 | 48 | 47.8 | 15.71 | 32.8661 | 37.704 | 48 | 48 | 100 | 37.704 |  | 37.704 | 37.704
|  |  | 100 | 49 | 97.4 | 32.26 | 33.1211 | 77.424 | 49 | 49 | 100 | 77.424 |  | 77.424 | 77.424
A | Oryzalin | 68663 | 5 | 49 | 3.8 | 2.16 | 56.8421 | 5.184 | 49 | 47 | 95.9184 | 5.184 | 5.184 | 5.184 | 5.184
|  |  | 10 | 57 | 12.3 | 5.68 | 46.1789 |  | 57 | 53 | 92.9825 |  |  | 13.632
|  |  | 50 | 48 | 52.1 | 15.62 | 29.9808 | 37.488 | 48 | 48 | 100 | 37.488 |  | 37.488 | 37.488
|  |  | 100 | 49 | 111.7 | 27.55 | 24.6643 | 66.12 | 49 | 49 | 100 | 66.12 |  | 66.12 | 66.12
9030 | Propanil | L903F | 5 | 49 | 2 | 2.53 | 126.5 |  | 49 | 26 | 53.0612 |  |  | 6.072
|  |  | 10 | 57 | 9.4 | 5.11 | 54.3617 |  | 57 | 48 | 84.2105 |  |  | 12.264
|  |  | 50 | 48 | 51.4 | 12 | 23.3463 | 28.8 | 48 | 48 | 100 | 28.8 | 28.8 | 28.8 | 28.8
|  |  | 100 | 49 | 114.1 | 23.47 | 20.5697 | 56.328 | 49 | 49 | 100 | 56.328 |  | 56.328 | 56.328
9030 | Prosulfuron | 61687 | 5 | 49 | 5.9 | 1.42 | 24.0678 | 3.408 | 49 | 49 | 100 | 3.408 | 3.408 | 3.408 | 3.408
|  |  | 10 | 57 | 13.2 | 5.73 | 43.4091 |  | 57 | 53 | 92.9825 |  |  | 13.752
|  |  | 50 | 48 | 56.3 | 17.94 | 31.865 | 43.056 | 48 | 48 | 100 | 43.056 |  | 43.056 | 43.056
|  |  | 100 | 49 | 112.2 | 34.7 | 30.9269 | 83.28 | 49 | 49 | 100 | 83.28 |  | 83.28 | 83.28
A | sec-Acetochlor OA | 68684 | 5 | 49 | 0.5 | 2.04 | 408 |  | 49 | 3 | 6.1224 |  |  | 4.896
|  |  | 10 | 57 | 3.4 | 5.05 | 148.5294 |  | 57 | 19 | 33.3333 |  |  | 12.12
|  |  | 50 | 48 | 50.5 | 10.65 | 21.0891 | 25.56 | 48 | 48 | 100 | 25.56 | 25.56 | 25.56 | 25.56
|  |  | 100 | 49 | 111.8 | 19.53 | 17.4687 | 46.872 | 49 | 49 | 100 | 46.872 |  | 46.872 | 46.872
A | sec-Alachlor OA | 68685 | 5 | 49 | 0.3 | 1.5 | 500 |  | 49 | 2 | 4.0816 |  |  | 3.6
|  |  | 10 | 57 | 3.1 | 6.32 | 203.871 |  | 57 | 15 | 26.3158 |  |  | 15.168
|  |  | 50 | 48 | 29.7 | 29.56 | 99.5286 |  | 48 | 31 | 64.5833 |  |  | 70.944
|  |  | 100 | 49 | 104.2 | 28.18 | 27.0441 | 67.632 | 49 | 49 | 100 | 67.632 | 67.632 | 67.632 | 67.632
A | Sulfentrazone | 68687 | 5 | 49 | 5.4 | 2.27 | 42.037 | 5.448 | 49 | 48 | 97.9592 | 5.448 | 5.448 | 5.448 | 5.448
|  |  | 10 | 57 | 13.2 | 6.4 | 48.4848 |  | 57 | 53 | 92.9825 |  |  | 15.36
|  |  | 50 | 48 | 56.7 | 18.78 | 33.1217 | 45.072 | 48 | 48 | 100 | 45.072 |  | 45.072 | 45.072
|  |  | 100 | 49 | 117.4 | 35.45 | 30.1959 | 85.08 | 49 | 49 | 100 | 85.08 |  | 85.08 | 85.08
A | Triclopyr | 68712 | 5 | 49 | 4.8 | 5.29 | 110.2083 |  | 49 | 30 | 61.2245 |  |  | 12.696
|  |  | 10 | 56 | 9.7 | 8.71 | 89.7938 |  | 56 | 41 | 73.2143 |  |  | 20.904
|  |  | 50 | 48 | 50.6 | 18.13 | 35.83 | 43.512 | 48 | 48 | 100 | 43.512 | 43.512 | 43.512 | 43.512
|  |  | 100 | 49 | 107.4 | 26.27 | 24.46 | 63.048 | 49 | 49 | 100 | 63.048 |  | 63.048 | 63.048