## Sheet1
Appendix 2. The types of bias or variability associated with the inorganic analyses and information on corrective actions taken by the U.S. Geological Survey's National Water Quality Laboratory, Denver, Colorado, during the combined sewer overflow study in Omaha, Nebraska.
[Data Source: U.S. Geological Survey, Branch of Quality System Inorganic Blind Sample Project, accessed December 12, 2008, at http://bqs.usgs.gov/bsp/; %, percent; parmcode, parameter code; mg/L, milligrams per liter; QC, quality control; ICP-MS, inductively coupled plasma-mass spectroscopy; NYSDOH, New York State Department of Health; NWQL, National Water Quality Laboratory; CVAF, cold vapor atomic fluorescence]
Constituent | Type of bias or variability noted | Date bias was noted | Date corrected | Corrective actions taken
Ammonia, whole-water recoverable, colorimetry, parmcode 00610CL017 | Positive bias (8.1%) | April 2007 | July 2007 | The sample data exhibiting a downward shift were analyzed just before the service provider performed semi-annual preventative maintenance on the instrument (completed May 7, 2007). The maintenance corrected the problem, but monitoring of the line continued.
Phosphorus, whole-water recoverable, Colorimetry, parmcode 0665D | Positive bias (12.6%) primarily from the low-concentration (0.04 mg/L) reference material. | October 2006 | December 2006 | Problems in maintenance were addressed and corrected in early December 2006.
Phosphorus, whole-water recoverable, Colorimetry, parmcode 0665H (Alkaline per-sulfate digestion) | Positive bias (12.6%) | December 2006 | February 2007 | New standards were prepared on January 22, 2007, and a new QC stock was received the week of February 5, 2007; analysis of subsequent samples showed no bias and good variability.
Barium, whole-water recoverable, ICP-MS, parmcode 1007D | Negative bias (7.1%) | February 2006 | November 2006 | Online QC charts do not show a similar bias for any of the low, middle, or high range checks.
Beryllium, whole-water recoverable, ICP-MS, parmcode 1012D | Negative bias (10.3%) | April 2006 | October 2006 | New calibration standards prepared, verified, and put into service by the end of October 2006.
Cobalt, whole-water recoverable, ICP-MS, parmcode 1037H | Negative bias (8.0%) | August 2006 | October 2006 | QC check of calibration standards using the latest NYSDOH sample and Standard Reference Water Samples did not indicate same bias.
Iron, whole-water recoverable, ICP, parmcode 1045C | Positive bias (7.2%) | May 2007 | September 2007 | In September 2007, the entire front-end of the ICP was replaced, after which QC charts indicated that the system has settled.
Manganese, whole-water recoverable, ICP, parmcode 1055D | Negative bias (5.9%) | May 2006 | November 2006 | New calibration standards prepared the last week in October 2006, after which QC charts indicated that the system has settled.
Manganese, whole-water recoverable, ICP-MS, parmcode 1055C | Negative bias (6.5%) | May 2006 | April 2007 | The bias appeared to be inherent to the specific ICP-MS instrument, the use of which was discontinued.
Manganese, whole-water recoverable, ICP, parmcode 01055PLA15 | Slight negative bias (6.4%) | March 2007 | November 2007 | The NWQL had the instrument manufacturer come out the week of October 15th, 2007, and do a complete check of the instrument. He optimized the instrument and subsequent runs indicated reduced variability.
Mercury, whole-water recoverable, CVAF, parmcode 71900D | Slight negative bias (6.2%) | May 2006 | October 2006 | New calibration standards were prepared the first week of October 2006. Analysis since then indicates that bias had been corrected.
Mercury, unfiltered, CVAF, parmcode 71900CV018 | Negative bias (9.3%) | May 2007 | September 2007 | New calibration standards were prepared in September 2007 and subsequent QC charts indicated that bias has been fixed. New calibration standards also were prepared in early December.
Zinc, whole-water recoverable, ICP-MS, parmcode 1092D | Negative bias (10.4%) | April 2006 | December 2006 | NWQL QC charts indicated a similar bias in November 2006, and new calibration standards were prepared. Analysis since then indicates that the bias has been corrected.