## Table 3-1 results
Table 3-1. Concentrations of elements in diluter water from sturgeon chronic exposures.
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51245 | C-Filter Blank | 2010-07-12 | 0 | < 0.04 | < 0.2 | < 0.024 | < 0.025
51246 | C-Cu-0 | 2010-07-12 | 0 | 0.26 | --- | --- | ---
51274 | C-Cu-0 unfiltered | 2010-07-12 | 0 | 0.31 | --- | --- | ---
51247 | C-Cu-0.5 | 2010-07-12 | 0 | 0.48 | --- | --- | ---
51248 | C-Cu-1.0 | 2010-07-12 | 0 | 0.9 | --- | --- | ---
51275 | C-Cu-MedLo unfiltered | 2010-07-12 | 0 | 0.99 | --- | --- | ---
51249 | C-Cu-2.0 | 2010-07-12 | 0 | 1.53 | --- | --- | ---
51270 | C-Cu-2.0 Duplicate | 2010-07-12 | 0 | 1.53 | --- | --- | ---
51250 | C-Cu-4.0 | 2010-07-12 | 0 | 3.02 | --- | --- | ---
51251 | C-Cu-8.0 | 2010-07-12 | 0 | 6.09 | --- | --- | ---
51252 | C-Zn-0 | 2010-07-12 | 0 | --- | 2.41 | --- | ---
51276 | C-Zn-0 unfiltered | 2010-07-12 | 0 | --- | 2.58 | --- | ---
51253 | C-Zn-25 | 2010-07-12 | 0 | --- | 21.4 | --- | ---
51254 | C-Zn-50 | 2010-07-12 | 0 | --- | 42.9 | --- | ---
51277 | C-Zn-MedLo unfiltered | 2010-07-12 | 0 | --- | 42.2 | --- | ---
51255 | C-Zn-100 | 2010-07-12 | 0 | --- | 86.6 | --- | ---
51271 | C-Zn-100 Duplicate | 2010-07-12 | 0 | --- | 87.8 | --- | ---
51256 | C-Zn-200 | 2010-07-12 | 0 | --- | 161 | --- | ---
51257 | C-Zn-400 | 2010-07-12 | 0 | --- | 365 | --- | ---
51258 | C-Cd-0 | 2010-07-12 | 0 | --- | --- | < 0.024 | ---
51278 | C-Cd-0 unfiltered | 2010-07-12 | 0 | --- | --- | < 0.024 | ---
51259 | C-Cd-0.75 | 2010-07-12 | 0 | --- | --- | 0.5 | ---
51260 | C-Cd-1.5 | 2010-07-12 | 0 | --- | --- | 1.04 | ---
51279 | C-Cd-MedLo unfiltered | 2010-07-12 | 0 | --- | --- | 1.03 | ---
51261 | C-Cd-3.0 | 2010-07-12 | 0 | --- | --- | 2.37 | ---
51272 | C-Cd-3.0 Duplicate | 2010-07-12 | 0 | --- | --- | 2.34 | ---
51262 | C-Cd-6.0 | 2010-07-12 | 0 | --- | --- | 4.55 | ---
51263 | C-Cd-12.0 | 2010-07-12 | 0 | --- | --- | 10.4 | ---
51264 | C-Pb-0 | 2010-07-12 | 0 | --- | --- | --- | < 0.025
51280 | C-Pb-0 unfiltered | 2010-07-12 | 0 | --- | --- | --- | < 0.025
51265 | C-Pb-5 | 2010-07-12 | 0 | --- | --- | --- | 2.18
51266 | C-Pb-10 | 2010-07-12 | 0 | --- | --- | --- | 4.26
51281 | C-Pb-MedLo unfiltered | 2010-07-12 | 0 | --- | --- | --- | 4.92
51267 | C-Pb-20 | 2010-07-12 | 0 | --- | --- | --- | 8.37
51273 | C-Pb-20 Duplicate | 2010-07-12 | 0 | --- | --- | --- | 8.66
51268 | C-Pb-40 | 2010-07-12 | 0 | --- | --- | --- | 21.1
51269 | C-Pb-80 | 2010-07-12 | 0 | --- | --- | --- | 50.6
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51530 | C-Filter Blank | 2010-08-02 | 3 | < 0.061 | < 0.4 | < 0.009 | < 0.031
51531 | C-Cu-0ppb | 2010-08-02 | 3 | 0.21 | --- | --- | ---
51532 | C-Cu-0.5ppb | 2010-08-02 | 3 | 0.47 | --- | --- | ---
51533 | C-Cu-1.0ppb | 2010-08-02 | 3 | 1.06 | --- | --- | ---
51534 | C-Cu-2.0ppb | 2010-08-02 | 3 | 1.9 | --- | --- | ---
51555 | C-Cu-2ppb duplicate | 2010-08-02 | 3 | 1.83 | --- | --- | ---
51535 | C-Cu-4.0ppb | 2010-08-02 | 3 | 3.69 | --- | --- | ---
51536 | C-Cu-8.0ppb | 2010-08-02 | 3 | 7.62 | --- | --- | ---
51537 | C-Zn-0ppb | 2010-08-02 | 3 | --- | 1.13 | --- | ---
51538 | C-Zn-25ppb | 2010-08-02 | 3 | --- | 22.6 | --- | ---
51539 | C-Zn-50ppb | 2010-08-02 | 3 | --- | 46.6 | --- | ---
51540 | C-Zn-100ppb | 2010-08-02 | 3 | --- | 90.7 | --- | ---
51556 | C-Zn-100ppb duplicate | 2010-08-02 | 3 | --- | 91.9 | --- | ---
51541 | C-Zn-200ppb | 2010-08-02 | 3 | --- | 175 | --- | ---
51542 | C-Zn-400ppb | 2010-08-02 | 3 | --- | 402 | --- | ---
51543 | C-Cd-0ppb | 2010-08-02 | 3 | --- | --- | < 0.009 | ---
51544 | C-Cd-0.75ppb | 2010-08-02 | 3 | --- | --- | 0.66 | ---
51545 | C-Cd-1.5ppb | 2010-08-02 | 3 | --- | --- | 1.3 | ---
51546 | C-Cd-3.0ppb | 2010-08-02 | 3 | --- | --- | 2.75 | ---
51557 | C-Cd-3ppb duplicate | 2010-08-02 | 3 | --- | --- | 2.76 | ---
51547 | C-Cd-6.0ppb | 2010-08-02 | 3 | --- | --- | 5.05 | ---
51548 | C-Cd-12ppb | 2010-08-02 | 3 | --- | --- | 11.5 | ---
51549 | C-Pb-0ppb | 2010-08-02 | 3 | --- | --- | --- | < 0.031
51550 | C-Pb-5ppb | 2010-08-02 | 3 | --- | --- | --- | 3.26
51551 | C-Pb-10ppb | 2010-08-02 | 3 | --- | --- | --- | 6.24
51552 | C-Pb-20ppb | 2010-08-02 | 3 | --- | --- | --- | 13.2
51558 | C-Pb-20ppb duplicate | 2010-08-02 | 3 | --- | --- | --- | 13
51553 | C-Pb-40ppb | 2010-08-02 | 3 | --- | --- | --- | 28.1
51554 | C-Pb-80ppb | 2010-08-02 | 3 | --- | --- | --- | 61
51326 | C-Cu-0ppb | 2010-07-16 | 4 | 0.15 | --- | --- | ---
51327 | C-Cu-0.5ppb | 2010-07-16 | 4 | 0.37 | --- | --- | ---
51328 | C-Cu-1.0ppb | 2010-07-16 | 4 | 0.87 | --- | --- | ---
51329 | C-Cu-2.0ppb | 2010-07-16 | 4 | 1.73 | --- | --- | ---
51330 | C-Cu-4.ppb | 2010-07-16 | 4 | 3.27 | --- | --- | ---
51331 | C-Cu-8.0ppb | 2010-07-16 | 4 | 7.24 | --- | --- | ---
51332 | C-Zn-400ppb | 2010-07-16 | 4 | --- | 371 | --- | ---
51333 | C-Cd-12ppb | 2010-07-16 | 4 | --- | --- | 11 | ---
51334 | C-Pb-80ppb | 2010-07-16 | 4 | --- | --- | --- | 56.8
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51335 | C-Filter Blank | 2010-07-19 | 7 | < 0.087 | < 0.4 | < 0.009 | < 0.019
51336 | C-Cu-0 | 2010-07-19 | 7 | 0.22 | --- | --- | ---
51337 | C-Cu-0.5 | 2010-07-19 | 7 | 0.46 | --- | --- | ---
51338 | C-Cu-1.0 | 2010-07-19 | 7 | 0.98 | --- | --- | ---
51339 | C-Cu-2.0 | 2010-07-19 | 7 | 1.81 | --- | --- | ---
51360 | C-Cu-2.0 Duplicate | 2010-07-19 | 7 | 1.6 | --- | --- | ---
51340 | C-Cu-4.0 | 2010-07-19 | 7 | 3.39 | --- | --- | ---
51341 | C-Cu-8.0 | 2010-07-19 | 7 | 7.41 | --- | --- | ---
51342 | C-Zn-0 | 2010-07-19 | 7 | --- | 0.66 | --- | ---
51343 | C-Zn-25 | 2010-07-19 | 7 | --- | 21.8 | --- | ---
51344 | C-Zn-50 | 2010-07-19 | 7 | --- | 44.7 | --- | ---
51345 | C-Zn-100 | 2010-07-19 | 7 | --- | 90 | --- | ---
51361 | C-Zn-100 Duplicate | 2010-07-19 | 7 | --- | 87.6 | --- | ---
51346 | C-Zn-200 | 2010-07-19 | 7 | --- | 169 | --- | ---
51347 | C-Zn-400 | 2010-07-19 | 7 | --- | 372 | --- | ---
51348 | C-Cd-0 | 2010-07-19 | 7 | --- | --- | 0.011 | ---
51349 | C-Cd-0.75 | 2010-07-19 | 7 | --- | --- | 0.59 | ---
51350 | C-Cd-1.5 | 2010-07-19 | 7 | --- | --- | 1.25 | ---
51351 | C-Cd-3.0 | 2010-07-19 | 7 | --- | --- | 2.52 | ---
51362 | C-Cd-3.0 Duplicate | 2010-07-19 | 7 | --- | --- | 2.56 | ---
51352 | C-Cd-6.0 | 2010-07-19 | 7 | --- | --- | 4.93 | ---
51353 | C-Cd-12.0 | 2010-07-19 | 7 | --- | --- | 11.4 | ---
51354 | C-Pb-0 | 2010-07-19 | 7 | --- | --- | --- | < 0.019
51355 | C-Pb-5 | 2010-07-19 | 7 | --- | --- | --- | 2.55
51356 | C-Pb-10 | 2010-07-19 | 7 | --- | --- | --- | 4.91
51357 | C-Pb-20 | 2010-07-19 | 7 | --- | --- | --- | 10.2
51363 | C-Pb-20 Duplicate | 2010-07-19 | 7 | --- | --- |  | 10.1
51358 | C-Pb-40 | 2010-07-19 | 7 | --- | --- | --- | 25.6
51359 | C-Pb-80 | 2010-07-19 | 7 | --- | --- | --- | 57.6
51673 | C-Filter Blank | 2010-08-16 | 10 | < 0.019 | < 0.2 | < 0.017 | < 0.031
51674 | C-Cu-0 | 2010-08-16 | 10 | 0.26 | --- | --- | ---
51675 | C-Cu-0.5 | 2010-08-16 | 10 | 0.51 | --- | --- | ---
51676 | C-Cu-1.0 | 2010-08-16 | 10 | 0.89 | --- | --- | ---
51677 | C-Cu-2.0 | 2010-08-16 | 10 | 1.79 | --- | --- | ---
51698 | C-Cu-2.0 Duplicate | 2010-08-16 | 10 | 1.72 | --- | --- | ---
51678 | C-Cu-4.0 | 2010-08-16 | 10 | 3.07 | --- | --- | ---
51679 | C-Cu-8.0 | 2010-08-16 | 10 | 6.94 | --- | --- | ---
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51680 | C-Zn-0 | 2010-08-16 | 10 | --- | 1.27 | --- | ---
51681 | C-Zn-25 | 2010-08-16 | 10 | --- | 24.1 | --- | ---
51682 | C-Zn-50 | 2010-08-16 | 10 | --- | 48.4 | --- | ---
51683 | C-Zn-100 | 2010-08-16 | 10 | --- | 94.3 | --- | ---
51699 | C-Zn-100 Duplicate | 2010-08-16 | 10 | --- | 94.3 | --- | ---
51684 | C-Zn-200 | 2010-08-16 | 10 | --- | 188 | --- | ---
51685 | C-Zn-400 | 2010-08-16 | 10 | --- | 400 | --- | ---
51686 | C-Cd-0 | 2010-08-16 | 10 | --- | --- | < 0.017 | ---
51687 | C-Cd-0.75 | 2010-08-16 | 10 | --- | --- | 0.62 | ---
51688 | C-Cd-1.5 | 2010-08-16 | 10 | --- | --- | 1.29 | ---
51689 | C-Cd-3.0 | 2010-08-16 | 10 | --- | --- | 2.76 | ---
51700 | C-Cd-3.0 Duplicate | 2010-08-16 | 10 | --- | --- | 2.73 | ---
51690 | C-Cd-6.0 | 2010-08-16 | 10 | --- | --- | 5.33 | ---
51691 | C-Cd-12.0 | 2010-08-16 | 10 | --- | --- | 11.5 | ---
51692 | C-Pb-0 | 2010-08-16 | 10 | --- | --- | --- | < 0.031
51693 | C-Pb-5 | 2010-08-16 | 10 | --- | --- | --- | 2.68
51694 | C-Pb-10 | 2010-08-16 | 10 | --- | --- | --- | 5.71
51695 | C-Pb-20 | 2010-08-16 | 10 | --- | --- | --- | 12.1
51701 | C-Pb-20 Duplicate | 2010-08-16 | 10 | --- | --- | --- | 12.3
51696 | C-Pb-40 | 2010-08-16 | 10 | --- | --- | --- | 26.9
51697 | C-Pb-80 | 2010-08-16 | 10 | --- | --- | --- | 52.5
51406 | C-Filter Blank | 2010-07-26 | 14 | < 0.087 | < 0.4 | < 0.009 | < 0.019
51407 | C-Cu-0 | 2010-07-26 | 14 | 0.27 | --- | --- | ---
51408 | C-Cu 0.5 | 2010-07-26 | 14 | 0.25 | --- | --- | ---
51409 | C-Cu-1.0 | 2010-07-26 | 14 | 0.84 | --- | --- | ---
51410 | C-Cu-2.0 | 2010-07-26 | 14 | 1.58 | --- | --- | ---
51431 | C-Cu-2.0 Duplicate | 2010-07-26 | 14 | 1.76 | --- | --- | ---
51411 | C-Cu-4.0 | 2010-07-26 | 14 | 3.15 | --- | --- | ---
51412 | C-Cu-8.0 | 2010-07-26 | 14 | 6.77 | --- | --- | ---
51413 | C-Zn-0 | 2010-07-26 | 14 | --- | < 0.4 | --- | ---
51414 | C-Zn-25 | 2010-07-26 | 14 | --- | 20.7 | --- | ---
51415 | C-Zn-50 | 2010-07-26 | 14 | --- | 41.3 | --- | ---
51416 | C-Zn-100 | 2010-07-26 | 14 | --- | 87.1 | --- | ---
51432 | C-Zn-100 Duplicate | 2010-07-26 | 14 | --- | 91.6 |  | ---
51417 | C-Zn-200 | 2010-07-26 | 14 | --- | 167 | --- | ---
51418 | C-Zn-400 | 2010-07-26 | 14 | --- | 367 | --- | ---
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51419 | C-Cd-0 | 2010-07-26 | 14 | --- | --- | < 0.009 | ---
51420 | C-Cd-0.75 | 2010-07-26 | 14 | --- | --- | 0.57 | ---
51421 | C-Cd-1.5 | 2010-07-26 | 14 | --- | --- | 1.21 | ---
51422 | C-Cd-3.0 | 2010-07-26 | 14 | --- | --- | 2.68 | ---
51433 | C-Cd-3.0 Duplicate | 2010-07-26 | 14 | --- | --- | 2.59 | ---
51423 | C-Cd-6.0 | 2010-07-26 | 14 | --- | --- | 4.92 | ---
51424 | C-Cd-12.0 | 2010-07-26 | 14 | --- | --- | 10.9 | ---
51425 | C-Pb-0 | 2010-07-26 | 14 | --- | --- | --- | < 0.019
51426 | C-Pb-5 | 2010-07-26 | 14 | --- | --- | --- | 2.61
51427 | C-Pb-10 | 2010-07-26 | 14 | --- | --- | --- | 5.03
51428 | C-Pb-20 | 2010-07-26 | 14 | --- | --- | --- | 10.7
51434 | C-Pb-20 Duplicate | 2010-07-26 | 14 | --- | --- | --- | 11
51429 | C-Pb-40 | 2010-07-26 | 14 | --- | --- | --- | 25.4
51430 | C-Pb-80 | 2010-07-26 | 14 | --- | --- | --- | 58.4
51876 | C-Filter Blank | 2010-08-23 | 17 | < 0.084 | < 1.2 | < 0.033 | < 0.008
51877 | C-S2-Cu-0 | 2010-08-23 | 17 | 0.21 | --- | --- | ---
51878 | C-S2-Cu-0.5 | 2010-08-23 | 17 | 0.44 | --- | --- | ---
51879 | C-S2-Cu-1.0 | 2010-08-23 | 17 | 0.97 | --- | --- | ---
51880 | C-S2-Cu-2.0 | 2010-08-23 | 17 | 1.85 | --- | --- | ---
51901 | C-S2-Cu-2.0 Duplicate | 2010-08-23 | 17 | 1.84 | --- | --- | ---
51881 | C-S2-Cu-4.0 | 2010-08-23 | 17 | 3.29 | --- | --- | ---
51882 | C-S2-Cu-8.0 | 2010-08-23 | 17 | 7.22 | --- | --- | ---
51883 | C-S2-Zn-0 | 2010-08-23 | 17 | --- | 1.85 | --- | ---
51884 | C-S2-Zn-25 | 2010-08-23 | 17 | --- | 22.9 | --- | ---
51885 | C-S2-Zn-50 | 2010-08-23 | 17 | --- | 45.3 | --- | ---
51886 | C-S2-Zn-100 | 2010-08-23 | 17 | --- | 98.2 | --- | ---
51902 | C-S2-Zn-100 Duplicate | 2010-08-23 | 17 | --- | 99.5 | --- | ---
51887 | C-S2-Zn-200 | 2010-08-23 | 17 | --- | 205 | --- | ---
51888 | C-S2-Zn-400 | 2010-08-23 | 17 | --- | 438 | --- | ---
51889 | C-S2-Cd-0 | 2010-08-23 | 17 | --- | --- | < 0.033 | ---
51890 | C-S2-Cd-0.75 | 2010-08-23 | 17 | --- | --- | 0.71 | ---
51891 | C-S2-Cd-1.5 | 2010-08-23 | 17 | --- | --- | 1.34 | ---
51892 | C-S2-Cd-3.0 | 2010-08-23 | 17 | --- | --- | 2.78 | ---
51903 | C-S2-Cd-3.0 Duplicate | 2010-08-23 | 17 | --- | --- | 2.79 | ---
51893 | C-S2-Cd-6.0 | 2010-08-23 | 17 | --- | --- | 5.4 | ---
51894 | C-S2-Cd-12.0 | 2010-08-23 | 17 | --- | --- | 11.5 | ---
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51895 | C-S2-Pb-0 | 2010-08-23 | 17 | --- | --- | --- | 0.013
51896 | C-S2-Pb-5 | 2010-08-23 | 17 | --- | --- | --- | 3.08
51897 | C-S2-Pb-10 | 2010-08-23 | 17 | --- | --- | --- | 6.24
51898 | C-S2-Pb-20 | 2010-08-23 | 17 | --- | --- | --- | 14.4
51904 | C-S2-Pb-20 Duplicate | 2010-08-23 | 17 | --- | --- | --- | 13.5
51899 | C-S2-Pb-40 | 2010-08-23 | 17 | --- | --- | --- | 29.5
51900 | C-S2-Pb-80 | 2010-08-23 | 17 | --- | --- | --- | 59.1
51479 | C-Filter Blank | 2010-08-02 | 21 | < 0.044 | < 0.3 | < 0.018 | < 0.036
51480 | C-Cu-0 | 2010-08-02 | 21 | 0.19 | --- | --- | ---
51481 | C-Cu-0.5 | 2010-08-02 | 21 | 0.41 | --- | --- | ---
51482 | C-Cu-1.0 | 2010-08-02 | 21 | 0.88 | --- | --- | ---
51483 | C-Cu-2.0 | 2010-08-02 | 21 | 1.73 | --- | --- | ---
51504 | C-Cu-2.0 Duplicate | 2010-08-02 | 21 | 1.74 | --- | --- | ---
51484 | C-Cu-4.0 | 2010-08-02 | 21 | 3.31 | --- | --- | ---
51485 | C-Cu-8.0 | 2010-08-02 | 21 | 7.17 | --- | --- | ---
51486 | C-Zn-0 | 2010-08-02 | 21 | --- | 2.47 | --- | ---
51487 | C-Zn-25 | 2010-08-02 | 21 | --- | 22.1 | --- | ---
51488 | C-Zn-50 | 2010-08-02 | 21 | --- | 46.9 | --- | ---
51489 | C-Zn-100 | 2010-08-02 | 21 | --- | 93.9 | --- | ---
51505 | C-Zn-100 Duplicate | 2010-08-02 | 21 | --- | 90.4 | --- | ---
51490 | C-Zn-200 | 2010-08-02 | 21 | --- | 179 | --- | ---
51491 | C-Zn-400 | 2010-08-02 | 21 | --- | 388 | --- | ---
51492 | C-Cd-0 | 2010-08-02 | 21 | --- | --- | < 0.018 | ---
51493 | C-Cd-0.75 | 2010-08-02 | 21 | --- | --- | 0.6 | ---
51494 | C-Cd-1.5 | 2010-08-02 | 21 | --- | --- | 1.2 | ---
51495 | C-Cd-3.0 | 2010-08-02 | 21 | --- | --- | 2.66 | ---
51506 | C-Cd-3.0 Duplicate | 2010-08-02 | 21 | --- | --- | 2.68 | ---
51496 | C-Cd-6.0 | 2010-08-02 | 21 | --- | --- | 4.91 | ---
51497 | C-Cd-12.0 | 2010-08-02 | 21 | --- | --- | 11.2 | ---
51498 | C-Pb-0 | 2010-08-02 | 21 | --- | --- | --- | < 0.036
51499 | C-Pb-5 | 2010-08-02 | 21 | --- | --- | --- | 2.94
51500 | C-Pb-10 | 2010-08-02 | 21 | --- | --- | --- | 5.84
51501 | C-Pb-20 | 2010-08-02 | 21 | --- | --- | --- | 12.4
51507 | C-Pb-20 Duplicate | 2010-08-02 | 21 | --- | --- | --- | 12.2
51502 | C-Pb-40 | 2010-08-02 | 21 | --- | --- | --- | 27.2
51503 | C-Pb-80 | 2010-08-02 | 21 | --- | --- | --- | 60
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51978 | C-Filter Blank | 2010-08-30 | 24 | < 0.068 | < 0.56 | < 0.018 | < 0.013
51979 | C-S2-Cu-0 | 2010-08-30 | 24 | 0.67 | --- | --- | ---
51980 | C-S2-Cu-0.5 | 2010-08-30 | 24 | 0.87 | --- | --- | ---
51981 | C-S2-Cu-1.0 | 2010-08-30 | 24 | 1.8 | --- | --- | ---
51982 | C-S2-Cu-2.0 | 2010-08-30 | 24 | 2.04 | --- | --- | ---
52003 | C-S2-Cu-2.0 Duplicate | 2010-08-30 | 24 | 1.93 | --- | --- | ---
51983 | C-S2-Cu-4.0 | 2010-08-30 | 24 | 3.64 | --- | --- | ---
51984 | C-S2-Cu-8.0 | 2010-08-30 | 24 | 7.65 | --- | --- | ---
51985 | C-S2-Zn-0 | 2010-08-30 | 24 | --- | 1.41 | --- | ---
51986 | C-S2-Zn-25 | 2010-08-30 | 24 | --- | 24.1 | --- | ---
51987 | C-S2-Zn-50 | 2010-08-30 | 24 | --- | 47.9 | --- | ---
51988 | C-S2-Zn-100 | 2010-08-30 | 24 | --- | 94.7 | --- | ---
52004 | C-S2-Zn-100 Duplicate | 2010-08-30 | 24 | --- | 93.9 | --- | ---
51989 | C-S2-Zn-200 | 2010-08-30 | 24 | --- | 185 | --- | ---
51990 | C-S2-Zn-400 | 2010-08-30 | 24 | --- | 392 | --- | ---
51991 | C-S2-Cd-0 | 2010-08-30 | 24 | --- | --- | < 0.018 | ---
51992 | C-S2-Cd-0.75 | 2010-08-30 | 24 | --- | --- | 0.67 | ---
51993 | C-S2-Cd-1.5 | 2010-08-30 | 24 | --- | --- | 1.31 | ---
51994 | C-S2-Cd-3.0 | 2010-08-30 | 24 | --- | --- | 2.69 | ---
52005 | C-S2-Cd-3.0 Duplicate | 2010-08-30 | 24 | --- | --- | 2.69 | ---
51995 | C-S2-Cd-6.0 | 2010-08-30 | 24 | --- | --- | 5.59 | ---
51996 | C-S2-Cd-12.0 | 2010-08-30 | 24 | --- | --- | 12.5 | ---
51997 | C-S2-Pb-0 | 2010-08-30 | 24 | --- | --- | --- | < 0.013
51998 | C-S2-Pb-5 | 2010-08-30 | 24 | --- | --- | --- | 3.38
51999 | C-S2-Pb-10 | 2010-08-30 | 24 | --- | --- | --- | 6.58
52000 | C-S2-Pb-20 | 2010-08-30 | 24 | --- | --- | --- | 13.5
52006 | C-S2-Pb-20 Duplicate | 2010-08-30 | 24 | --- | --- | --- | 13.5
52001 | C-S2-Pb-40 | 2010-08-30 | 24 | --- | --- | --- | 28.8
52002 | C-S2-Pb-80 | 2010-08-30 | 24 | --- | --- | --- | 62.2
52011 | C-Filter Blank | 2010-09-02 | 27 | < 0.062 | < 0.9 | < 0.045 | < 0.008
52012 | C-S2-Cu-0 | 2010-09-02 | 27 | 0.24 | --- | --- | ---
52013 | C-S2-Cu-0.5 | 2010-09-02 | 27 | 0.55 | --- | --- | ---
52014 | C-S2-Cu-1.0 | 2010-09-02 | 27 | 1.02 | --- | --- | ---
52015 | C-S2-Cu-2.0 | 2010-09-02 | 27 | 1.65 | --- | --- | ---
52036 | C-S2-Cu-2.0 Duplicate | 2010-09-02 | 27 | 1.87 | --- | --- | ---
52016 | C-S2-Cu-4.0 | 2010-09-02 | 27 | 3.22 | --- | --- | ---
52017 | C-S2-Cu-8.0 | 2010-09-02 | 27 | 7.46 | --- | --- | ---
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
52018 | C-S2-Zn-0 | 2010-09-02 | 27 | --- | 0.86 | --- | ---
52019 | C-S2-Zn-25 | 2010-09-02 | 27 | --- | 25 | --- | ---
52020 | C-S2-Zn-50 | 2010-09-02 | 27 | --- | 49.5 | --- | ---
52021 | C-S2-Zn-100 | 2010-09-02 | 27 | --- | 96.7 | --- | ---
52037 | C-S2-Zn-100 Duplicate | 2010-09-02 | 27 | --- | 95.6 | --- | ---
52022 | C-S2-Zn-200 | 2010-09-02 | 27 | --- | 190 | --- | ---
52023 | C-S2-Zn-400 | 2010-09-02 | 27 | --- | 401 | --- | ---
52024 | C-S2-Cd-0 | 2010-09-02 | 27 | --- | --- | < 0.045 | ---
52025 | C-S2-Cd-0.75 | 2010-09-02 | 27 | --- | --- | 0.73 | ---
52026 | C-S2-Cd-1.5 | 2010-09-02 | 27 | --- | --- | 1.42 | ---
52027 | C-S2-Cd-3.0 | 2010-09-02 | 27 | --- | --- | 2.8 | ---
52038 | C-S2-Cd-3.0 Duplicate | 2010-09-02 | 27 | --- | --- | 2.72 | ---
52028 | C-S2-Cd-6.0 | 2010-09-02 | 27 | --- | --- | 5.47 | ---
52029 | C-S2-Cd-12.0 | 2010-09-02 | 27 | --- | --- | 11.9 | ---
52030 | C-S2-Pb-0 | 2010-09-02 | 27 | --- | --- | --- | < 0.008
52031 | C-S2-Pb-5 | 2010-09-02 | 27 | --- | --- | --- | 3.33
52032 | C-S2-Pb-10 | 2010-09-02 | 27 | --- | --- | --- | 6.34
52033 | C-S2-Pb-20 | 2010-09-02 | 27 | --- | --- | --- | 14.1
52039 | C-S2-Pb-20 Duplicate | 2010-09-02 | 27 | --- | --- | --- | 13.6
52034 | C-S2-Pb-40 | 2010-09-02 | 27 | --- | --- | --- | 28.8
52035 | C-S2-Pb-80 | 2010-09-02 | 27 | --- | --- | --- | 63.1
51559 | C-Cu-2ppb | 2010-08-02 | 28 | 1.84 | --- | --- | ---
51560 | C-Zn-100ppb | 2010-08-02 | 28 | --- | 92.1 | --- | ---
51561 | C-Cd-3ppb | 2010-08-02 | 28 | --- | --- | 2.48 | ---
51562 | C-Pb-20ppb | 2010-08-02 | 28 | --- | --- | --- | 12.1
51702 | C-Cu-2ppb | 2010-08-16 | 35 | 1.64 | --- | --- | ---
51703 | C-Zn-100ppb | 2010-08-16 | 35 | --- | 93.8 | --- | ---
51704 | C-Cd-3ppb | 2010-08-16 | 35 | --- | --- | 2.71 | ---
51705 | C-Pb-20ppb | 2010-08-16 | 35 | --- | --- | --- | 11.3
51905 | C-Cu-2ppb | 2010-08-23 | 42 | 1.83 | --- | --- | ---
51906 | C-Zn-100ppb | 2010-08-23 | 42 | --- | 97.4 | --- | ---
51907 | C-Cd-3ppb | 2010-08-23 | 42 | --- | --- | 2.83
51908 | C-Pb-20ppb | 2010-08-23 | 42 | --- | --- | --- | 12.1
52007 | C-Cu-2ppb | 2010-08-30 | 49 | 2.05 | --- | --- | ---
52008 | C-Zn-100ppb | 2010-08-30 | 49 | --- | 97.3 | --- | ---
52009 | C-Cd-3ppb | 2010-08-30 | 49 | --- | --- | 2.67 | ---
52010 | C-Pb-20ppb | 2010-08-30 | 49 | --- | --- | --- | 12.6
Table 3-1. Concentrations of elements in diluter water from chronic sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; Pb, lead; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd | Pb
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
52040 | C-Cu-2ppb | 2010-09-02 | 52 | 1.92 | --- | --- | ---
52041 | C-Zn-100ppb | 2010-09-02 | 52 | --- | 97.7 | --- | ---
52042 | C-Cd-3ppb | 2010-09-02 | 52 | --- | --- | 2.86 | ---
52043 | C-Pb-20ppb | 2010-09-02 | 52 | --- | --- | --- | 12.4

## Table 3-2 results
Table 3-2. Concentrations of elements in diluter water from sturgeon acute exposures.
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51217 | A-Filter Blank | 2010-07-16 | 0 | < 0.067 | < 1.5 | < 0.014
51218 | A-Cu-0ppb Control | 2010-07-12 | 0 | 0.2 | --- | ---
51239 | A-Cu-0ppb unfiltered | 2010-07-12 | 0 | 0.18 | --- | ---
51219 | A-Cu-1.56ppb Low | 2010-07-12 | 0 | 1.44 | --- | ---
51220 | A-Cu-3.125ppb Med-Low | 2010-07-12 | 0 | 2.58 | --- | ---
51240 | A-Cu Med-Low UF | 2010-07-12 | 0 | 2.68 | --- | ---
51221 | A-Cu-6.25ppb Med | 2010-07-12 | 0 | 5.23 | --- | ---
51222 | DUP-A-Cu-6.25ppb Med | 2010-07-12 | 0 | 5.15 | --- | ---
51223 | A-Cu-12.5ppb Med-Hi | 2010-07-12 | 0 | 10.3 | --- | ---
51224 | A-Cu-25ppb Hi | 2010-07-12 | 0 | 23.7 | --- | ---
51225 | A-Zn-0ppb | 2010-07-12 | 0 | --- | < 1.5 | ---
51241 | A-Zn 0ppb unfiltered | 2010-07-12 | 0 | --- | < 1.5 | ---
51226 | A-Zn-31.25ppb Low | 2010-07-12 | 0 | --- | 21.1 | ---
51227 | A-Zn-62.5ppb Med-Low | 2010-07-12 | 0 | --- | 36.3 | ---
51242 | A-Zn Med-Low UF | 2010-07-12 | 0 | --- | 28.3 | ---
51228 | A-Zn-125ppb Med | 2010-07-12 | 0 | --- | 66.8 | ---
51229 | DUP-A-Zn-125ppb Med | 2010-07-12 | 0 | --- | 67.2 | ---
51230 | A-Zn-250ppb Med-Hi | 2010-07-12 | 0 | --- | 113 | ---
51231 | A-Zn-500ppb Hi | 2010-07-12 | 0 | --- | 274 | ---
51232 | A-Cd-0ppb Control | 2010-07-12 | 0 | --- | --- | < 0.014
51243 | A-Cd 0ppb unfiltered | 2010-07-12 | 0 | --- | --- | < 0.014
51233 | A-Cd-3.125ppb Low | 2010-07-12 | 0 | --- | --- | 1.78
51234 | A-Cd-6.25ppb Med-Low | 2010-07-12 | 0 | --- | --- | 3.93
51244 | A-Cd Med-Low UF | 2010-07-12 | 0 | --- | --- | 4.11
51235 | A-Cd-12.5ppb Med | 2010-07-12 | 0 | --- | --- | 8.48
51236 | DUP-A-Cd-12.5ppb Med | 2010-07-12 | 0 | --- | --- | 8.76
51237 | A-Cd-25ppb Med-Hi | 2010-07-12 | 0 | --- | --- | 17.4
51238 | A-Cd-50ppb Hi | 2010-07-12 | 0 | --- | --- | 45.8
51375 | A-Filter Blank | 2010-07-26 | 0 | < 0.067 | < 1.5 | < 0.014
51376 | A-Cu-0ppb unfiltered | 2010-07-26 | 0 | 0.43 | --- | ---
51377 | A-Cu-0ppb Control | 2010-07-26 | 0 | 0.18 | --- | ---
51378 | A-Cu-4.7ppb Low | 2010-07-26 | 0 | 4.02 | --- | ---
51379 | A-Cu-9.4ppb Med-Low | 2010-07-26 | 0 | 7.81 | --- | ---
51380 | A-Cu-9.4ppb UF | 2010-07-26 | 0 | 8.1 | --- | ---
51381 | A-Cu-18.75ppb Med | 2010-07-26 | 0 | 15.6 | --- | ---
51382 | DUP-A-Cu-18.75ppb Med | 2010-07-26 | 0 | 15.8 | --- | ---
51383 | A-Cu-37.5ppb Med-Hi | 2010-07-26 | 0 | 31.8 | --- | ---
51384 | A-Cu-75ppb Hi | 2010-07-26 | 0 | 73.5 | --- | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51385 | A-Zn-0ppb UF | 2010-07-26 | 0 | --- | < 1.5 | ---
51386 | A-Zn-0ppb Control | 2010-07-26 | 0 | --- | < 1.5 | ---
51387 | A-Zn-93.75ppb Low | 2010-07-26 | 0 | --- | 77.2 | ---
51388 | A-Zn-187.5ppb Med-Low | 2010-07-26 | 0 | --- | 138 | ---
51389 | A-Zn-187.5ppb UF | 2010-07-26 | 0 | --- | 139 | ---
51390 | A-Zn-375ppb Med | 2010-07-26 | 0 | --- | 268 | ---
51391 | DUP-A-Zn-375ppb Med | 2010-07-26 | 0 | --- | 268 | ---
51392 | A-Zn-750ppb Med-Hi | 2010-07-26 | 0 | --- | 505 | ---
51393 | A-Zn-1500ppb Hi | 2010-07-26 | 0 | --- | 1510 | ---
51394 | A-Cd-0ppb UF | 2010-07-26 | 0 | --- | --- | < 0.014
51395 | A-Cd-0ppb Control | 2010-07-26 | 0 | --- | --- | < 0.014
51396 | A-Cd-12.5ppb Low | 2010-07-26 | 0 | --- | --- | 6.06
51397 | A-Cd-25ppb Med-Low | 2010-07-26 | 0 | --- | --- | 14.2
51398 | A-Cd-25ppb UF | 2010-07-26 | 0 | --- | --- | 14.5
51399 | A-Cd-50ppb Med | 2010-07-26 | 0 | --- | --- | 31.1
51400 | DUP-A-Cd-50ppb Med | 2010-07-26 | 0 | --- | --- | 31.1
51401 | A-Cd-100ppb Med-Hi | 2010-07-26 | 0 | --- | --- | 67.4
51402 | A-Cd-200ppb Hi | 2010-07-26 | 0 | --- | --- | 179
51403 | A-Cu-Stock 150mg/L | 2010-07-26 | --- | 149000 | --- | ---
51404 | A-Zn-Stock 3150mg/L | 2010-07-26 | --- | --- | 3190000 | ---
51405 | A-Cd-Stock 400mg/L | 2010-07-26 | --- | --- | --- | 388200
51508 | A-Filter Blank | 2010-08-09 | 0 | < 0.061 | < 0.4 | < 0.009
51509 | A-Cu-0ppb Control | 2010-08-09 | 0 | 0.18 | --- | ---
51510 | A-Cu-3.125ppb Low | 2010-08-09 | 0 | 2.13 | --- | ---
51511 | A-Cu-6.25ppb Med-Low | 2010-08-09 | 0 | 4.71 | --- | ---
51512 | A-Cu-12.5ppb Med | 2010-08-09 | 0 | 9.35 | --- | ---
51513 | DUP-A-Cu-12.5ppb Med | 2010-08-09 | 0 | 9.43 | --- | ---
51514 | A-Cu-25ppb Med-Hi | 2010-08-09 | 0 | 19 | --- | ---
51515 | A-Cu-50ppb Hi | 2010-08-09 | 0 | 44.8 | --- | ---
51516 | A-Zn 0ppb Control | 2010-08-09 | 0 | --- | 1.5 | ---
51517 | A-Zn 156.25ppb Low | 2010-08-09 | 0 | --- | 102 | ---
51518 | A-Zn 312.5ppb Med-Low | 2010-08-09 | 0 | --- | 206 | ---
51519 | A-Zn 625ppb Med | 2010-08-09 | 0 | --- | 411 | ---
51520 | DUP-A-Zn 625ppb Med | 2010-08-09 | 0 | --- | 417 | ---
51521 | A-Zn 1250ppb Med-Hi | 2010-08-09 | 0 | --- | 824 | ---
51522 | A-Zn-2500ppb Hi | 2010-08-09 | 0 | --- | 2490 | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51523 | A-Cd 0ppb Control | 2010-08-09 | 0 | --- | --- | < 0.009
51524 | A-Cd 25ppb Low | 2010-08-09 | 0 | --- | --- | 14
51525 | A-Cd 50ppb Med-Low | 2010-08-09 | 0 | --- | --- | 30.5
51526 | A-Cd 100ppb Med | 2010-08-09 | 0 | --- | --- | 65.5
51527 | DUP-A-Cd 100ppb Med | 2010-08-09 | 0 | --- | --- | 64.6
51528 | A-Cd 200ppb Med-Hi | 2010-08-09 | 0 | --- | --- | 168
51529 | A-Cd 400ppb Hi | 2010-08-09 | 0 | --- | --- | 337
52044 | A-Filter Blank | 2010-09-09 | 0 | < 0.068 | < 0.6 | < 0.018
52045 | A-Cu-0ppb Control | 2010-09-09 | 0 | 0.15 | --- | ---
52046 | A-Cu-12.5ppb Low | 2010-09-09 | 0 | 7.25 | --- | ---
52047 | A-Cu-25ppb Med-Low | 2010-09-09 | 0 | 14.3 | --- | ---
52048 | A-Cu-50ppb Med | 2010-09-09 | 0 | 43.6 | --- | ---
52049 | DUP-A-Cu-50ppb Med | 2010-09-09 | 0 | 43.6 | --- | ---
52050 | A-Cu-100ppb Med-Hi | 2010-09-09 | 0 | 89.2 | --- | ---
52051 | A-Cu-200ppb High | 2010-09-09 | 0 | 130 | --- | ---
52052 | A-Zn-0ppb Control | 2010-09-09 | 0 | --- | 0.64 | ---
52053 | A-Zn-312ppb Low | 2010-09-09 | 0 | --- | 260 | ---
52054 | A-Zn-625ppb Med-Low | 2010-09-09 | 0 | --- | 542 | ---
52055 | A-Zn-1250ppb Med | 2010-09-09 | 0 | --- | 1070 | ---
52056 | DUP-A-Zn-1250ppb Med | 2010-09-09 | 0 | --- | 1060 | ---
52057 | A-Zn-2500ppb Med-Hi | 2010-09-09 | 0 | --- | 2170 | ---
52058 | A-Zn-5000ppb High | 2010-09-09 | 0 | --- | 4380 | ---
52059 | A-Cd-0ppb Control | 2010-09-09 | 0 | --- | --- | < 0.018
52060 | A-Cd-37.5ppb Low | 2010-09-09 | 0 | --- | --- | 32.7
52061 | A-Cd-75ppb Med-Low | 2010-09-09 | 0 | --- | --- | 59.3
52062 | A-Cd-150ppb Med | 2010-09-09 | 0 | --- | --- | 134
52063 | DUP-A-Cd-150ppb Med | 2010-09-09 | 0 | --- | --- | 131
52064 | A-Cd-300ppb Med-Hi | 2010-09-09 | 0 | --- | --- | 268
52065 | A-Cd-600ppb High | 2010-09-09 | 0 | --- | --- | 565
52066 | A-Cu-Stock 740mg/L | 2010-09-09 | 0 | 748000 | --- | ---
52067 | A-Zn-Stock 3750mg/L | 2010-09-09 | 0 | --- | 3640000 | ---
52068 | A-Cd-Stock 456mg/L | 2010-09-09 | 0 | --- | --- | 456000
52069 | A-BE-Cu-Stock 389.2mg/L | 2010-09-09 | 0 | 387000 | --- | ---
52070 | A-BE-Cu-0ppb Control | 2010-09-09 | 0 | 0.16 | --- | ---
52071 | A-BE-Cu-12.5ppb Low | 2010-09-09 | 0 | 10.7 | --- | ---
52072 | A-BE-Cu-50ppb Med | 2010-09-09 | 0 | 44.4 | --- | ---
52073 | DUP-A-BE-Cu 50ppb Med | 2010-09-09 | 0 | 43.7 | --- | ---
52074 | A-BE-Cu 200ppb High | 2010-09-09 | 0 | 201 | --- | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
52267 | A-Filter Blank | 2010-09-20 | 0 | < 0.11 | < 0.3 | < 0.027
52268 | A-Cu-0ppb Control | 2010-09-20 | 0 | 0.17 | --- | ---
52269 | A-Cu-12.5ppb Low | 2010-09-20 | 0 | 7.51 | --- | ---
52270 | A-Cu-25ppb Med-Low | 2010-09-20 | 0 | 21.4 | --- | ---
52271 | A-Cu-50ppb Med | 2010-09-20 | 0 | 44 | --- | ---
52272 | A-DUP-Cu-50ppb Med | 2010-09-20 | 0 | 43.6 | --- | ---
52273 | A-Cu-100ppb Med-Hi | 2010-09-20 | 0 | 82.2 | --- | ---
52274 | A-Cu-200ppb High | 2010-09-20 | 0 | 183 | --- | ---
52275 | A-Zn-0ppb Control | 2010-09-20 | 0 | --- | 0.57 | ---
52276 | A-Zn-468.75ppb Low | 2010-09-20 | 0 | --- | 382 | ---
52277 | A-Zn-937.5ppb Med-Low | 2010-09-20 | 0 | --- | 790 | ---
52278 | A-Zn-1875ppb Med | 2010-09-20 | 0 | --- | 1640 | ---
52279 | A-DUP-Zn-1875ppb Med | 2010-09-20 | 0 | --- | 1680 | ---
52280 | A-Zn-3750ppb Med-Hi | 2010-09-20 | 0 | --- | 3240 | ---
52281 | A-Zn-7500ppb High | 2010-09-20 | 0 | --- | 5990 | ---
52282 | A-Cd-0ppb Control | 2010-09-20 | 0 | --- | --- | < 0.027
52283 | A-Cd-9.375 Low | 2010-09-20 | 0 | --- | --- | 8.99
52284 | A-Cd-18.75ppb Med-Low | 2010-09-20 | 0 | --- | --- | 17.9
52285 | A-Cd-37.5ppb Med | 2010-09-20 | 0 | --- | --- | 37.5
52286 | A-DUP-Cd-37.5ppb Med | 2010-09-20 | 0 | --- | --- | 37
52287 | A-Cd-75ppb Med-Hi | 2010-09-20 | 0 | --- | --- | 72.1
52288 | A-Cd-150ppb High | 2010-09-20 | 0 | --- | --- | 151
52289 | A-Zn-Stock 5625mg/L | 2010-09-20 | 0 | --- | 5740000 | ---
52290 | A-Cd-Stock 570mg/L | 2010-09-20 | 0 | --- | --- | 568000
52471 | A-Filter Blank | 2010-10-07 | 0 | < 0.05 | < 0.17 | < 0.017
52472 | A-Cu-0ppb Control | 2010-10-07 | 0 | 0.19 | --- | ---
52473 | A-Cu-18.75ppb Low | 2010-10-07 | 0 | 12.5 | --- | ---
52474 | A-Cu-37.5ppb Med-Low | 2010-10-07 | 0 | 29.5 | --- | ---
52475 | A-Cu-75ppb Med | 2010-10-07 | 0 | 61.5 | --- | ---
52476 | A-DUP-Cu-75ppb Med | 2010-10-07 | 0 | 60.1 | --- | ---
52477 | A-Cu-150ppb Med-Hi | 2010-10-07 | 0 | 116 | --- | ---
52478 | A-Cu-300ppb High | 2010-10-07 | 0 | 256 | --- | ---
52479 | A-Zn-0ppb Control | 2010-10-07 | 0 | --- | 0.57 | ---
52480 | A-Zn-625ppb Low | 2010-10-07 | 0 | --- | 579 | ---
52481 | A-Zn-1250ppb Med-Low | 2010-10-07 | 0 | --- | 1190 | ---
52482 | A-Zn-2500ppb Med | 2010-10-07 | 0 | --- | 2380 | ---
52483 | A-DUP-Zn-2500ppb Med | 2010-10-07 | 0 | --- | 2400 | ---
52484 | A-Zn-5000ppb Med-High | 2010-10-07 | 0 | --- | 4520 | ---
52485 | A-Zn-10000ppb High | 2010-10-07 | 0 | --- | 9330 | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
52486 | A-Cd-0ppb Control | 2010-10-07 | 0 | --- | --- | < 0.017
52487 | A-Cd-18.75ppb Low | 2010-10-07 | 0 | --- | --- | 15.9
52488 | A-Cd-37.5ppb Med-Low | 2010-10-07 | 0 | --- | --- | 31.1
52489 | A-Cd-75ppb Med | 2010-10-07 | 0 | --- | --- | 66.2
52490 | A-DUP-Cd-75ppb Med | 2010-10-07 | 0 | --- | --- | 67.5
52491 | A-Cd-150ppb Med-Hi | 2010-10-07 | 0 | --- | --- | 112
52492 | A-Cd-300ppb High | 2010-10-07 | 0 | --- | --- | 258
52493 | A-Cu-Stock 1110mg/L | 2010-10-07 | 0 | 1090000 | --- | ---
52494 | A-Cd-Stock 1140mg/L | 2010-10-07 | 0 | --- | --- | 1130000
52495 | A-Zn-Stock 7500mg/L | 2010-10-07 | 0 | --- | 7370000 | ---
51298 | A-Filter Blank | 2010-07-16 | 4 | < 0.077 | < 0.8 | < 0.013
51299 | A-Cu-0ppb unfiltered | 2010-07-16 | 4 | 0.16 | --- | ---
51300 | A-Cu-0ppb Control | 2010-07-16 | 4 | 0.17 | --- | ---
51301 | A-Cu-1.56ppb Low | 2010-07-16 | 4 | 1.41 | --- | ---
51302 | A-Cu-Med-Low UF | 2010-07-16 | 4 | 2.82 | --- | ---
51303 | A-Cu-3.125ppb Med-Low | 2010-07-16 | 4 | 2.74 | --- | ---
51304 | A-Cu-6.25ppb Med | 2010-07-16 | 4 | 5.34 | --- | ---
51305 | DUP-A-Cu-6.25ppb Med | 2010-07-16 | 4 | 5.45 | --- | ---
51306 | A-Cu-12.5ppb Med-Hi | 2010-07-16 | 4 | 10.7 | --- | ---
51307 | A-Cu-25ppb Hi | 2010-07-16 | 4 | 23.4 | --- | ---
51308 | A-Zn-0ppb unfiltered | 2010-07-16 | 4 | --- | 0.81 | ---
51309 | A-Zn-0ppb Control | 2010-07-16 | 4 | --- | 1.29 | ---
51310 | A-Zn-31.25ppb Low | 2010-07-16 | 4 | --- | 29.9 | ---
51311 | A-Zn-62.5ppb Med-Low | 2010-07-16 | 4 | --- | 59.5 | ---
51312 | A-Zn-Med-Low UF | 2010-07-16 | 4 | --- | 59.2 | ---
51313 | A-Zn-125ppb Med | 2010-07-16 | 4 | --- | 116 | ---
51314 | DUP-A-Zn-125ppb Med | 2010-07-16 | 4 | --- | 117 | ---
51315 | A-Zn-250ppb Med-Hi | 2010-07-16 | 4 | --- | 225 | ---
51316 | A-Zn-500ppb Hi | 2010-07-16 | 4 | --- | 634 | ---
51317 | A-Cd-0ppb unfiltered | 2010-07-16 | 4 | --- | --- | < 0.013
51318 | A-Cd-0ppb Control | 2010-07-16 | 4 | --- | --- | < 0.013
51319 | A-Cd-3.125ppb Low | 2010-07-16 | 4 | --- | --- | 2.42
51320 | A-Cd-6.25ppb Med-Low | 2010-07-16 | 4 | --- | --- | 5.04
51321 | A-Cd Med-Low UF | 2010-07-16 | 4 | --- | --- | 5.05
51322 | A-Cd 12.5ppb Med | 2010-07-16 | 4 | --- | --- | 10.9
51323 | DUP-A-Cd 12.5ppb Med | 2010-07-16 | 4 | --- | --- | 10.5
51324 | A-Cd 25ppb Med-Hi | 2010-07-16 | 4 | --- | --- | 21.6
51325 | A-Cd 50ppb Hi | 2010-07-16 | 4 | --- | --- | 48.6
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51451 | A-Filter Blank | 2010-07-30 | 4 | < 0.044 | < 0.3 | < 0.018
51452 | A-Cu-0ppb unfiltered | 2010-07-30 | 4 | 0.18 | --- | ---
51453 | A-Cu-0ppb Control | 2010-07-30 | 4 | 0.19 | --- | ---
51454 | A-Cu-4.7ppb Low | 2010-07-30 | 4 | 4.59 | --- | ---
51455 | A-Cu-9.4ppb Med-Low | 2010-07-30 | 4 | 9.02 | --- | ---
51456 | A-Cu-9.4ppb Med-Low UF | 2010-07-30 | 4 | 9.07 | --- | ---
51457 | A-Cu-18.75ppb Med | 2010-07-30 | 4 | 18.1 | --- | ---
51458 | DUP-A-Cu-18.75ppb Med | 2010-07-30 | 4 | 17 | --- | ---
51459 | A-Cu-37.5ppb Med-Hi | 2010-07-30 | 4 | 35.4 | --- | ---
51460 | A-Cu-75ppb Hi | 2010-07-30 | 4 | 80 | --- | ---
51461 | A-Zn-0ppb unfiltered | 2010-07-30 | 4 | --- | 2.1 | ---
51462 | A-Zn-0ppb Control | 2010-07-30 | 4 | --- | 1.3 | ---
51463 | A-Zn-93.75ppb Low | 2010-07-30 | 4 | --- | 86.8 | ---
51464 | A-Zn-187.5ppb Med-Low | 2010-07-30 | 4 | --- | 175 | ---
51465 | A-Zn-187.5ppb Med-Low UF | 2010-07-30 | 4 | --- | 171 | ---
51466 | A-Zn-375ppb Med | 2010-07-30 | 4 | --- | 331 | ---
51467 | DUP-A-Zn-375ppb Med | 2010-07-30 | 4 | --- | 332 | ---
51468 | A-Zn-750ppb Med-Hi | 2010-07-30 | 4 | --- | 611 | ---
51469 | A-Zn-1500ppb Hi | 2010-07-30 | 4 | --- | 1640 | ---
51470 | A-Cd-0ppb unfiltered | 2010-07-30 | 4 | --- | --- | < 0.018
51471 | A-Cd-0ppb Control | 2010-07-30 | 4 | --- | --- | < 0.018
51472 | A-Cd-12.5ppb Low | 2010-07-30 | 4 | --- | --- | 10.1
51473 | A-Cd-25ppb Med-Low | 2010-07-30 | 4 | --- | --- | 21.6
51474 | A-Cd-25ppb Med-Low UF | 2010-07-30 | 4 | --- | --- | 21.5
51475 | A-Cd 50ppb Med | 2010-07-30 | 4 | --- | --- | 44.4
51476 | DUP-A-Cd 50ppb Med | 2010-07-30 | 4 | --- | --- | 43.6
51477 | A-Cd 100ppb Med-Hi | 2010-07-30 | 4 | --- | --- | 92.5
51478 | A-Cd 200ppb Hi | 2010-07-30 | 4 | --- | --- | 194
51642 | A-Filter Blank | 2010-08-13 | 4 | < 0.034 | < 0.8 | < 0.01
51643 | A-Cu-0ppb unfiltered | 2010-08-13 | 4 | 0.22 | --- | ---
51644 | A-Cu-0ppb Control | 2010-08-13 | 4 | 0.17 | --- | ---
51645 | A-Cu-3.125ppb Low | 2010-08-13 | 4 | 2.87 | --- | ---
51646 | A-Cu-6.25ppb Med-Low | 2010-08-13 | 4 | 5.6 | --- | ---
51647 | A-Cu-6.25ppb Med-Low UF | 2010-08-13 | 4 | 5.85 | --- | ---
51648 | A-Cu-12.5ppb Med | 2010-08-13 | 4 | 11.5 | --- | ---
51649 | DUP-A-Cu-12.5ppb Med | 2010-08-13 | 4 | 11.6 | --- | ---
51650 | A-Cu-25ppb Med-Hi | 2010-08-13 | 4 | 23.9 | --- | ---
51651 | A-Cu-50ppb Hi | 2010-08-13 | 4 | 52.6 | --- | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51652 | A-Zn-0ppb unfiltered | 2010-08-13 | 4 | --- | 1.48 | ---
51653 | A-Zn-0ppb Control | 2010-08-13 | 4 | --- | 1.37 | ---
51654 | A-Zn-156.25ppb Low | 2010-08-13 | 4 | --- | 137 | ---
51655 | A-Zn-312.5ppb Med-Low | 2010-08-13 | 4 | --- | 277 | ---
51656 | A-Zn-312.5ppb Med-Low UF | 2010-08-13 | 4 | --- | 269 | ---
51657 | A-Zn-625ppb Med | 2010-08-13 | 4 | --- | 517 | ---
51658 | DUP-A-Zn-625ppb Med | 2010-08-13 | 4 | --- | 521 | ---
51659 | A-Zn-1250ppb Med-Hi | 2010-08-13 | 4 | --- | 976 | ---
51660 | A-Zn-2500ppb Hi | 2010-08-13 | 4 | --- | 2640 | ---
51661 | A-Cd-0ppb unfiltered | 2010-08-13 | 4 | --- | --- | < 0.01
51662 | A-Cd-0ppb Control | 2010-08-13 | 4 | --- | --- | < 0.01
51663 | A-Cd-25ppb Low | 2010-08-13 | 4 | --- | --- | 19.5
51664 | A-Cd-50ppb Med-Low | 2010-08-13 | 4 | --- | --- | 40.2
51665 | A-Cd-50ppb Med-Low UF | 2010-08-13 | 4 | --- | --- | 39.6
51666 | A-Cd 100ppb Med | 2010-08-13 | 4 | --- | --- | 85.4
51667 | DUP-A-Cd 100ppb Med | 2010-08-13 | 4 | --- | --- | 84.8
51668 | A-Cd 200ppb Med-Hi | 2010-08-13 | 4 | --- | --- | 167
51669 | A-Cd 400ppb Hi | 2010-08-13 | 4 | --- | --- | 373
51670 | A-Zn-Stock 1050mg/L | 2010-08-13 | 4 | --- | 1100000 | ---
51671 | A-Cu-Stock 97.3mg/L | 2010-08-13 | 4 | 94000 | --- | ---
51672 | A-Cd-Stock 759.2mg/L | 2010-08-13 | 4 | --- | --- | 749000
51950 | A-Filter Blank | 2010-08-27 | 4 | < 0.062 | < 0.9 | < 0.045
51951 | A-Cu-0ppb unfiltered | 2010-08-27 | 4 | 0.14 | --- | ---
51952 | A-Cu-0ppb Control | 2010-08-27 | 4 | 0.14 | --- | ---
51953 | A-Cu-3.125ppb Low | 2010-08-27 | 4 | 2.84 | --- | ---
51954 | A-Cu-6.25ppb Med-Low | 2010-08-27 | 4 | 5.4 | --- | ---
51955 | A-Cu-6.25ppb Med-Low UF | 2010-08-27 | 4 | 5.62 | --- | ---
51956 | A-Cu-12.5ppb Med | 2010-08-27 | 4 | 10.7 | --- | ---
51957 | DUP-A-Cu-12.5ppb Med | 2010-08-27 | 4 | 11.1 | --- | ---
51958 | A-Cu-25ppb Med-Hi | 2010-08-27 | 4 | 22.6 | --- | ---
51959 | A-Cu-50ppb Hi | 2010-08-27 | 4 | 49.8 | --- | ---
51960 | A-Zn-0ppb unfiltered | 2010-08-27 | 4 | --- | < 0.9 | ---
51961 | A-Zn-0ppb Control | 2010-08-27 | 4 | --- | 2.18 | ---
51962 | A-Zn-156.25ppb Low | 2010-08-27 | 4 | --- | 128 | ---
51963 | A-Zn-312.5ppb Med-Low | 2010-08-27 | 4 | --- | 261 | ---
51964 | A-Zn-312.5ppb Med-Low UF | 2010-08-27 | 4 | --- | 264 | ---
51965 | A-Zn-625ppb Med | 2010-08-27 | 4 | --- | 511 | ---
51966 | DUP-A-Zn-625ppb Med | 2010-08-27 | 4 | --- | 506 | ---
51967 | A-Zn-1250ppb Med-Hi | 2010-08-27 | 4 | --- | 968 | ---
51968 | A-Zn-2500ppb Hi | 2010-08-27 | 4 | --- | 2610 | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51969 | A-Cd-0ppb unfiltered | 2010-08-27 | 4 | --- | --- | < 0.045
51970 | A-Cd-0ppb Control | 2010-08-27 | 4 | --- | --- | < 0.045
51971 | A-Cd-25ppb Low | 2010-08-27 | 4 | --- | --- | 19.7
51972 | A-Cd-50ppb Med-Low | 2010-08-27 | 4 | --- | --- | 40.4
51973 | A-Cd-50ppb Med-Low UF | 2010-08-27 | 4 | --- | --- | 40.4
51974 | A-Cd 100ppb Med | 2010-08-27 | 4 | --- | --- | 80.4
51975 | DUP-A-Cd 100ppb Med | 2010-08-27 | 4 | --- | --- | 80.9
51976 | A-Cd 200ppb Med-Hi | 2010-08-27 | 4 | --- | --- | 167
51977 | A-Cd 400ppb Hi | 2010-08-27 | 4 | --- | --- | 368
52114 | A-Filter Blank | 2010-09-13 | 4 | < 0.11 | < 0.3 | < 0.027
52115 | A-Cu-0ppb Control | 2010-09-13 | 4 | 0.15 | --- | ---
52116 | A-Cu-0ppb Control UF | 2010-09-13 | 4 | 0.18 | --- | ---
52117 | A-Cu-12.5ppb Low | 2010-09-13 | 4 | 7.66 | --- | ---
52118 | A-Cu-25ppb Med-Low | 2010-09-13 | 4 | 22 | --- | ---
52119 | Cu-25ppb Med-Low UF | 2010-09-13 | 4 | 22.9 | --- | ---
52120 | Cu-50ppb Med | 2010-09-13 | 4 | 44.5 | --- | ---
52121 | A-DUP-Cu-50ppb Med | 2010-09-13 | 4 | 45.8 | --- | ---
52122 | A-Cu-100ppb Med-Hi | 2010-09-13 | 4 | 91.1 | --- | ---
52123 | A-Zn-0ppb Control | 2010-09-13 | 4 | --- | 0.35 | ---
52124 | A-Zn-0ppb Control UF | 2010-09-13 | 4 | --- | 0.62 | ---
52125 | A-Zn-312.5ppb Low | 2010-09-13 | 4 | --- | 246 | ---
52126 | A-Zn-625ppb Med-Low | 2010-09-13 | 4 | --- | 536 | ---
52127 | A-Zn-625ppb Med-Low UF | 2010-09-13 | 4 | --- | 519 | ---
52128 | A-Zn-1250ppb Med | 2010-09-13 | 4 | --- | 1070 | ---
52129 | A-DUP-Zn-1250ppb Med | 2010-09-13 | 4 | --- | 1050 | ---
52130 | A-Zn-2500ppb Med-Hi | 2010-09-13 | 4 | --- | 2080 | ---
52131 | A-Zn-5000ppb Hi | 2010-09-13 | 4 | --- | 4210 | ---
52132 | A-Cd-0ppb Control | 2010-09-13 | 4 | --- | --- | < 0.027
52133 | A-Cd-0ppb Control UF | 2010-09-13 | 4 | --- | --- | < 0.027
52134 | A-Cd-37.5ppb Low | 2010-09-13 | 4 | --- | --- | 36.1
52135 | A-Cd-75ppb Med-Low | 2010-09-13 | 4 | --- | --- | 67.6
52136 | A-Cd-75ppb Med-Low UF | 2010-09-13 | 4 | --- | --- | 67.9
52137 | A-Cd-150ppb Med | 2010-09-13 | 4 | --- | --- | 141
52138 | A-DUP-Cd-150ppb Med | 2010-09-13 | 4 | --- | --- | 140
52138 | A-DUP-Cd-150ppb Med | 2010-09-13 | 4 | --- | --- | 148
52139 | A-Cd-300ppb Med-Hi | 2010-09-13 | 4 | --- | --- | 284
52140 | A-Cd 600ppb High | 2010-09-13 | 4 | --- | --- | 585
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
52141 | A-BE-Cu 0ppb Control | 2010-09-13 | 4 | 0.14 | --- | ---
52142 | A-BE-Cu 0ppb Control UF | 2010-09-13 | 4 | 0.17 | --- | ---
52143 | A-BE-Cu-12.5 Low | 2010-09-13 | 4 | 10.6 | --- | ---
52144 | A-BE-Cu-12.5 Low UF | 2010-09-13 | 4 | 11.1 | --- | ---
52145 | A-BE-Cu-50 Med | 2010-09-13 | 4 | 44 | --- | ---
52146 | A-DUP-BE-Cu-50 Med | 2010-09-13 | 4 | 44.1 | --- | ---
52351 | A-Filter Blank Day 4 | 2010-09-24 | 4 | < 0.054 | < 0.25 | < 0.01
52352 | A-Cu-0ppb Control UF | 2010-09-24 | 4 | 0.16 | --- | ---
52353 | A-Cu-0ppb Control | 2010-09-24 | 4 | 0.16 | --- | ---
52354 | A-Cu-12.5ppb Low | 2010-09-24 | 4 | 8.61 | --- | ---
52355 | A-Cu-25ppb Med-Low UF | 2010-09-24 | 4 | 23.4 | --- | ---
52356 | Cu-25ppb Med-Low | 2010-09-24 | 4 | 22.3 | --- | ---
52357 | Cu-50ppb Med | 2010-09-24 | 4 | 44.5 | --- | ---
52358 | A-DUP-Cu-50ppb Med | 2010-09-24 | 4 | 44.5 | --- | ---
52359 | A-Cu-100ppb Med-Hi | 2010-09-24 | 4 | 90.6 | --- | ---
52360 | A-Zn-0ppb Control UF | 2010-09-24 | 4 | --- | 0.51 | ---
52361 | A-Zn-0ppb Control | 2010-09-24 | 4 | --- | 0.42 | ---
52362 | A-Zn-468.75ppb Low | 2010-09-24 | 4 | --- | 400 | ---
52363 | A-Zn-937.55ppb Med-Low UF | 2010-09-24 | 4 | --- | 829 | ---
52364 | A-Zn-937.5ppb Med-Low | 2010-09-24 | 4 | --- | 829 | ---
52365 | A-Zn-1875ppb Med | 2010-09-24 | 4 | --- | 1650 | ---
52366 | A-DUP-Zn-1875ppb Med | 2010-09-24 | 4 | --- | 1660 | ---
52367 | A-Zn-3750ppb Med-Hi | 2010-09-24 | 4 | --- | 3310 | ---
52368 | A-Zn-7500ppb Hi | 2010-09-24 | 4 | --- | 6740 | ---
52369 | A-Cd-0ppb Control UF | 2010-09-24 | 4 | --- | --- | < 0.01
52370 | A-Cd-0ppb Control | 2010-09-24 | 4 | --- | --- | < 0.01
52371 | A-Cd-9.375ppb Low | 2010-09-24 | 4 | --- | --- | 8.99
52372 | A-Cd-18.75ppb Med-Low UF | 2010-09-24 | 4 | --- | --- | 17.7
52373 | A-Cd-18.75ppb Med-Low | 2010-09-24 | 4 | --- | --- | 17.5
52374 | A-Cd-37.5ppb Med | 2010-09-24 | 4 | --- | --- | 36.4
52375 | A-DUP-Cd-37.5ppb Med | 2010-09-24 | 4 | --- | --- | 36
52376 | A-Cd-75ppb Med-Hi | 2010-09-24 | 4 | --- | --- | 71.6
52377 | A-Cd 150ppb High | 2010-09-24 | 4 | --- | --- | 148
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
52496 | A-Filter Blank | 2010-10-11 | 4 | < 0.05 | 0.19 | < 0.017
52497 | A-Cu-0ppb Control | 2010-10-11 | 4 | 0.12 | --- | ---
52498 | A-Cu 0ppb Control UF | 2010-10-11 | 4 | 0.19 | --- | ---
52499 | A-Cu 18.75ppb Low | 2010-10-11 | 4 | 13.8 | --- | ---
52500 | A-Cu-37.5 Med-Low | 2010-10-11 | 4 | 32 | --- | ---
52501 | A-Cu-37.5 Med-Low UF | 2010-10-11 | 4 | 33.6 | --- | ---
52502 | A-Cu-75 Med | 2010-10-11 | 4 | 64.5 | --- | ---
52503 | A-DUP-Cu-75 Med | 2010-10-11 | 4 | 65.3 | --- | ---
52504 | A-Cu-150ppb Med-Hi | 2010-10-11 | 4 | 132 | --- | ---
52505 | A-Zn-0ppb Control | 2010-10-11 | 4 | --- | 0.99 | ---
52506 | A-Zn-0ppb Control UF | 2010-10-11 | 4 | --- | 2.14 | ---
52507 | A-Zn-625ppb Low | 2010-10-11 | 4 | --- | 592 | ---
52508 | A-Zn-1250ppb Med-Low | 2010-10-11 | 4 | --- | 1230 | ---
52509 | A-Zn-1250ppb Med-Low UF | 2010-10-11 | 4 | --- | 1250 | ---
52510 | A-Zn-2500ppb Med | 2010-10-11 | 4 | --- | 2450 | ---
52511 | A-DUP-Zn-2500ppb Med | 2010-10-11 | 4 | --- | 2480 | ---
52512 | A-Zn-5000ppb Med-Hi | 2010-10-11 | 4 | --- | 4910 | ---
52513 | A-Cd-0ppb Control | 2010-10-11 | 4 | --- | --- | < 0.017
52514 | A-Cd-0ppb Control UF | 2010-10-11 | 4 | --- | --- | 0.041
52515 | A-Cd-18.75ppb Low | 2010-10-11 | 4 | --- | --- | 18.5
52516 | A-Cd-37.5ppb Med-Low | 2010-10-11 | 4 | --- | --- | 35.3
52517 | A-Cd-37.5ppb Med-Low UF | 2010-10-11 | 4 | --- | --- | 36.9
52518 | A-Cd-75ppb Med | 2010-10-11 | 4 | --- | --- | 72.4
52519 | A-DUP-Cd-75ppb Med | 2010-10-11 | 4 | --- | --- | 70.8
52520 | A-Cd-150ppb Med-Hi | 2010-10-11 | 4 | --- | --- | 139
52521 | A-Cd-300ppb High | 2010-10-11 | 4 | --- | --- | 289
51851 | A-Filter Blank | 2010-08-23 | --- | < 0.084 | < 1.2 | < 0.033
51852 | A-Cu-0ppb Control | 2010-08-23 | --- | 0.18 | --- | ---
51853 | A-Cu-3.125ppb Low | 2010-08-23 | --- | 1.67 | --- | ---
51854 | A-Cu-6.25ppb Med-Low | 2010-08-23 | --- | 3.25 | --- | ---
51855 | A-Cu-12.5ppb Med | 2010-08-23 | --- | 6.25 | --- | ---
51856 | DUP-A-Cu-12.5ppb Med | 2010-08-23 | --- | 6.31 | --- | ---
51857 | A-Cu-25ppb Med-Hi | 2010-08-23 | --- | 12.1 | --- | ---
51858 | A-Cu-50ppb Hi | 2010-08-23 | --- | 26.4 | --- | ---
51859 | A-Zn-0ppb Control | 2010-08-23 | --- | --- | 2.28 | ---
51860 | A-Zn-156.25ppb Low | 2010-08-23 | --- | --- | 69.3 | ---
51861 | A-Zn-312.5ppb Med-Low | 2010-08-23 | --- | --- | 155 | ---
51862 | A-Zn-625ppb Med | 2010-08-23 | --- | --- | 311 | ---
51863 | DUP-A-Zn-625ppb Med | 2010-08-23 | --- | --- | 317 | ---
51864 | A-Zn-1250ppb Med-Hi | 2010-08-23 | --- | --- | 610 | ---
51865 | A-Zn-2500ppb Hi | 2010-08-23 | --- | --- | 2040 | ---
Table 3-2. Concentrations of elements in diluter water from acute sturgeon exposures. --- Continued
[ID; identification, Cu, copper; Zn, zinc; Cd, cadmium; ng, nanograms; mL, milliliter. Bold and italicized values are less than the method quantitation limit and have greater relative uncertainty]
Sample |  | Collection | Test | Cu | Zn | Cd
Number | Sample ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL)
51866 | A-Cd-0ppb Control | 2010-08-23 | --- | --- | --- | < 0.033
51867 | A-Cd-25ppb Low | 2010-08-23 | --- | --- | --- | 9.21
51868 | A-Cd-50ppb Med-Low | 2010-08-23 | --- | --- | --- | 21.3
51869 | A-Cd 100ppb Med | 2010-08-23 | --- | --- | --- | 46.9
51870 | DUP-A-Cd 100ppb Med | 2010-08-23 | --- | --- | --- | 47.1
51871 | A-Cd 200ppb Med-Hi | 2010-08-23 | --- | --- | --- | 101
51872 | A-Cd 400ppb Hi | 2010-08-23 | --- | --- | --- | 253
51873 | A-Zn-Stock-1050mg/L | 2010-08-23 | --- | --- | 1000000 | ---
51874 | A-Cu-Stock-97.3mg/L | 2010-08-23 | --- | 95900 | --- | ---
51875 | A-Cd-Stock-759mg/L | 2010-08-23 | --- | --- | --- | 768000

## Table 3-3 QC
Table 3-3. Concentrations of elements in filter blanks from sturgeon acute
and chronic exposure studies.
Acute
CERC | Lab | Collection |  | Cu | Zn | Cd | Pb
ID | ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51851 | A-Filter Blank | 2010-08-23 | --- | < 0.084 | < 1.2 | < 0.033 | ---
51217 | A-Filter Blank | 2010-07-16 | 0 | < 0.067 | < 1.5 | < 0.014 | ---
51375 | A-Filter Blank | 2010-07-26 | 0 | < 0.067 | < 1.5 | < 0.014 | ---
51508 | A-Filter Blank | 2010-08-09 | 0 | < 0.061 | < 0.4 | < 0.009 | ---
52044 | A-Filter Blank | 2010-09-09 | 0 | < 0.068 | < 0.6 | < 0.018 | ---
52267 | A-Filter Blank | 2010-09-20 | 0 | < 0.11 | < 0.3 | < 0.027 | ---
52471 | A-Filter Blank | 2010-10-07 | 0 | < 0.050 | < 0.17 | < 0.017 | ---
51298 | A-Filter Blank | 2010-07-16 | 4 | < 0.077 | < 0.8 | < 0.013 | ---
51451 | A-Filter Blank | 2010-07-30 | 4 | < 0.044 | < 0.3 | < 0.018 | ---
51642 | A-Filter Blank | 2010-08-13 | 4 | < 0.034 | < 0.8 | < 0.010 | ---
51950 | A-Filter Blank | 2010-08-27 | 4 | < 0.062 | < 0.9 | < 0.045 | ---
52114 | A-Filter Blank | 2010-09-13 | 4 | < 0.11 | < 0.3 | < 0.027 | ---
52351 | A-Filter Blank | 2010-09-24 | 4 | < 0.054 | < 0.3 | < 0.010 | ---
52496 | A-Filter Blank | 2010-10-11 | 4 | < 0.050 | 0.19 | < 0.017 | ---
Chronic
CERC | Lab | Collection |  | Cu | Zn | Cd | Pb
ID | ID | Date | Day | (ng/mL) | (ng/mL) | (ng/mL) | (ng/mL)
51245 | C-Filter Blank | 2010-07-12 | 0 | < 0.04 | < 0.2 | < 0.024 | < 0.025
51530 | C-Filter Blank | 2010-08-02 | 3 | < 0.061 | < 0.4 | < 0.009 | < 0.031
51335 | C-Filter Blank | 2010-07-19 | 7 | < 0.087 | < 0.4 | < 0.009 | < 0.019
51673 | C-Filter Blank | 2010-08-16 | 10 | < 0.019 | < 0.2 | < 0.017 | < 0.031
51406 | C-Filter Blank | 2010-07-26 | 14 | < 0.087 | < 0.4 | < 0.009 | < 0.019
51876 | C-Filter Blank | 2010-08-23 | 17 | < 0.084 | < 1.2 | < 0.033 | < 0.008
51479 | C-Filter Blank | 2010-08-02 | 21 | < 0.044 | < 0.3 | < 0.018 | < 0.036
51978 | C-Filter Blank | 2010-08-30 | 24 | < 0.068 | < 0.56 | < 0.018 | < 0.013
52011 | C-Filter Blank | 2010-09-02 | 27 | < 0.062 | < 0.9 | < 0.045 | < 0.008

## Table 3-4 QC
Table 3-4. Concentrations of elements in a continuing calibration blank (CCB) and
independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-07-22 | Cu | 0.004259 | 14.9407 | 99.6049 | 2010-07-22 | Cu | -0.001182 | 14.2908 | 95.2722
Run #1 | Zn | -0.005137 | 203.6628 | 101.8314 | Run #9 | Zn | -0.00075 | 202.4886 | 101.2443
| Cd | -0.000326 | 4.0954 | 102.385 |  | Cd | -4.4e-05 | 3.8182 | 95.4546
| Pb | 0.003585 | 15.166 | 101.1066 |  | Pb | -0.000637 | 14.4105 | 96.0698
2010-07-22 | Cu | 0.001055 | 14.739 | 98.2603 | 2010-08-17 | Cu | 0.002922 | 14.4006 | 96.0037
Run #2 | Zn | 0.016938 | 203.7284 | 101.8642 | Run #1 | Zn | 0.001218 | 196.5224 | 98.2612
| Cd | 3.8e-05 | 3.9988 | 99.9691 |  | Cd | 0.000598 | 3.8865 | 97.1632
| Pb | 0.001737 | 15.0607 | 100.4049
[5 empty] | 2010-08-17 | Cu | 0.001267 | 15.219 | 101.4597
2010-07-22 | Cu | 0.000863 | 14.4331 | 96.2205 | Run #2 | Zn | -0.000386 | 206.5908 | 103.2954
Run #3 | Zn | -0.010429 | 199.4687 | 99.7343 |  | Cd | 0.00084 | 3.8545 | 96.3621
| Cd | -0.000864 | 3.984 | 99.6003
| Pb | -0.000615 | 15.0218 | 100.1454 | 2010-08-17 | Cu | 0.000936 | 15.2424 | 101.6161
[5 empty] | Run #3 | Zn | 0.004609 | 206.5694 | 103.2847
2010-07-22 | Cu | 0.001497 | 14.313 | 95.4203 |  | Cd | 0.000483 | 3.8285 | 95.7116
Run #4 | Zn | -0.009965 | 197.9328 | 98.9664
| Cd | -0.000239 | 4.0082 | 100.2039 | 2010-08-17 | Cu | 0.001957 | 14.6618 | 97.7451
| Pb | -0.00112 | 14.9024 | 99.3496 | Run #4 | Zn | 0.009224 | 200.1259 | 100.063
[6 empty] | Cd | 0.000303 | 3.8873 | 97.183
2010-07-22 | Cu | 0.002619 | 14.9122 | 99.4149
Run #5 | Zn | -0.003769 | 205.003 | 102.5015 | 2010-08-17 | Cu | -0.000199 | 14.5662 | 97.1081
| Cd | 0.000225 | 4.039 | 100.9753 | Run #5 | Zn | 0.007006 | 199.5953 | 99.7976
| Pb | 0.002757 | 15.2388 | 101.5918 |  | Cd | 0.000571 | 3.9218 | 98.0451
2010-07-22 | Cu | 0.001353 | 14.4673 | 96.4487 | 2010-08-17 | Cu | 0.000109 | 14.4017 | 96.0114
Run #6 | Zn | -0.011972 | 198.7328 | 99.3664 | Run #6 | Zn | 0.005183 | 198.1244 | 99.0622
| Cd | -0.000594 | 3.9659 | 99.1486 |  | Cd | 0.000854 | 3.9281 | 98.2024
| Pb | 0.001189 | 15.0995 | 100.6631
[5 empty] | 2010-08-17 | Cu | 0.000784 | 14.2934 | 95.2893
2010-07-22 | Cu | 0.001863 | 14.7058 | 98.0387 | Run #7 | Zn | 0.007167 | 196.6096 | 98.3048
Run #7 | Zn | 0.005366 | 203.1659 | 101.583 |  | Cd | 0.000415 | 3.9493 | 98.7332
| Cd | 0.000833 | 4.0369 | 100.9232
| Pb | 0.003658 | 15.1495 | 100.9965
2010-07-22 | Cu | 0.003857 | 14.5708 | 97.1384
Run #8 | Zn | 0.16341 | 202.4168 | 101.2084
| Cd | 0.001032 | 3.9401 | 98.5022
| Pb | 0.00307 | 14.9885 | 99.9235
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4b QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-07-27 | Cu | 0.00042 | 14.7403 | 98.2689 | 2010-08-18 | Cu | 0.002183 | 14.7982 | 98.6544
Run #1 | Zn | 0.00253 | 206.7279 | 103.364 | Run #1 | Zn | 0.006953 | 201.9549 | 100.9775
| Cd | 0.001351 | 3.9794 | 99.4855 |  | Cd | -0.000563 | 4.0075 | 100.1882
| Pb | 0.001273 | 14.934 | 99.56 |  | Pb | 0.002509 | 15.2498 | 101.665
2010-07-27 | Cu | 0.00116 | 14.7289 | 98.1929 | 2010-08-18 | Cu | 0.00381 | 14.6912 | 97.9411
Run #2 | Zn | -0.008216 | 199.7894 | 99.8947 | Run #2 | Zn | 0.004107 | 202.5047 | 101.2523
| Cd | 0.000632 | 3.9657 | 99.1417 |  | Cd | 0.000679 | 4.0348 | 100.8692
| Pb | 0.002128 | 15.1801 | 101.2007 |  | Pb | 0.002088 | 15.1826 | 101.2174
2010-07-27 | Cu | -0.002186 | 14.273 | 95.1536 | 2010-08-18 | Cu | 0.000467 | 14.3913 | 95.9418
Run #3 | Zn | -0.008825 | 195.3967 | 97.6984 | Run #3 | Zn | 0.00346 | 200.0845 | 100.0422
| Cd | 6.2e-05 | 3.9162 | 97.9047 |  | Cd | 0.000108 | 3.9231 | 98.0772
| Pb | -0.000999 | 14.9653 | 99.7685 |  | Pb | -0.000559 | 14.9544 | 99.6963
2010-07-27 | Cu | 0.001395 | 14.741 | 98.273 | 2010-08-18 | Cu | 0.000718 | 14.3247 | 95.498
Run #4 | Zn | 0.001136 | 202.7216 | 101.3608 | Run #4 | Zn | 0.003341 | 196.2344 | 98.1172
| Cd | 0.000533 | 3.9382 | 98.4539 |  | Cd | -4.9e-05 | 3.9028 | 97.5692
| Pb | 0.002367 | 14.8943 | 99.2953 |  | Pb | -0.000396 | 14.7915 | 98.6099
2010-07-27 | Cu | -3.7e-05 | 15.1439 | 100.9593 | 2010-08-18 | Cu | -0.002461 | 14.8292 | 98.8612
Run #5 | Zn | -0.025048 | 206.3442 | 103.1721 | Run #5 | Zn | -0.099559 | 202.7912 | 101.3956
| Cd | -3.2e-05 | 4.1401 | 103.5014 |  | Cd | -0.000488 | 3.983 | 99.574
| Pb | 0.000615 | 15.3368 | 102.2455 |  | Pb | -0.001269 | 15.0057 | 100.0379
2010-07-27 | Cu | -0.001433 | 14.328 | 95.5203 | 2010-08-18 | Cu | -3.5e-05 | 14.7532 | 98.3547
Run #6 | Zn | 0.008735 | 196.6999 | 98.35 | Run #6 | Zn | -0.023077 | 201.5983 | 100.7992
| Cd | 0.000378 | 3.9336 | 98.3393 |  | Cd | 0.000152 | 4.005 | 100.1259
| Pb | -0.001558 | 14.9234 | 99.4891 |  | Pb | -0.000997 | 14.9385 | 99.5903
2010-07-27 | Cu | -0.001029 | 14.2969 | 95.3124 | 2010-08-18 | Cu | -0.000627 | 14.1546 | 94.3643
Run #7 | Zn | 0.004307 | 195.049 | 97.5245 | Run #7 | Zn | -0.015444 | 194.1111 | 97.0555
| Cd | -1e-05 | 3.9323 | 98.3081 |  | Cd | -0.000358 | 3.8528 | 96.3191
| Pb | 0.000349 | 14.7213 | 98.1419 |  | Pb | -0.001323 | 14.4308 | 96.2057
2010-07-27 | Cu | -0.001483 | 14.2508 | 95.0056
Run #8 | Zn | 0.004391 | 196.3732 | 98.1866
| Cd | 3.3e-05 | 3.9482 | 98.7046
| Pb | -0.001527 | 14.6671 | 97.7805
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4c QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-07-29 | Cu | 0.002104 | 14.868 | 99.1202 | 2010-10-05 | Cu | 0.000746 | 14.7626 | 98.4174
Run #1 | Zn | -0.004876 | 202.6239 | 101.312 | Run #1 | Zn | 0.005358 | 202.7479 | 101.3739
| Cd | -0.000238 | 4.0132 | 100.3301 |  | Cd | 0.000359 | 3.9963 | 99.9079
2010-07-29 | Cu | 0.00045 | 14.7306 | 98.204 | 2010-10-05 | Cu | -0.001448 | 14.8899 | 99.2661
Run #2 | Zn | -0.004944 | 201.7547 | 100.8773 | Run #2 | Zn | -0.003016 | 203.5431 | 101.7716
| Cd | 0.000374 | 3.9691 | 99.2282 |  | Cd | -0.000408 | 4.0763 | 101.9068
2010-07-29 | Cu | -0.0003 | 14.5667 | 97.1111 | 2010-10-05 | Cu | -0.000963 | 14.8514 | 99.0096
Run #3 | Zn | 0.004404 | 200.5738 | 100.2869 | Run #3 | Zn | 0.001475 | 201.0857 | 100.5429
| Cd | 9.9e-05 | 3.9854 | 99.6358 |  | Cd | 0.000422 | 4.052 | 101.3008
2010-07-29 | Cu | 0.000565 | 14.5035 | 96.6899 | 2010-10-05 | Cu | -0.00131 | 14.7685 | 98.4567
Run #4 | Zn | -0.005758 | 197.746 | 98.873 | Run #4 | Zn | 0.001899 | 200.0235 | 100.0117
| Cd | -0.000798 | 3.9479 | 98.6966 |  | Cd | 0.000296 | 4.0168 | 100.421
2010-07-29 | Cu | 0.001308 | 14.5705 | 97.1368 | 2010-10-05 | Cu | -0.002019 | 14.4421 | 96.2806
Run #5 | Zn | -0.013722 | 200.6536 | 100.3268 | Run #5 | Zn | 1.9e-05 | 197.4897 | 98.7449
| Cd | 0.001244 | 3.9612 | 99.0289 |  | Cd | 0.000335 | 3.9231 | 98.078
2010-07-29 | Cu | -0.000131 | 14.5071 | 96.7137 | 2010-10-05 | Cu | -0.000534 | 14.3096 | 95.3971
Run #6 | Zn | -0.005414 | 200.6144 | 100.3072 | Run #6 | Zn | -0.010301 | 193.9009 | 96.9504
| Cd | 0.000538 | 3.942 | 98.5502 |  | Cd | 0.00051 | 3.8972 | 97.4305
2010-07-29 | Cu | -0.000221 | 14.3689 | 95.7929
Run #7 | Zn | -0.00294 | 197.2428 | 98.6214
| Cd | 0.001036 | 3.9027 | 97.567
2010-07-29 | Cu | 0.001397 | 14.4953 | 96.6356
Run #8 | Zn | -0.002675 | 199.875 | 99.9375
| Cd | 0.000982 | 3.9293 | 98.2334
2010-07-29 | Cu | 0.001756 | 14.7021 | 98.0143
Run #9 | Zn | -0.005457 | 198.1833 | 99.0916
| Cd | -0.000706 | 3.8721 | 96.802
2010-07-29 | Cu | 0.000634 | 14.3507 | 95.6712
Run #10 | Zn | -0.004437 | 197.4525 | 98.7263
| Cd | -0.000405 | 3.9223 | 98.0587
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4d QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-08-03 | Cu | -0.003899 | 14.6925 | 97.9503 | 2010-08-03 | Cu | 8.6e-05 | 15.2283 | 101.5218
Run #1 | Zn | -0.000614 | 200.6274 | 100.3137 | Run #9 | Zn | -0.013407 | 207.5265 | 103.7632
| Cd | 0.000568 | 3.9811 | 99.5265 |  | Cd | -0.000387 | 3.9491 | 98.7278
| Pb | 0.002784 | 15.0159 | 100.1062 |  | Pb | -0.000381 | 15.0212 | 100.1414
2010-08-03 | Cu | 0.000467 | 14.922 | 99.48 | 2010-08-03 | Cu | 0.000488 | 14.9262 | 99.5083
Run #2 | Zn | -0.021927 | 203.2598 | 101.6299 | Run #10 | Zn | -0.014338 | 205.528 | 102.764
| Cd | 0.0002 | 3.9697 | 99.2428 |  | Cd | -0.000291 | 3.8972 | 97.4309
| Pb | 0.002411 | 15.0456 | 100.3041 |  | Pb | 0.000123 | 15.0614 | 100.4094
2010-08-03 | Cu | -0.005732 | 14.5395 | 96.9302 | 2010-08-03 | Cu | -0.000994 | 14.8493 | 98.9956
Run #3 | Zn | -0.032051 | 199.8968 | 99.9484 | Run #11 | Zn | -0.011846 | 203.6985 | 101.8492
| Cd | -0.000512 | 3.939 | 98.4761 |  | Cd | 0.000168 | 3.9199 | 97.9987
| Pb | -0.000565 | 14.7516 | 98.3437 |  | Pb | -0.000734 | 15.2016 | 101.3442
2010-08-03 | Cu | -0.005364 | 14.4359 | 96.2392
Run #4 | Zn | -0.02619 | 198.4038 | 99.2019
| Cd | -0.00044 | 3.9751 | 99.3765
| Pb | -0.000747 | 14.9291 | 99.5276
2010-08-03 | Cu | -0.010749 | 14.1719 | 94.4792
Run #5 | Zn | -0.040331 | 194.2394 | 97.1197
| Cd | -0.000419 | 3.9242 | 98.1042
| Pb | -0.00027 | 14.8873 | 99.2485
2010-08-03 | Cu | -0.014305 | 14.0007 | 93.3382
Run #6 | Zn | -0.058972 | 190.3111 | 95.1556
| Cd | -0.000596 | 3.9591 | 98.978
| Pb | 9.6e-05 | 14.8781 | 99.1875
2010-08-03 | Cu | -0.016735 | 13.9136 | 92.7573
Run #7 | Zn | -0.05727 | 190.0305 | 95.0153
| Cd | -0.000168 | 3.9711 | 99.2776
| Pb | -0.001344 | 14.5699 | 97.1325
2010-08-03 | Cu | 0.000975 | 14.6979 | 97.9863
Run #8 | Zn | -0.019499 | 199.3175 | 99.6588
| Cd | 0.000884 | 3.9775 | 99.4367
| Pb | 0.002154 | 15.1804 | 101.203
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4e QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-08-09 | Cu | 0.004057 | 14.7772 | 98.5148 | 2010-08-09 | Cu | 0.000999 | 14.3778 | 95.8522
Run #1 | Zn | 0.000233 | 202.5575 | 101.2788 | Run #9 | Zn | -0.000804 | 200.6312 | 100.3156
| Cd | 0.000424 | 3.9549 | 98.8719 |  | Cd | -0.00024 | 3.9199 | 97.998
| Pb | 0.004281 | 15.0739 | 100.493 |  | Pb | 0.002487 | 15.1043 | 100.6951
2010-08-09 | Cu | 0.000885 | 15.3265 | 102.1765 | 2010-08-09 | Cu | -0.000742 | 14.1055 | 94.0366
Run #2 | Zn | 0.005914 | 210.3877 | 105.1939 | Run #10 | Zn | -0.007471 | 196.1188 | 98.0594
| Cd | 0.000461 | 4.0394 | 100.9859 |  | Cd | -4.4e-05 | 3.9214 | 98.0344
| Pb | -0.000363 | 15.2305 | 101.5366 |  | Pb | 0.001176 | 14.7204 | 98.1359
2010-08-09 | Cu | 0.000122 | 14.9272 | 99.5147 | 2010-08-09 | Cu | 0.001098 | 14.2241 | 94.827
Run #3 | Zn | -0.004135 | 205.227 | 102.6135 | Run #11 | Zn | -0.000309 | 197.0204 | 98.5102
| Cd | 0.00031 | 4.0256 | 100.6403 |  | Cd | -0.000246 | 3.9295 | 98.2369
| Pb | -0.000802 | 15.037 | 100.2465 |  | Pb | 0.000998 | 14.7667 | 98.445
2010-08-09 | Cu | 0.001096 | 14.7015 | 98.0102 | 2010-08-09 | Cu | -0.000153 | 13.7537 | 91.6913
Run #4 | Zn | 0.004826 | 202.3987 | 101.1994 | Run #12 | Zn | -0.007063 | 191.0998 | 95.5499
| Cd | 0.000572 | 3.9522 | 98.8039 |  | Cd | -0.000526 | 3.7358 | 93.3948
| Pb | -0.000992 | 14.8278 | 98.852 |  | Pb | 0.000127 | 14.2457 | 94.9712
2010-08-09 | Cu | 0.000181 | 14.0969 | 93.9792
Run #5 | Zn | -0.008211 | 193.6373 | 96.8186
| Cd | 0.000336 | 3.8494 | 96.236
| Pb | -0.000839 | 14.3933 | 95.9553
2010-08-09 | Cu | 0.002277 | 14.5545 | 97.0301
Run #6 | Zn | 0.00371 | 201.8842 | 100.9421
| Cd | 0.000541 | 3.9211 | 98.0274
| Pb | 0.002183 | 15.0545 | 100.3635
2010-08-09 | Cu | 0.00036 | 14.3664 | 95.7758
Run #7 | Zn | -0.001005 | 199.3725 | 99.6862
| Cd | 0.000527 | 3.8251 | 95.6269
| Pb | -0.000712 | 14.6536 | 97.6908
2010-08-09 | Cu | 3.3e-05 | 14.0358 | 93.5717
Run #8 | Zn | -0.004796 | 193.7979 | 96.899
| Cd | 0.000653 | 3.8454 | 96.136
| Pb | -0.000712 | 14.4068 | 96.0455
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4f QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-08-13 | Cu | 0.00246 | 15.4671 | 103.1143 | 2010-08-13 | Cu | 0.000229 | 14.5566 | 97.0443
Run #1 | Zn | -0.001082 | 213.9201 | 106.9601 | Run #9 | Zn | -0.009796 | 198.9074 | 99.4537
| Cd | 0.00041 | 3.8947 | 97.3675 |  | Cd | 0.000652 | 3.9018 | 97.5442
| Pb | 0.002375 | 14.9391 | 99.5939 |  | Pb | -0.001295 | 14.5444 | 96.9625
2010-08-13 | Cu | 0.002477 | 14.8136 | 98.7576 | 2010-08-13 | Cu | -3.6e-05 | 14.4671 | 96.4471
Run #2 | Zn | -0.007126 | 202.2552 | 101.1276 | Run #10 | Zn | 0.003198 | 199.4346 | 99.7173
| Cd | 0.001362 | 3.9517 | 98.7929 |  | Cd | 0.000364 | 3.9109 | 97.773
| Pb | 0.002639 | 15.0798 | 100.5317 |  | Pb | 0.002871 | 14.7773 | 98.5152
2010-08-13 | Cu | -0.000627 | 14.5503 | 97.0017 | 2010-08-13 | Cu | 0.000399 | 14.5038 | 96.6918
Run #3 | Zn | -0.017168 | 204.0979 | 102.0489 | Run #11 | Zn | 0.000538 | 197.6547 | 98.8274
| Cd | 0.000467 | 3.9847 | 99.6163 |  | Cd | 0.000271 | 3.8815 | 97.0372
| Pb | -0.001465 | 15.0816 | 100.544 |  | Pb | 0.001439 | 14.7078 | 98.0518
2010-08-13 | Cu | 0.00063 | 14.6499 | 97.6662
Run #4 | Zn | -0.016874 | 199.2902 | 99.6451
| Cd | 2.2e-05 | 3.9995 | 99.9864
| Pb | -0.000484 | 14.9653 | 99.7687
2010-08-13 | Cu | 0.000965 | 15.0325 | 100.2168
Run #5 | Zn | -0.023135 | 209.0453 | 104.5227
| Cd | 0.000778 | 3.9326 | 98.3143
| Pb | -0.001255 | 14.974 | 99.8265
2010-08-13 | Cu | -0.000356 | 14.3423 | 95.6153
Run #6 | Zn | -0.018074 | 198.4545 | 99.2272
| Cd | 0.000704 | 3.9818 | 99.5448
| Pb | -0.001905 | 14.901 | 99.3399
2010-08-13 | Cu | 0.000186 | 14.5796 | 97.1974
Run #7 | Zn | -0.014956 | 202.6193 | 101.3096
| Cd | 0.00087 | 3.8442 | 96.1046
| Pb | -0.000452 | 14.7974 | 98.6493
2010-08-13 | Cu | -0.000948 | 14.495 | 96.6336
Run #8 | Zn | -0.011826 | 200.7812 | 100.3906
| Cd | 0.000453 | 3.892 | 97.2994
| Pb | -0.001261 | 14.6258 | 97.5056
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4g QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-08-26 | Cu | 0.00077 | 14.7763 | 98.5089 | 2010-08-26 | Cu | -0.000662 | 14.7713 | 98.4751
Run #1 | Zn | 0.001339 | 200.0718 | 100.0359 | Run #9 | Zn | 0.000592 | 205.7543 | 102.8771
| Cd | 0.000939 | 3.9234 | 98.0841 |  | Cd | 0.000607 | 3.8648 | 96.6196
| Pb | 0.001867 | 14.9721 | 99.8139 |  | Pb | 0.000106 | 14.6942 | 97.9614
2010-08-26 | Cu | -0.000201 | 14.664 | 97.7603 | 2010-08-26 | Cu | -0.001232 | 14.7633 | 98.4221
Run #2 | Zn | -0.014109 | 199.9921 | 99.9961 | Run #10 | Zn | 0.003517 | 204.0177 | 102.0089
| Cd | 0.000656 | 4.0205 | 100.5126 |  | Cd | -9.5e-05 | 3.8737 | 96.8438
| Pb | 0.001689 | 14.7159 | 98.1059 |  | Pb | 0.000346 | 14.5156 | 96.7706
2010-08-26 | Cu | 0.002943 | 14.3879 | 95.9194 | 2010-08-26 | Cu | -0.001182 | 14.5456 | 96.9706
Run #3 | Zn | -0.007504 | 201.3885 | 100.6942 | Run #11 | Zn | -0.002461 | 201.6758 | 100.8379
| Cd | -0.000166 | 3.8895 | 97.2382 |  | Cd | 8.6e-05 | 3.8111 | 95.2766
| Pb | 0.00165 | 14.5203 | 96.8019 |  | Pb | -0.000262 | 14.3913 | 95.9423
2010-08-26 | Cu | 0.000637 | 14.4931 | 96.6208
Run #4 | Zn | 0.005435 | 197.1703 | 98.5852
| Cd | 5.5e-05 | 3.9316 | 98.2912
| Pb | 0.001944 | 14.9819 | 99.8796
2010-08-26 | Cu | 0.001612 | 14.6109 | 97.4057
Run #5 | Zn | 0.009968 | 192.7199 | 96.36
| Cd | 0.000447 | 3.9191 | 97.9772
| Pb | 0.001776 | 14.8342 | 98.8946
2010-08-26 | Cu | -0.00012 | 14.5967 | 97.3116
Run #6 | Zn | -0.004018 | 190.0766 | 95.0383
| Cd | -0.000171 | 3.9474 | 98.6846
| Pb | -0.00089 | 14.8852 | 99.2345
2010-08-26 | Cu | 0.000891 | 14.901 | 99.3402
Run #7 | Zn | -0.001614 | 203.2668 | 101.6334
| Cd | 0.000169 | 4.0273 | 100.682
| Pb | 0.003335 | 15.0551 | 100.3676
2010-08-26 | Cu | -0.003016 | 15.1849 | 101.2325
Run #8 | Zn | -0.000333 | 207.1766 | 103.5883
| Cd | 0.000167 | 3.8367 | 95.9186
| Pb | -0.00122 | 14.9492 | 99.6615
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4h QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-09-07 | Cu | 0.001091 | 14.6821 | 97.8805 | 2010-09-07 | Cu | 0.001511 | 14.5428 | 96.9519
Run #1 | Zn | -0.00305 | 200.4696 | 100.2348 | Run #9 | Zn | 0.059086 | 201.2728 | 100.6364
| Cd | -0.000327 | 3.9309 | 98.2725 |  | Cd | -0.000258 | 3.9383 | 98.4572
| Pb | 0.002826 | 15.0354 | 100.2361 |  | Pb | 0.000414 | 14.7452 | 98.3015
2010-09-07 | Cu | -0.001757 | 14.6219 | 97.4795 | 2010-09-07 | Cu | 0.001527 | 14.6265 | 97.51
Run #2 | Zn | -0.002391 | 200.5254 | 100.2627 | Run #10 | Zn | 0.06327 | 200.9195 | 100.4598
| Cd | -0.000692 | 3.998 | 99.9488 |  | Cd | -0.000291 | 3.9415 | 98.5376
| Pb | 0.000374 | 14.981 | 99.8733 |  | Pb | 0.001216 | 14.6884 | 97.9224
2010-09-07 | Cu | -0.002058 | 14.5191 | 96.7939 | 2010-09-07 | Cu | 0.000755 | 14.4464 | 96.3093
Run #3 | Zn | 0.000135 | 198.7777 | 99.3889 | Run #11 | Zn | 0.054247 | 201.7248 | 100.8624
| Cd | -0.000724 | 3.9978 | 99.9453 |  | Cd | 0.000219 | 3.867 | 96.6756
| Pb | 0.00053 | 14.879 | 99.1936 |  | Pb | 0.000571 | 14.7939 | 98.6262
2010-09-07 | Cu | -0.002973 | 14.4781 | 96.5204
Run #4 | Zn | -0.005352 | 197.8136 | 98.9068
| Cd | -1.8e-05 | 3.9816 | 99.5392
| Pb | 0.000629 | 14.7896 | 98.5974
2010-09-07 | Cu | -0.001533 | 15.1016 | 100.6775
Run #5 | Zn | 0.006077 | 206.3345 | 103.1672
| Cd | -0.000289 | 4.0841 | 102.1018
| Pb | 0.000329 | 15.4944 | 103.2957
2010-09-07 | Cu | 0.00285 | 14.7173 | 98.1155
Run #6 | Zn | -0.005993 | 203.0318 | 101.5159
| Cd | 0.000606 | 4.0061 | 100.1538
| Pb | 0.002271 | 14.8782 | 99.1878
2010-09-07 | Cu | -0.0015 | 14.7156 | 98.1038
Run #7 | Zn | -0.006206 | 202.0899 | 101.045
| Cd | 0.000312 | 3.9538 | 98.8457
| Pb | -0.000864 | 14.8998 | 99.3321
2010-09-07 | Cu | 0.001309 | 14.9048 | 99.365
Run #8 | Zn | 0.054211 | 204.9908 | 102.4954
| Cd | 0.000156 | 3.9588 | 98.9694
| Pb | -0.000515 | 14.8661 | 99.1073
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4i QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-09-10 | Cu | 0.002554 | 14.8469 | 98.9795 | 2010-09-10 | Cu | -0.000183 | 14.666 | 97.7735
Run #1 | Zn | -0.00276 | 200.4756 | 100.2378 | Run #9 | Zn | 0.003456 | 196.8348 | 98.4174
| Cd | 0.00021 | 3.9906 | 99.7655 |  | Cd | 0.000151 | 3.9196 | 97.9893
| Pb | 0.002529 | 15.2791 | 101.8609 |  | Pb | -0.000826 | 14.9856 | 99.9039
2010-09-10 | Cu | 0.000522 | 14.7143 | 98.0953 | 2010-09-10 | Cu | 0.00065 | 14.6765 | 97.8435
Run #2 | Zn | -0.007943 | 200.1234 | 100.0617 | Run #10 | Zn | 0.005357 | 195.8481 | 97.924
| Cd | 0.000378 | 4.0012 | 100.0309 |  | Cd | 0.000167 | 3.8488 | 96.2195
| Pb | 0.000539 | 15.2852 | 101.9013 |  | Pb | -0.000364 | 14.8908 | 99.2722
2010-09-10 | Cu | 0.000391 | 15.0379 | 100.2525 | 2010-09-10 | Cu | 0.001339 | 14.5173 | 96.7818
Run #3 | Zn | 0.000505 | 204.009 | 102.0045 | Run #11 | Zn | 0.001708 | 193.5251 | 96.7626
| Cd | -0.000346 | 4.0942 | 102.3541 |  | Cd | -0.000678 | 3.9328 | 98.3208
| Pb | 0.000161 | 15.5669 | 103.7792 |  | Pb | 0.000337 | 14.9836 | 99.8905
2010-09-10 | Cu | -0.001253 | 14.9257 | 99.5044 | 2010-09-10 | Cu | 8.5e-05 | 14.5933 | 97.2885
Run #4 | Zn | 0.003539 | 201.2321 | 100.6161 | Run #12 | Zn | 0.003359 | 197.5301 | 98.7651
| Cd | 2.8e-05 | 4.0768 | 101.9212 |  | Cd | -0.000351 | 3.9508 | 98.7709
| Pb | 0.000537 | 15.3827 | 102.5512 |  | Pb | 0.000127 | 14.8389 | 98.9259
2010-09-10 | Cu | -0.00053 | 14.4108 | 96.0723
Run #5 | Zn | -0.006 | 195.1914 | 97.5957
| Cd | -0.000345 | 3.9687 | 99.2182
| Pb | 0.000192 | 15.0307 | 100.2044
2010-09-10 | Cu | 0.001778 | 14.7774 | 98.5159
Run #6 | Zn | -0.006745 | 196.9037 | 98.4518
| Cd | -0.000188 | 3.9491 | 98.7273
| Pb | 0.00178 | 15.1442 | 100.961
2010-09-10 | Cu | 0.001751 | 14.8007 | 98.6715
Run #7 | Zn | -0.007032 | 197.4738 | 98.7369
| Cd | 0.000572 | 3.9281 | 98.2013
| Pb | 0.000847 | 14.9515 | 99.6763
2010-09-10 | Cu | -0.000972 | 14.8917 | 99.2781
Run #8 | Zn | -0.003301 | 199.4118 | 99.7059
| Cd | 0.000231 | 3.9044 | 97.6098
| Pb | -0.00078 | 14.988 | 99.9202
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4j QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-09-29 | Cu | 0.000311 | 14.8243 | 98.8288 | 2010-09-29 | Cu | -0.001499 | 14.8669 | 99.1129
Run #1 | Zn | 0.001507 | 204.3103 | 102.1551 | Run #9 | Zn | -0.004835 | 202.2308 | 101.1154
| Cd | 0.000272 | 3.8756 | 96.8903 |  | Cd | 0.000747 | 3.9814 | 99.5347
| Pb | 0.002446 | 14.9281 | 99.5208 |  | Pb | -0.000676 | 15.1148 | 100.7651
2010-09-29 | Cu | -0.002126 | 15.1871 | 101.2473 | 2010-09-29 | Cu | -0.000152 | 14.7653 | 98.4353
Run #2 | Zn | 0.007851 | 209.1668 | 104.5834 | Run #10 | Zn | -0.00235 | 200.0124 | 100.0062
| Cd | -0.000953 | 3.9005 | 97.5132 |  | Cd | 0.000257 | 3.9435 | 98.587
| Pb | -0.00093 | 14.5935 | 97.2902 |  | Pb | -0.000363 | 14.9964 | 99.9762
2010-09-29 | Cu | 0.000359 | 14.5062 | 96.7079 | 2010-09-29 | Cu | -0.000919 | 14.6492 | 97.661
Run #3 | Zn | -0.00491 | 199.2719 | 99.6359 | Run #11 | Zn | -0.009807 | 198.377 | 99.1885
| Cd | -1.3e-05 | 3.9601 | 99.0023 |  | Cd | 0.000681 | 3.8743 | 96.8567
| Pb | 0.001551 | 15.041 | 100.2732 |  | Pb | -0.000351 | 14.7921 | 98.6138
2010-09-29 | Cu | -0.004344 | 14.3175 | 95.4498
Run #4 | Zn | -0.022853 | 196.3123 | 98.1562
| Cd | -0.000434 | 3.8198 | 95.4953
| Pb | -0.001366 | 14.6878 | 97.9188
2010-09-29 | Cu | 0.001734 | 14.5123 | 96.7486
Run #5 | Zn | 0.005047 | 199.2734 | 99.6367
| Cd | 0.000556 | 3.9549 | 98.8728
| Pb | 0.001841 | 14.9603 | 99.7352
2010-09-29 | Cu | 0.000195 | 14.3672 | 95.7814
Run #6 | Zn | -0.007208 | 197.3403 | 98.6701
| Cd | -0.000108 | 3.9182 | 97.956
| Pb | -0.001671 | 14.7276 | 98.1842
2010-09-29 | Cu | 0.000174 | 14.4729 | 96.4861
Run #7 | Zn | -0.014039 | 196.9268 | 98.4634
| Cd | 4.3e-05 | 3.8819 | 97.0483
| Pb | -0.001248 | 14.646 | 97.6402
2010-09-29 | Cu | 0.000569 | 14.9549 | 99.6994
Run #8 | Zn | -0.009665 | 202.5055 | 101.2528
| Cd | 0.000417 | 4.0265 | 100.6626
| Pb | 0.002094 | 15.166 | 101.1067
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-4k QC
Table 3-4 (continued). Concentrations of elements in a continuing calibration blank
(CCB) and independent calibration verification standard (ICVS) ran every 10
samples during water analyses from sturgeon tests.
Results expressed as ng/mL.
|  |  |  | % Rec |  |  |  |  | % Rec
BIDa | Element | CCBb | ICVS | (ICVS)c | BIDa | Element | CCBb | ICVS | (ICVS)c
2010-10-14 | Cu | 0.001067 | 14.8172 | 98.7814 | 2010-10-14 | Cu | 0.000473 | 15.008 | 100.0531
Run #1 | Zn | -0.01024 | 202.9141 | 101.457 | Run #9 | Zn | -0.012835 | 202.5147 | 101.2573
| Cd | 3.8e-05 | 4.0321 | 100.8018 |  | Cd | 0.00028 | 3.9052 | 97.631
| Pb | 0.002356 | 15.1661 | 101.1071 |  | Pb | 0.001702 | 14.7475 | 98.3166
2010-10-14 | Cu | 0.000817 | 14.7567 | 98.3782 | 2010-10-14 | Cu | 0.000463 | 14.908 | 99.3865
Run #2 | Zn | -0.021283 | 202.0058 | 101.0029 | Run #10 | Zn | -0.019754 | 200.7845 | 100.3922
| Cd | 0.000178 | 4.0138 | 100.3456 |  | Cd | 0.000516 | 3.883 | 97.074
| Pb | -5.5e-05 | 14.9942 | 99.9615 |  | Pb | 0.000284 | 14.8302 | 98.8677
2010-10-14 | Cu | 0.001885 | 14.7057 | 98.0379 | 2010-10-14 | Cu | 0.001125 | 15.0196 | 100.1306
Run #3 | Zn | -0.020729 | 201.796 | 100.898 | Run #11 | Zn | -0.018661 | 203.0523 | 101.5261
| Cd | 0.00021 | 3.9345 | 98.3634 |  | Cd | 0.000712 | 3.8929 | 97.3213
| Pb | -0.000222 | 14.7758 | 98.5053 |  | Pb | -0.000344 | 14.6973 | 97.9823
2010-10-14 | Cu | -0.000366 | 14.8326 | 98.8838 | 2010-10-14 | Cu | -0.000997 | 14.4839 | 96.5591
Run #4 | Zn | -0.017277 | 200.1364 | 100.0682 | Run #12 | Zn | -0.028852 | 198.3275 | 99.1638
| Cd | -0.000108 | 3.9116 | 97.7908 |  | Cd | 0.000616 | 3.8441 | 96.1036
| Pb | 0.000724 | 14.6158 | 97.4387 |  | Pb | -0.000816 | 14.3747 | 95.831
2010-10-14 | Cu | -0.000343 | 14.6258 | 97.5053
Run #5 | Zn | -0.023393 | 199.6276 | 99.8138
| Cd | 0.000119 | 3.9183 | 97.9584
| Pb | -0.000127 | 14.5633 | 97.0884
2010-10-14 | Cu | -0.000699 | 14.4133 | 96.0884
Run #6 | Zn | -0.028918 | 197.8905 | 98.9452
| Cd | 0.000372 | 3.849 | 96.2252
| Pb | -0.000285 | 14.4994 | 96.6625
2010-10-14 | Cu | -0.000733 | 14.9346 | 99.5638
Run #7 | Zn | -0.01662 | 203.0602 | 101.5301
| Cd | 0.00012 | 3.9654 | 99.1352
| Pb | 2.3e-05 | 14.9014 | 99.3425
2010-10-14 | Cu | 0.000528 | 15.0795 | 100.5302
Run #8 | Zn | -0.015518 | 205.1772 | 102.5886
| Cd | -0.00033 | 3.9614 | 99.034
| Pb | -0.000441 | 14.8637 | 99.091
aBID = Block Initiation Date: a date assigned to each member of a
group of samples that will identify the sample as a member of the group or "block."
bacceptance criteria for CCB is +/- 3 X IDL for each element.
cacceptance criteria for ICVS = +/- 10% (90% - 110%).
ICVS = 15ppb for Ni,Cu; 200ppb for Zn, and 4ppb for Cd.

## Table 3-5 QC
Table 3-5. Recoveries of elements from a reference solution used as a
control sample in the ICP-MS quantitative analysis of diluter water from
sturgeon exposure studies.
| Analysis | Reference |  | Actual | Meas.
BID | Date | Material | Element | Conc. | Conc. | % Reca | ISOP | Oper. Init.
2010-07-22 | 2010-07-22 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.2922 | 95 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | Spex ICS -1c | Zn | 200 +/- 20 | 209.1844 | 100 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 4.1468 | 102 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.6436 | 104 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.7783 | 97 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | Spex ICS -1c | Zn | 200 +/- 20 | 199.0976 | 100 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.915 | 100 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.201 | 100 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.9804 | 98 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | Spex ICS -1c | Zn | 200 +/- 20 | 199.9083 | 100 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.9313 | 100 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.922 | 98 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | Spex ICS -1c | Zn | 200 +/- 20 | 198.2852 | 100 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8304 | 98 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.0606 | 100 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 22.2092 | 99 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | Spex ICS -1c | Zn | 200 +/- 20 | 204.8299 | 100 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.9717 | 100 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.2968 | 101 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.0823 | 94 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | Spex ICS -1c | Zn | 200 +/- 20 | 206.3901 | 100 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 4.2977 | 106 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 13.0603 | 108 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.879 | 98 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | Spex ICS -1c | Zn | 200 +/- 20 | 203.5427 | 100 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8059 | 97 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 11.7975 | 98 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.4558 | 96 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | Spex ICS -1c | Zn | 200 +/- 20 | 200.5398 | 100 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.9517 | 100 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.3095 | 101 | P.241 | MJW/TWM
a%Rec = 100% if within range, otherwise calculated based on upper or lower limit of range.
aSpex ICS-1 = SPEX ClaritasPPT Instrument Check Standard 1; Cat No. CL-ICS-1; Spec Certiprep, Metuchen, NJ;
Solution used as laboratory control sample.
bNIST 1643e = National Institute of Standards and Technology Standard Reference Material Trace Elements in Water
1643e. Concentration results expressed as ng/mL. Solution used as laboratory control sample.

## Table 3-5b QC
Table 3-5 (continued). Recoveries of elements from a reference solution used as a
control sample in the ICP-MS quantitative analysis of diluter water from
sturgeon exposure studies.
| Analysis | Reference |  | Actual | Meas.
BID | Date | Material | Element | Conc. | Conc. | % Reca | ISOP | Oper. Init.
2010-08-26 | 2010-08-24 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.4697 | 96 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | Spex ICS -1c | Zn | 200 +/- 20 | 203.9832 | 100 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8328 | 98 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 11.9808 | 100 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 22.1286 | 98 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | Spex ICS -1c | Zn | 200 +/- 20 | 208.278 | 100 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 4.0373 | 100 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.6071 | 104 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.8874 | 98 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | Spex ICS -1c | Zn | 200 +/- 20 | 211.2102 | 100 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 4.0112 | 100 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.7401 | 105 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.246 | 94 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | Spex ICS -1c | Zn | 200 +/- 20 | 207.6303 | 100 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8463 | 98 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 11.8924 | 99 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.7282 | 100 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | Spex ICS -1c | Zn | 200 +/- 20 | 199.5296 | 100 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8784 | 99 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.035 | 100 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | NIST 1643eb | Cu | 22.76 +/- 0.31 | 21.8072 | 97 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | Spex ICS -1c | Zn | 200 +/- 20 | 201.3837 | 100 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | NIST 1640ad | Cd | 3.992 +/- 0.074 | 3.8779 | 99 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | NIST 1640ad | Pb | 12.101 +/- 0.050 | 12.0216 | 100 | P.241 | MJW/TWM
a%Rec = 100% if within range, otherwise calculated based on upper or lower limit of range.
aSpex ICS-1 = SPEX ClaritasPPT Instrument Check Standard 1; Cat No. CL-ICS-1; Spec Certiprep, Metuchen, NJ;
Solution used as laboratory control sample.
bNIST 1643e = National Institute of Standards and Technology Standard Reference Material Trace Elements in Water
1643e. Concentration results expressed as ng/mL. Solution used as laboratory control sample.

## Table 3-6 QC
Table 3-6. Recoveries of elements from a low level calibration check
in the ICP-MS quantitative analysis of diluter water from sturgeon
exposure studies.
Mean
| Analysis |  | Actual | Meas. Conc.
BID | Date | Element | Conc. (ppb) | (n=3) | SD | % Reca | ISOP | Oper. Init.
2010-07-22 | 2010-07-22 | Cu | 0.5 | 0.4936 | 0.0016221 | 98 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | Zn | 7.5 | 7.5737 | 0.047867 | 101 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | Cd | 0.15 | 0.15127 | 0.0053859 | 100 | P.241 | MJW/TWM
2010-07-22 | 2010-07-22 | Pb | 0.5 | 0.50112 | 0.0067503 | 100 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | Cu | 0.5 | 0.5452 | 0.0086959 | 110 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | Zn | 7.5 | 8.5298 | 0.077524 | 113 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | Cd | 0.15 | 0.1463 | 0.001951 | 100 | P.241 | MJW/TWM
2010-07-27 | 2010-07-27 | Pb | 0.5 | 0.49684 | 0.0060023 | 100 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | Cu | 0.5 | 0.5312 | 0.008608 | 106 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | Zn | 7.5 | 8.25 | 0.036295 | 109 | P.241 | MJW/TWM
2010-07-29 | 2010-07-29 | Cd | 0.15 | 0.15088 | 0.0028553 | 100 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | Cu | 0.5 | 0.54735 | 0.012782 | 110 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | Zn | 7.5 | 8.6772 | 0.12498 | 116 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | Cd | 0.15 | 0.14662 | 0.0028329 | 100 | P.241 | MJW/TWM
2010-08-03 | 2010-08-03 | Pb | 0.5 | 0.50311 | 0.0013776 | 100 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | Cu | 0.5 | 0.53117 | 0.0089585 | 106 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | Zn | 7.5 | 7.9878 | 0.028898 | 107 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | Cd | 0.15 | 0.1554 | 0.0052798 | 107 | P.241 | MJW/TWM
2010-08-09 | 2010-08-09 | Pb | 0.5 | 0.50991 | 0.0046224 | 102 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | Cu | 0.5 | 0.48736 | 0.0077129 | 98 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | Zn | 7.5 | 7.5332 | 0.054303 | 100 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | Cd | 0.15 | 0.14745 | 0.0022193 | 100 | P.241 | MJW/TWM
2010-08-13 | 2010-08-13 | Pb | 0.5 | 0.49478 | 0.0044003 | 98 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | Cu | 0.5 | 0.52454 | 0.006938 | 104 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | Zn | 7.5 | 8.451 | 0.043462 | 113 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | Cd | 0.15 | 0.16777 | 0.0016456 | 113 | P.241 | MJW/TWM
2010-08-17 | 2010-08-17 | Pb | 0.5 | 0.49818 | 0.0039384 | 100 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | Cu | 0.5 | 0.53155 | 0.0026728 | 106 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | Zn | 7.5 | 8.1861 | 0.0028365 | 109 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | Cd | 0.15 | 0.14656 | 0.0054845 | 100 | P.241 | MJW/TWM
2010-08-18 | 2010-08-18 | Pb | 0.5 | 0.5 | 0.008693 | 100 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | Cu | 0.5 | 0.51441 | 0.0077484 | 102 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | Zn | 7.5 | 7.9381 | 0.045954 | 105 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | Cd | 0.15 | 0.14496 | 0.0040618 | 93 | P.241 | MJW/TWM
2010-08-26 | 2010-08-24 | Pb | 0.5 | 0.48311 | 0.0017881 | 96 | P.241 | MJW/TWM
a%Rec = 100% if within range, otherwise calculated based on upper or lower limit of range.

## Table 3-6b QC
Table 3-6 (continued). Recoveries of elements from a low level calibration check
in the ICP-MS quantitative analysis of diluter water from sturgeon
exposure studies.
Mean
| Analysis |  | Actual | Meas. Conc.
BID | Date | Element | Conc. (ppb) | (n=3) | SD | % Reca | ISOP | Oper. Init.
2010-09-07 | 2010-09-07 | Cu | 0.5 | 0.49965 | 0.0022247 | 100 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | Zn | 7.5 | 7.8343 | 0.0055325 | 103 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | Cd | 0.15 | 0.1504 | 0.0021112 | 100 | P.241 | MJW/TWM
2010-09-07 | 2010-09-07 | Pb | 0.5 | 0.5042 | 0.0021543 | 100 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | Cu | 0.5 | 0.4997 | 0.0016982 | 100 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | Zn | 7.5 | 7.7011 | 0.051655 | 103 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | Cd | 0.15 | 0.15439 | 5.0239e-05 | 100 | P.241 | MJW/TWM
2010-09-10 | 2010-09-10 | Pb | 0.5 | 0.5032 | 0.004171 | 100 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | Cu | 0.5 | 0.51201 | 0.0059311 | 102 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | Zn | 7.5 | 8.3171 | 0.055771 | 111 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | Cd | 0.15 | 0.14717 | 0.0019508 | 100 | P.241 | MJW/TWM
2010-09-29 | 2010-09-29 | Pb | 0.5 | 0.48968 | 0.0050839 | 98 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | Cu | 0.5 | 0.50803 | 0.0080675 | 102 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | Zn | 7.5 | 7.9438 | 0.040498 | 105 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | Cd | 0.15 | 0.1565 | 0.00019653 | 107 | P.241 | MJW/TWM
2010-10-05 | 2010-10-06 | Pb | 0.5 | 0.50816 | 0.0055942 | 102 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | Cu | 0.5 | 0.5089 | 0.002272 | 102 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | Zn | 7.5 | 7.6082 | 0.040567 | 101 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | Cd | 0.15 | 0.1463 | 0.001544 | 100 | P.241 | MJW/TWM
2010-10-14 | 2010-10-15 | Pb | 0.5 | 0.49162 | 0.002073 | 98 | P.241 | MJW/TWM
a%Rec = 100% if within range, otherwise calculated based on upper or lower limit of range.

## Table 3-7 QC
Table 3-7. Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-07-22 | 51247 - Analytical | water | Cu | 19.2803 | 19.3231 | 0.042804 | 19.3017 | 0.22176 | P.241 | MJW/TWM
2010-07-22 | 51247 - Analytical | water | Zn | 152.4896 | 152.3295 | 0.16018 | 152.4095 | 0.1051 | P.241 | MJW/TWM
2010-07-22 | 51247 - Analytical | water | Cd | 5.9709 | 5.9507 | 0.020202 | 5.9608 | 0.33891 | P.241 | MJW/TWM
2010-07-22 | 51247 - Analytical | water | Pb | 19.8047 | 19.7144 | 0.090345 | 19.7596 | 0.45722 | P.241 | MJW/TWM
2010-07-22 | 51260 - Analytical | water | Cu | 19.3975 | 19.5965 | 0.19905 | 19.497 | 1.0209 | P.241 | MJW/TWM
2010-07-22 | 51260 - Analytical | water | Zn | 153.1619 | 153.5675 | 0.40553 | 153.3647 | 0.26442 | P.241 | MJW/TWM
2010-07-22 | 51260 - Analytical | water | Cd | 6.1117 | 6.1031 | 0.008615 | 6.1074 | 0.14106 | P.241 | MJW/TWM
2010-07-22 | 51260 - Analytical | water | Pb | 19.8539 | 19.8627 | 0.008809 | 19.8583 | 0.044359 | P.241 | MJW/TWM
2010-07-22 | 51265 - Analytical | water | Cu | 19.4094 | 19.1511 | 0.25831 | 19.2802 | 1.3398 | P.241 | MJW/TWM
2010-07-22 | 51265 - Analytical | water | Zn | 153.493 | 152.0178 | 1.4751 | 152.7554 | 0.96569 | P.241 | MJW/TWM
2010-07-22 | 51265 - Analytical | water | Cd | 5.9578 | 5.9536 | 0.00418 | 5.9557 | 0.070185 | P.241 | MJW/TWM
2010-07-22 | 51265 - Analytical | water | Pb | 19.7576 | 19.9445 | 0.18689 | 19.8511 | 0.94145 | P.241 | MJW/TWM
2010-07-27 | 51301 - Analytical | water | Cu | 19.4968 | 19.3242 | 0.17261 | 19.4105 | 0.88926 | P.241 | MJW/TWM
2010-07-27 | 51301 - Analytical | water | Zn | 152.2272 | 149.8284 | 2.3989 | 151.0278 | 1.5884 | P.241 | MJW/TWM
2010-07-27 | 51301 - Analytical | water | Cd | 5.9196 | 5.8754 | 0.044275 | 5.8975 | 0.75074 | P.241 | MJW/TWM
2010-07-27 | 51301 - Analytical | water | Pb | 19.7716 | 19.5039 | 0.26768 | 19.6378 | 1.3631 | P.241 | MJW/TWM
2010-07-27 | 51310 - Analytical | water | Cu | 19.2464 | 19.3233 | 0.076864 | 19.2849 | 0.39857 | P.241 | MJW/TWM
2010-07-27 | 51310 - Analytical | water | Zn | 154.4392 | 154.8729 | 0.43372 | 154.6561 | 0.28044 | P.241 | MJW/TWM
2010-07-27 | 51310 - Analytical | water | Cd | 5.9541 | 5.9549 | 0.000812 | 5.9545 | 0.013637 | P.241 | MJW/TWM
2010-07-27 | 51310 - Analytical | water | Pb | 19.6449 | 19.2324 | 0.41252 | 19.4386 | 2.1222 | P.241 | MJW/TWM
2010-07-27 | 51319 - Analytical | water | Cu | 18.593 | 18.8586 | 0.2656 | 18.7258 | 1.4184 | P.241 | MJW/TWM
2010-07-27 | 51319 - Analytical | water | Zn | 149.5908 | 150.0074 | 0.41658 | 149.7991 | 0.27809 | P.241 | MJW/TWM
2010-07-27 | 51319 - Analytical | water | Cd | 6.1704 | 6.1954 | 0.025009 | 6.1829 | 0.40449 | P.241 | MJW/TWM
2010-07-27 | 51319 - Analytical | water | Pb | 19.3169 | 19.4379 | 0.12102 | 19.3774 | 0.62453 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7b QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-07-29 | 51219 - Analytical | water | Cu | 19.5444 | 19.4588 | 0.085656 | 19.5016 | 0.43923 | P.241 | MJW/TWM
2010-07-29 | 51219 - Analytical | water | Zn | 154.3996 | 152.5044 | 1.8952 | 153.452 | 1.2351 | P.241 | MJW/TWM
2010-07-29 | 51219 - Analytical | water | Cd | 5.9207 | 5.9398 | 0.019159 | 5.9302 | 0.32307 | P.241 | MJW/TWM
2010-07-29 | 51226 - Analytical | water | Cu | 19.4998 | 19.2584 | 0.24136 | 19.3791 | 1.2455 | P.241 | MJW/TWM
2010-07-29 | 51226 - Analytical | water | Zn | 155.3671 | 154.1397 | 1.2274 | 154.7534 | 0.79316 | P.241 | MJW/TWM
2010-07-29 | 51226 - Analytical | water | Cd | 5.8846 | 5.9544 | 0.069829 | 5.9195 | 1.1796 | P.241 | MJW/TWM
2010-07-29 | 51387 - Analytical | water | Cu | 18.7677 | 18.7374 | 0.030255 | 18.7526 | 0.16134 | P.241 | MJW/TWM
2010-07-29 | 51387 - Analytical | water | Zn | 155.9621 | 155.6856 | 0.27653 | 155.8239 | 0.17746 | P.241 | MJW/TWM
2010-07-29 | 51387 - Analytical | water | Cd | 5.7396 | 5.7376 | 0.002041 | 5.7386 | 0.035566 | P.241 | MJW/TWM
2010-07-29 | 51396 - Analytical | water | Cu | 19.1996 | 19.2961 | 0.096476 | 19.2478 | 0.50123 | P.241 | MJW/TWM
2010-07-29 | 51396 - Analytical | water | Zn | 151.9924 | 152.9863 | 0.99389 | 152.4893 | 0.65178 | P.241 | MJW/TWM
2010-07-29 | 51396 - Analytical | water | Cd | 6.4207 | 6.5155 | 0.094886 | 6.4681 | 1.467 | P.241 | MJW/TWM
2010-08-03 | 51337 - Analytical | water | Cu | 19.6591 | 19.5894 | 0.069744 | 19.6243 | 0.3554 | P.241 | MJW/TWM
2010-08-03 | 51337 - Analytical | water | Zn | 153.7833 | 153.1945 | 0.58883 | 153.4889 | 0.38363 | P.241 | MJW/TWM
2010-08-03 | 51337 - Analytical | water | Cd | 5.8972 | 5.9293 | 0.032121 | 5.9132 | 0.54321 | P.241 | MJW/TWM
2010-08-03 | 51337 - Analytical | water | Pb | 19.7896 | 19.7785 | 0.011082 | 19.784 | 0.056015 | P.241 | MJW/TWM
2010-08-03 | 51343 - Analytical | water | Cu | 18.8363 | 18.7734 | 0.062952 | 18.8048 | 0.33476 | P.241 | MJW/TWM
2010-08-03 | 51343 - Analytical | water | Zn | 150.0721 | 148.3354 | 1.7366 | 149.2038 | 1.1639 | P.241 | MJW/TWM
2010-08-03 | 51343 - Analytical | water | Cd | 5.9004 | 5.9183 | 0.017879 | 5.9093 | 0.30256 | P.241 | MJW/TWM
2010-08-03 | 51343 - Analytical | water | Pb | 19.189 | 19.2375 | 0.048452 | 19.2132 | 0.25218 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7c QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-08-03 | 51420 - Analytical | water | Cu | 19.7692 | 19.7428 | 0.02637 | 19.756 | 0.13348 | P.241 | MJW/TWM
2010-08-03 | 51420 - Analytical | water | Zn | 153.8051 | 154.7332 | 0.92804 | 154.2691 | 0.60157 | P.241 | MJW/TWM
2010-08-03 | 51420 - Analytical | water | Cd | 5.998 | 6.0643 | 0.066306 | 6.0312 | 1.0994 | P.241 | MJW/TWM
2010-08-03 | 51420 - Analytical | water | Pb | 19.7537 | 19.8261 | 0.072429 | 19.7899 | 0.36599 | P.241 | MJW/TWM
2010-08-03 | 51426 - Analytical | water | Cu | 20.0554 | 20.1165 | 0.061172 | 20.086 | 0.30455 | P.241 | MJW/TWM
2010-08-03 | 51426 - Analytical | water | Zn | 157.6012 | 158.9295 | 1.3284 | 158.2654 | 0.83933 | P.241 | MJW/TWM
2010-08-03 | 51426 - Analytical | water | Cd | 5.8269 | 5.8393 | 0.012344 | 5.8331 | 0.21162 | P.241 | MJW/TWM
2010-08-03 | 51426 - Analytical | water | Pb | 20.1051 | 20.0024 | 0.10265 | 20.0537 | 0.51186 | P.241 | MJW/TWM
2010-08-09 | 51454 - Analytical | water | Cu | 20.4852 | 20.5112 | 0.026004 | 20.4982 | 0.12686 | P.241 | MJW/TWM
2010-08-09 | 51454 - Analytical | water | Zn | 157.8966 | 158.784 | 0.88735 | 158.3403 | 0.5604 | P.241 | MJW/TWM
2010-08-09 | 51454 - Analytical | water | Cd | 6.0822 | 6.0987 | 0.016499 | 6.0905 | 0.2709 | P.241 | MJW/TWM
2010-08-09 | 51454 - Analytical | water | Pb | 19.9757 | 20.1467 | 0.171 | 20.0612 | 0.8524 | P.241 | MJW/TWM
2010-08-09 | 51463 - Analytical | water | Cu | 18.3753 | 18.3726 | 0.002662 | 18.374 | 0.014488 | P.241 | MJW/TWM
2010-08-09 | 51463 - Analytical | water | Zn | 152.5304 | 151.9807 | 0.54972 | 152.2556 | 0.36105 | P.241 | MJW/TWM
2010-08-09 | 51463 - Analytical | water | Cd | 5.6137 | 5.6573 | 0.04363 | 5.6355 | 0.7742 | P.241 | MJW/TWM
2010-08-09 | 51463 - Analytical | water | Pb | 18.3656 | 18.49 | 0.12439 | 18.4278 | 0.67503 | P.241 | MJW/TWM
2010-08-09 | 51493 - Analytical | water | Cu | 19.0527 | 18.8693 | 0.18338 | 18.961 | 0.96716 | P.241 | MJW/TWM
2010-08-09 | 51493 - Analytical | water | Zn | 153.212 | 151.6814 | 1.5305 | 152.4467 | 1.004 | P.241 | MJW/TWM
2010-08-09 | 51493 - Analytical | water | Cd | 5.9497 | 5.8975 | 0.052179 | 5.9236 | 0.88087 | P.241 | MJW/TWM
2010-08-09 | 51493 - Analytical | water | Pb | 19.6179 | 19.5542 | 0.063765 | 19.5861 | 0.32556 | P.241 | MJW/TWM
2010-08-09 | 51499 - Analytical | water | Cu | 18.6869 | 18.6808 | 0.006018 | 18.6839 | 0.03221 | P.241 | MJW/TWM
2010-08-09 | 51499 - Analytical | water | Zn | 150.3767 | 150.3983 | 0.0216 | 150.3875 | 0.014363 | P.241 | MJW/TWM
2010-08-09 | 51499 - Analytical | water | Cd | 5.7879 | 5.8198 | 0.031912 | 5.8039 | 0.54984 | P.241 | MJW/TWM
2010-08-09 | 51499 - Analytical | water | Pb | 19.4561 | 19.6407 | 0.18463 | 19.5484 | 0.94446 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7d QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-08-13 | 51510 - Analytical | water | Cu | 19.9917 | 19.8778 | 0.11383 | 19.9348 | 0.57103 | P.241 | MJW/TWM
2010-08-13 | 51510 - Analytical | water | Zn | 154.501 | 154.0163 | 0.48466 | 154.2587 | 0.31419 | P.241 | MJW/TWM
2010-08-13 | 51510 - Analytical | water | Cd | 6.008 | 6.0536 | 0.045604 | 6.0308 | 0.75618 | P.241 | MJW/TWM
2010-08-13 | 51510 - Analytical | water | Pb | 19.6354 | 19.6788 | 0.043344 | 19.6571 | 0.2205 | P.241 | MJW/TWM
2010-08-13 | 51517 - Analytical | water | Cu | 19.6052 | 19.4652 | 0.14 | 19.5352 | 0.71668 | P.241 | MJW/TWM
2010-08-13 | 51517 - Analytical | water | Zn | 166.1035 | 164.8909 | 1.2127 | 165.4972 | 0.73274 | P.241 | MJW/TWM
2010-08-13 | 51517 - Analytical | water | Cd | 5.8173 | 5.7955 | 0.021809 | 5.8064 | 0.3756 | P.241 | MJW/TWM
2010-08-13 | 51517 - Analytical | water | Pb | 19.324 | 19.4282 | 0.10417 | 19.3761 | 0.53762 | P.241 | MJW/TWM
2010-08-13 | 51544 - Analytical | water | Cu | 19.0151 | 19.1672 | 0.15208 | 19.0911 | 0.79661 | P.241 | MJW/TWM
2010-08-13 | 51544 - Analytical | water | Zn | 149.2033 | 150.7303 | 1.527 | 149.9668 | 1.0182 | P.241 | MJW/TWM
2010-08-13 | 51544 - Analytical | water | Cd | 6.5285 | 6.2262 | 0.30225 | 6.3774 | 4.7395 | P.241 | MJW/TWM
2010-08-13 | 51544 - Analytical | water | Pb | 19.0879 | 19.1509 | 0.063037 | 19.1194 | 0.3297 | P.241 | MJW/TWM
2010-08-13 | 51552 - Analytical | water | Cu | 19.0103 | 19.0586 | 0.048344 | 19.0345 | 0.25398 | P.241 | MJW/TWM
2010-08-13 | 51552 - Analytical | water | Zn | 148.8117 | 149.629 | 0.81727 | 149.2203 | 0.54769 | P.241 | MJW/TWM
2010-08-13 | 51552 - Analytical | water | Cd | 5.7289 | 5.7358 | 0.006863 | 5.7324 | 0.11972 | P.241 | MJW/TWM
2010-08-13 | 51552 - Analytical | water | Pb | 20.1667 | 20.1813 | 0.014599 | 20.174 | 0.072365 | P.241 | MJW/TWM
2010-08-17 | 51645 - Analytical | water | Cu | 19.8304 | 19.7738 | 0.05663 | 19.8021 | 0.28598 | P.241 | MJW/TWM
2010-08-17 | 51645 - Analytical | water | Zn | 154.178 | 153.8643 | 0.31377 | 154.0212 | 0.20372 | P.241 | MJW/TWM
2010-08-17 | 51645 - Analytical | water | Cd | 5.9527 | 5.9536 | 0.00088 | 5.9531 | 0.014782 | P.241 | MJW/TWM
2010-08-17 | 51645 - Analytical | water | Pb | 19.718 | 19.6952 | 0.022794 | 19.7066 | 0.11567 | P.241 | MJW/TWM
2010-08-17 | 51654 - Analytical | water | Cu | 19.1483 | 19.2443 | 0.095919 | 19.1963 | 0.49967 | P.241 | MJW/TWM
2010-08-17 | 51654 - Analytical | water | Zn | 162.7306 | 163.4128 | 0.68213 | 163.0717 | 0.4183 | P.241 | MJW/TWM
2010-08-17 | 51654 - Analytical | water | Cd | 5.8229 | 5.8529 | 0.030006 | 5.8379 | 0.51399 | P.241 | MJW/TWM
2010-08-17 | 51654 - Analytical | water | Pb | 19.5173 | 19.519 | 0.00166 | 19.5181 | 0.0085049 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7e QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-08-18 | 51675 - Analytical | water | Cu | 19.3671 | 19.5518 | 0.18467 | 19.4595 | 0.949 | P.241 | MJW/TWM
2010-08-18 | 51675 - Analytical | water | Zn | 151.9034 | 153.8435 | 1.9401 | 152.8734 | 1.2691 | P.241 | MJW/TWM
2010-08-18 | 51675 - Analytical | water | Cd | 5.9381 | 5.9966 | 0.058553 | 5.9673 | 0.98122 | P.241 | MJW/TWM
2010-08-18 | 51675 - Analytical | water | Pb | 19.6161 | 19.907 | 0.29082 | 19.7616 | 1.4717 | P.241 | MJW/TWM
2010-08-18 | 51693 - Analytical | water | Cu | 19.5027 | 19.5204 | 0.017736 | 19.5115 | 0.0909 | P.241 | MJW/TWM
2010-08-18 | 51693 - Analytical | water | Zn | 153.2195 | 152.1751 | 1.0444 | 152.6973 | 0.68399 | P.241 | MJW/TWM
2010-08-18 | 51693 - Analytical | water | Cd | 5.8406 | 5.9049 | 0.06429 | 5.8728 | 1.0947 | P.241 | MJW/TWM
2010-08-18 | 51693 - Analytical | water | Pb | 19.6425 | 19.7276 | 0.085099 | 19.6851 | 0.4323 | P.241 | MJW/TWM
2010-08-26 | 51853 - Analytical | water | Cu | 19.3638 | 19.2201 | 0.14371 | 19.2919 | 0.74493 | P.241 | MJW/TWM
2010-08-26 | 51853 - Analytical | water | Zn | 154.656 | 153.2626 | 1.3934 | 153.9593 | 0.90503 | P.241 | MJW/TWM
2010-08-26 | 51853 - Analytical | water | Cd | 5.83 | 5.8195 | 0.010538 | 5.8247 | 0.18092 | P.241 | MJW/TWM
2010-08-26 | 51853 - Analytical | water | Pb | 19.3172 | 19.1607 | 0.15653 | 19.2389 | 0.8136 | P.241 | MJW/TWM
2010-08-26 | 51860 - Analytical | water | Cu | 19.3993 | 19.438 | 0.038654 | 19.4186 | 0.19906 | P.241 | MJW/TWM
2010-08-26 | 51860 - Analytical | water | Zn | 151.3364 | 150.8994 | 0.43693 | 151.1179 | 0.28913 | P.241 | MJW/TWM
2010-08-26 | 51860 - Analytical | water | Cd | 5.9286 | 5.9282 | 0.000366 | 5.9284 | 0.0061737 | P.241 | MJW/TWM
2010-08-26 | 51860 - Analytical | water | Pb | 19.6991 | 19.6218 | 0.077315 | 19.6605 | 0.39325 | P.241 | MJW/TWM
2010-08-26 | 51890 - Analytical | water | Cu | 19.8469 | 19.758 | 0.088915 | 19.8024 | 0.44901 | P.241 | MJW/TWM
2010-08-26 | 51890 - Analytical | water | Zn | 157.0969 | 157.1055 | 0.008566 | 157.1012 | 0.0054525 | P.241 | MJW/TWM
2010-08-26 | 51890 - Analytical | water | Cd | 5.8791 | 5.9143 | 0.035206 | 5.8967 | 0.59705 | P.241 | MJW/TWM
2010-08-26 | 51890 - Analytical | water | Pb | 19.4532 | 19.3641 | 0.089139 | 19.4087 | 0.45927 | P.241 | MJW/TWM
2010-08-26 | 51897 - Analytical | water | Cu | 20.269 | 19.7651 | 0.50397 | 20.0171 | 2.5177 | P.241 | MJW/TWM
2010-08-26 | 51897 - Analytical | water | Zn | 161.174 | 157.0425 | 4.1315 | 159.1082 | 2.5967 | P.241 | MJW/TWM
2010-08-26 | 51897 - Analytical | water | Cd | 5.9202 | 5.7718 | 0.14836 | 5.846 | 2.5378 | P.241 | MJW/TWM
2010-08-26 | 51897 - Analytical | water | Pb | 20.2634 | 19.8025 | 0.46088 | 20.033 | 2.3006 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7f QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-09-07 | 51953 - Analytical | water | Cu | 19.6024 | 19.607 | 0.004529 | 19.6047 | 0.023102 | P.241 | MJW/TWM
2010-09-07 | 51953 - Analytical | water | Zn | 151.9599 | 150.4736 | 1.4863 | 151.2168 | 0.98287 | P.241 | MJW/TWM
2010-09-07 | 51953 - Analytical | water | Cd | 6.0027 | 5.9174 | 0.085283 | 5.9601 | 1.4309 | P.241 | MJW/TWM
2010-09-07 | 51953 - Analytical | water | Pb | 19.6977 | 19.6715 | 0.026197 | 19.6846 | 0.13308 | P.241 | MJW/TWM
2010-09-07 | 51962 - Analytical | water | Cu | 19.1885 | 19.137 | 0.051507 | 19.1627 | 0.26879 | P.241 | MJW/TWM
2010-09-07 | 51962 - Analytical | water | Zn | 162.5504 | 162.6948 | 0.14441 | 162.6226 | 0.088802 | P.241 | MJW/TWM
2010-09-07 | 51962 - Analytical | water | Cd | 5.8207 | 5.8155 | 0.005194 | 5.8181 | 0.089273 | P.241 | MJW/TWM
2010-09-07 | 51962 - Analytical | water | Pb | 19.2105 | 19.1415 | 0.068979 | 19.176 | 0.35971 | P.241 | MJW/TWM
2010-09-07 | 52025 - Analytical | water | Cu | 19.4791 | 19.4504 | 0.028687 | 19.4647 | 0.14738 | P.241 | MJW/TWM
2010-09-07 | 52025 - Analytical | water | Zn | 153.4144 | 153.5922 | 0.17784 | 153.5033 | 0.11586 | P.241 | MJW/TWM
2010-09-07 | 52025 - Analytical | water | Cd | 5.9488 | 5.9119 | 0.03695 | 5.9304 | 0.62307 | P.241 | MJW/TWM
2010-09-07 | 52025 - Analytical | water | Pb | 19.3962 | 19.3162 | 0.079948 | 19.3562 | 0.41304 | P.241 | MJW/TWM
2010-09-07 | 52031 - Analytical | water | Cu | 19.6538 | 19.3055 | 0.34834 | 19.4797 | 1.7882 | P.241 | MJW/TWM
2010-09-07 | 52031 - Analytical | water | Zn | 153.7277 | 152.9868 | 0.7409 | 153.3572 | 0.48312 | P.241 | MJW/TWM
2010-09-07 | 52031 - Analytical | water | Cd | 5.9569 | 5.858 | 0.098819 | 5.9075 | 1.6728 | P.241 | MJW/TWM
2010-09-07 | 52031 - Analytical | water | Pb | 19.6223 | 19.7208 | 0.098493 | 19.6715 | 0.50069 | P.241 | MJW/TWM
2010-09-10 | 51992 - Analytical | water | Cu | 19.425 | 19.6494 | 0.22434 | 19.5372 | 1.1483 | P.241 | MJW/TWM
2010-09-10 | 51992 - Analytical | water | Zn | 149.383 | 151.5725 | 2.1895 | 150.4777 | 1.455 | P.241 | MJW/TWM
2010-09-10 | 51992 - Analytical | water | Cd | 5.8227 | 5.9165 | 0.093731 | 5.8696 | 1.5969 | P.241 | MJW/TWM
2010-09-10 | 51992 - Analytical | water | Pb | 19.525 | 19.3519 | 0.17312 | 19.4385 | 0.89058 | P.241 | MJW/TWM
2010-09-10 | 51998 - Analytical | water | Cu | 18.9271 | 19.1114 | 0.18428 | 19.0192 | 0.96894 | P.241 | MJW/TWM
2010-09-10 | 51998 - Analytical | water | Zn | 148.4885 | 148.6194 | 0.13088 | 148.554 | 0.088102 | P.241 | MJW/TWM
2010-09-10 | 51998 - Analytical | water | Cd | 5.8648 | 5.8511 | 0.01365 | 5.8579 | 0.23302 | P.241 | MJW/TWM
2010-09-10 | 51998 - Analytical | water | Pb | 19.7592 | 19.8866 | 0.12742 | 19.8229 | 0.64281 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7g QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-09-10 | 52046 - Analytical | water | Cu | 19.9012 | 20.0823 | 0.18102 | 19.9918 | 0.90548 | P.241 | MJW/TWM
2010-09-10 | 52046 - Analytical | water | Zn | 150.2498 | 151.0387 | 0.78886 | 150.6442 | 0.52366 | P.241 | MJW/TWM
2010-09-10 | 52046 - Analytical | water | Cd | 5.9395 | 5.9427 | 0.003217 | 5.9411 | 0.054149 | P.241 | MJW/TWM
2010-09-10 | 52046 - Analytical | water | Pb | 19.5066 | 19.7572 | 0.25064 | 19.6319 | 1.2767 | P.241 | MJW/TWM
2010-09-10 | 52053 - Analytical | water | Cu | 19.5063 | 19.4423 | 0.063967 | 19.4743 | 0.32847 | P.241 | MJW/TWM
2010-09-10 | 52053 - Analytical | water | Zn | 176.0048 | 174.3562 | 1.6487 | 175.1805 | 0.94112 | P.241 | MJW/TWM
2010-09-10 | 52053 - Analytical | water | Cd | 5.8918 | 5.8515 | 0.040281 | 5.8717 | 0.68602 | P.241 | MJW/TWM
2010-09-10 | 52053 - Analytical | water | Pb | 19.6661 | 19.4756 | 0.1905 | 19.5709 | 0.97338 | P.241 | MJW/TWM
2010-09-29 | 52117 - Analytical | water | Cu | 19.9418 | 19.9966 | 0.054774 | 19.9692 | 0.27429 | P.241 | MJW/TWM
2010-09-29 | 52117 - Analytical | water | Zn | 151.5338 | 151.1963 | 0.33745 | 151.3651 | 0.22294 | P.241 | MJW/TWM
2010-09-29 | 52117 - Analytical | water | Cd | 5.8896 | 5.8196 | 0.069982 | 5.8546 | 1.1953 | P.241 | MJW/TWM
2010-09-29 | 52117 - Analytical | water | Pb | 19.387 | 19.3148 | 0.07219 | 19.3509 | 0.37306 | P.241 | MJW/TWM
2010-09-29 | 52125 - Analytical | water | Cu | 19.0894 | 19.2154 | 0.12605 | 19.1524 | 0.65815 | P.241 | MJW/TWM
2010-09-29 | 52125 - Analytical | water | Zn | 174.3165 | 175.306 | 0.98948 | 174.8112 | 0.56603 | P.241 | MJW/TWM
2010-09-29 | 52125 - Analytical | water | Cd | 5.8335 | 5.8588 | 0.02536 | 5.8462 | 0.43379 | P.241 | MJW/TWM
2010-09-29 | 52125 - Analytical | water | Pb | 19.1641 | 19.2538 | 0.089701 | 19.209 | 0.46697 | P.241 | MJW/TWM
2010-09-29 | 52269 - Analytical | water | Cu | 20.0832 | 20.2362 | 0.15297 | 20.1597 | 0.75879 | P.241 | MJW/TWM
2010-09-29 | 52269 - Analytical | water | Zn | 152.4623 | 153.5269 | 1.0646 | 152.9946 | 0.69587 | P.241 | MJW/TWM
2010-09-29 | 52269 - Analytical | water | Cd | 5.8668 | 5.8962 | 0.029366 | 5.8815 | 0.4993 | P.241 | MJW/TWM
2010-09-29 | 52269 - Analytical | water | Pb | 19.2184 | 19.4228 | 0.2044 | 19.3206 | 1.0579 | P.241 | MJW/TWM
2010-09-29 | 52283 - Analytical | water | Cu | 19.7815 | 19.7162 | 0.065257 | 19.7489 | 0.33043 | P.241 | MJW/TWM
2010-09-29 | 52283 - Analytical | water | Zn | 155.7136 | 154.7197 | 0.99384 | 155.2166 | 0.64029 | P.241 | MJW/TWM
2010-09-29 | 52283 - Analytical | water | Cd | 6.782 | 6.8269 | 0.044888 | 6.8044 | 0.65969 | P.241 | MJW/TWM
2010-09-29 | 52283 - Analytical | water | Pb | 19.5392 | 19.4055 | 0.13368 | 19.4723 | 0.68653 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-7h QC
Table 3-7 (continued). Relative percent difference for duplicate ICPMS analysis of water samples
from sturgeon exposure studies.
BIDa | Duplicate Type | Matrix | Element | Dup 1 | Dup 2 | Diffb | Mean | RPDc | ISOPd | Oper. Init.
2010-10-05 | 52354 - Analytical | water | Cu | 20.2992 | 20.5564 | 0.25722 | 20.4278 | 1.2592 | P.241 | MJW/TWM
2010-10-05 | 52354 - Analytical | water | Zn | 153.0433 | 155.111 | 2.0677 | 154.0772 | 1.342 | P.241 | MJW/TWM
2010-10-05 | 52354 - Analytical | water | Cd | 6.0489 | 6.0705 | 0.021582 | 6.0597 | 0.35615 | P.241 | MJW/TWM
2010-10-05 | 52354 - Analytical | water | Pb | 19.696 | 19.6691 | 0.026896 | 19.6825 | 0.13665 | P.241 | MJW/TWM
2010-10-05 | 52371 - Analytical | water | Cu | 19.1016 | 19.098 | 0.003612 | 19.0998 | 0.018911 | P.241 | MJW/TWM
2010-10-05 | 52371 - Analytical | water | Zn | 149.8755 | 149.9256 | 0.05011 | 149.9005 | 0.033429 | P.241 | MJW/TWM
2010-10-05 | 52371 - Analytical | water | Cd | 6.7839 | 6.8334 | 0.049475 | 6.8086 | 0.72665 | P.241 | MJW/TWM
2010-10-05 | 52371 - Analytical | water | Pb | 19.4712 | 19.2001 | 0.27119 | 19.3357 | 1.4026 | P.241 | MJW/TWM
2010-10-14 | 52473 - Analytical | water | Cu | 20.5955 | 20.694 | 0.098505 | 20.6448 | 0.47714 | P.241 | MJW/TWM
2010-10-14 | 52473 - Analytical | water | Zn | 151.792 | 152.3897 | 0.59775 | 152.0908 | 0.39302 | P.241 | MJW/TWM
2010-10-14 | 52473 - Analytical | water | Cd | 5.8654 | 5.8693 | 0.003931 | 5.8673 | 0.066998 | P.241 | MJW/TWM
2010-10-14 | 52473 - Analytical | water | Pb | 19.1825 | 19.1056 | 0.076829 | 19.144 | 0.40132 | P.241 | MJW/TWM
2010-10-14 | 52480 - Analytical | water | Cu | 19.8061 | 19.7153 | 0.090759 | 19.7607 | 0.45929 | P.241 | MJW/TWM
2010-10-14 | 52480 - Analytical | water | Zn | 211.873 | 210.7962 | 1.0768 | 211.3346 | 0.50953 | P.241 | MJW/TWM
2010-10-14 | 52480 - Analytical | water | Cd | 5.9993 | 5.9555 | 0.043811 | 5.9774 | 0.73294 | P.241 | MJW/TWM
2010-10-14 | 52480 - Analytical | water | Pb | 19.3876 | 19.341 | 0.046575 | 19.3643 | 0.24052 | P.241 | MJW/TWM
2010-10-14 | 52515 - Analytical | water | Cu | 19.7722 | 19.7579 | 0.014205 | 19.765 | 0.071869 | P.241 | MJW/TWM
2010-10-14 | 52515 - Analytical | water | Zn | 155.1433 | 154.9547 | 0.18857 | 155.049 | 0.12162 | P.241 | MJW/TWM
2010-10-14 | 52515 - Analytical | water | Cd | 7.7431 | 7.6903 | 0.052792 | 7.7167 | 0.68412 | P.241 | MJW/TWM
2010-10-14 | 52515 - Analytical | water | Pb | 19.1985 | 19.3045 | 0.10604 | 19.2515 | 0.55084 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the
sample as a member of the group or "block."
bDiff = Dup 1 - Dup 2; digestates spiked with mid-range standard prior to analysis.
cRPD = relative percent difference, calculated as Diff/Mean X 100; acceptance criteria +/- 10%.
dISOP = standard operating procedure used for instrumental analysis of sample ( SOP P.241).

## Table 3-8 QC
Table 3-8. Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-07-22 | Cu | 51247 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.051 | 392.1569 | 19.2187 | 95.8383 | P.241 | MJW/TWM
2010-07-22 | Zn | 51247 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.15343 | 977.6127 | 152.5523 | 101.5992 | P.241 | MJW/TWM
2010-07-22 | Cd | 51247 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.001723 | 3480 | 5.9581 | 99.3309 | P.241 | MJW/TWM
2010-07-22 | Pb | 51247 - Analytical | water | ng/mL | 100 | 5 | 20 | -0.000613 | 32600 | 19.8698 | 99.3521 | P.241 | MJW/TWM
2010-07-22 | Cu | 51260 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.0258 | 775.1938 | 19.2509 | 96.1255 | P.241 | MJW/TWM
2010-07-22 | Zn | 51260 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.19556 | 767.0163 | 150.8717 | 100.4507 | P.241 | MJW/TWM
2010-07-22 | Cd | 51260 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.10329 | 58.0906 | 6.0175 | 98.5709 | P.241 | MJW/TWM
2010-07-22 | Pb | 51260 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001188 | 16800 | 19.7449 | 98.7183 | P.241 | MJW/TWM
2010-07-22 | Cu | 51265 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.026059 | 767.4892 | 19.18 | 95.7695 | P.241 | MJW/TWM
2010-07-22 | Zn | 51265 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.18651 | 804.2507 | 151.6411 | 100.9697 | P.241 | MJW/TWM
2010-07-22 | Cd | 51265 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000461 | 13000 | 5.9437 | 99.0694 | P.241 | MJW/TWM
2010-07-22 | Pb | 51265 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.21725 | 92.0615 | 19.8862 | 98.3445 | P.241 | MJW/TWM
2010-07-27 | Cu | 51301 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.17791 | 112.4183 | 19.4286 | 96.2536 | P.241 | MJW/TWM
2010-07-27 | Zn | 51301 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.33258 | 451.0139 | 151.16 | 100.5516 | P.241 | MJW/TWM
2010-07-27 | Cd | 51301 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000667 | 9000 | 5.8797 | 98.0064 | P.241 | MJW/TWM
2010-07-27 | Pb | 51301 - Analytical | water | ng/mL | 100 | 5 | 20 | -0.000735 | 27200 | 19.6022 | 98.0149 | P.241 | MJW/TWM
2010-07-27 | Cu | 51310 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.033112 | 604.0106 | 19.2605 | 96.1368 | P.241 | MJW/TWM
2010-07-27 | Zn | 51310 - Analytical | water | ng/mL | 750 | 5 | 150 | 3.1415 | 47.7479 | 154.5559 | 100.943 | P.241 | MJW/TWM
2010-07-27 | Cd | 51310 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000206 | 29100 | 5.9074 | 98.4532 | P.241 | MJW/TWM
2010-07-27 | Pb | 51310 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.004039 | 4951.7207 | 19.4972 | 97.4658 | P.241 | MJW/TWM
2010-07-27 | Cu | 51319 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.0385 | 519.4805 | 18.8726 | 94.1705 | P.241 | MJW/TWM
2010-07-27 | Zn | 51319 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.28404 | 528.0872 | 148.1849 | 98.6006 | P.241 | MJW/TWM
2010-07-27 | Cd | 51319 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.25877 | 23.1866 | 6.1305 | 97.8616 | P.241 | MJW/TWM
2010-07-27 | Pb | 51319 - Analytical | water | ng/mL | 100 | 5 | 20 | -3e-06 | 6670000 | 19.4147 | 97.0736 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8b QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-07-29 | Cu | 51219 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.14569 | 137.274 | 19.5016 | 96.7794 | P.241 | MJW/TWM
2010-07-29 | Zn | 51219 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.19742 | 759.786 | 153.0613 | 101.9093 | P.241 | MJW/TWM
2010-07-29 | Cd | 51219 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.00071 | 8450 | 5.9814 | 99.7019 | P.241 | MJW/TWM
2010-07-29 | Cu | 51226 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.01923 | 1040.0416 | 19.2401 | 96.1043 | P.241 | MJW/TWM
2010-07-29 | Zn | 51226 - Analytical | water | ng/mL | 750 | 5 | 150 | 2.0023 | 74.9156 | 154.3575 | 101.5701 | P.241 | MJW/TWM
2010-07-29 | Cd | 51226 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.00281 | 2140 | 5.9204 | 98.6257 | P.241 | MJW/TWM
2010-07-29 | Cu | 51387 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.025367 | 788.4259 | 18.8696 | 94.2213 | P.241 | MJW/TWM
2010-07-29 | Zn | 51387 - Analytical | water | ng/mL | 750 | 5 | 150 | 7.6029 | 19.7294 | 155.919 | 98.8775 | P.241 | MJW/TWM
2010-07-29 | Cd | 51387 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001141 | 5260 | 5.7705 | 96.1555 | P.241 | MJW/TWM
2010-07-29 | Cu | 51396 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.026071 | 767.1359 | 19.0224 | 94.9819 | P.241 | MJW/TWM
2010-07-29 | Zn | 51396 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.27943 | 536.8148 | 151.2121 | 100.6218 | P.241 | MJW/TWM
2010-07-29 | Cd | 51396 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.60866 | 9.8577 | 6.4667 | 97.6342 | P.241 | MJW/TWM
2010-08-03 | Cu | 51337 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.081616 | 245.05 | 19.6168 | 97.6759 | P.241 | MJW/TWM
2010-08-03 | Zn | 51337 - Analytical | water | ng/mL | 750 | 5 | 150 | -0.026086 | 5750.2108 | 153.9512 | 102.6515 | P.241 | MJW/TWM
2010-08-03 | Cd | 51337 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.001039 | 5780 | 5.9621 | 99.3857 | P.241 | MJW/TWM
2010-08-03 | Pb | 51337 - Analytical | water | ng/mL | 100 | 5 | 20 | -1.5e-05 | 1300000 | 19.8919 | 99.4596 | P.241 | MJW/TWM
2010-08-03 | Cu | 51343 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.008832 | 2260 | 18.5949 | 92.9302 | P.241 | MJW/TWM
2010-08-03 | Zn | 51343 - Analytical | water | ng/mL | 750 | 5 | 150 | 2.121 | 70.7205 | 148.2705 | 97.433 | P.241 | MJW/TWM
2010-08-03 | Cd | 51343 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000866 | 6930 | 5.8877 | 98.1138 | P.241 | MJW/TWM
2010-08-03 | Pb | 51343 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000255 | 78400 | 19.4039 | 97.018 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8c QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-08-03 | Cu | 51420 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.02273 | 879.8944 | 19.8275 | 99.0236 | P.241 | MJW/TWM
2010-08-03 | Zn | 51420 - Analytical | water | ng/mL | 750 | 5 | 150 | -0.001452 | 103000 | 154.5341 | 103.0237 | P.241 | MJW/TWM
2010-08-03 | Cd | 51420 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.056534 | 106.1308 | 6.0415 | 99.7494 | P.241 | MJW/TWM
2010-08-03 | Pb | 51420 - Analytical | water | ng/mL | 100 | 5 | 20 | -0.000292 | 68500 | 19.8874 | 99.4383 | P.241 | MJW/TWM
2010-08-03 | Cu | 51426 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.022229 | 899.7256 | 20.1247 | 100.5122 | P.241 | MJW/TWM
2010-08-03 | Zn | 51426 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.11956 | 1260 | 158.4921 | 105.5817 | P.241 | MJW/TWM
2010-08-03 | Cd | 51426 - Analytical | water | ng/mL | 30 | 5 | 6 | 3.6e-05 | 167000 | 5.9159 | 98.5976 | P.241 | MJW/TWM
2010-08-03 | Pb | 51426 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.26796 | 74.638 | 20.0548 | 98.9341 | P.241 | MJW/TWM
2010-08-09 | Cu | 51454 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.44728 | 44.7144 | 20.3912 | 99.7197 | P.241 | MJW/TWM
2010-08-09 | Zn | 51454 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.089586 | 1670 | 157.4266 | 104.8913 | P.241 | MJW/TWM
2010-08-09 | Cd | 51454 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000388 | 15500 | 6.0895 | 101.4857 | P.241 | MJW/TWM
2010-08-09 | Pb | 51454 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001581 | 12700 | 20.0404 | 100.1943 | P.241 | MJW/TWM
2010-08-09 | Cu | 51463 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.019955 | 1000 | 19.2536 | 96.1681 | P.241 | MJW/TWM
2010-08-09 | Zn | 51463 - Analytical | water | ng/mL | 750 | 5 | 150 | 8.4546 | 17.7419 | 162.6883 | 102.8225 | P.241 | MJW/TWM
2010-08-09 | Cd | 51463 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000138 | 43500 | 5.849 | 97.4858 | P.241 | MJW/TWM
2010-08-09 | Pb | 51463 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001684 | 11900 | 19.6231 | 98.1073 | P.241 | MJW/TWM
2010-08-09 | Cu | 51493 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.022577 | 885.8573 | 18.8308 | 94.0411 | P.241 | MJW/TWM
2010-08-09 | Zn | 51493 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.069093 | 2170.9869 | 151.1404 | 100.7142 | P.241 | MJW/TWM
2010-08-09 | Cd | 51493 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.0599 | 100.1669 | 5.9629 | 98.383 | P.241 | MJW/TWM
2010-08-09 | Pb | 51493 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001245 | 16100 | 19.478 | 97.3836 | P.241 | MJW/TWM
2010-08-09 | Cu | 51499 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.024314 | 822.5714 | 18.7747 | 93.7517 | P.241 | MJW/TWM
2010-08-09 | Zn | 51499 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.073471 | 2040 | 151.5269 | 100.969 | P.241 | MJW/TWM
2010-08-09 | Cd | 51499 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000304 | 19700 | 5.8783 | 97.9761 | P.241 | MJW/TWM
2010-08-09 | Pb | 51499 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.29881 | 66.9326 | 19.736 | 97.1861 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8d QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-08-13 | Cu | 51510 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.23248 | 86.0296 | 19.6729 | 97.2022 | P.241 | MJW/TWM
2010-08-13 | Zn | 51510 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.27161 | 552.2665 | 152.3376 | 101.3773 | P.241 | MJW/TWM
2010-08-13 | Cd | 51510 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000479 | 12500 | 6.0047 | 100.0702 | P.241 | MJW/TWM
2010-08-13 | Pb | 51510 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000492 | 40700 | 19.7317 | 98.6559 | P.241 | MJW/TWM
2010-08-13 | Cu | 51517 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.020343 | 983.1392 | 19.3787 | 96.7917 | P.241 | MJW/TWM
2010-08-13 | Zn | 51517 - Analytical | water | ng/mL | 750 | 5 | 150 | 10.0446 | 14.9333 | 164.2745 | 102.8199 | P.241 | MJW/TWM
2010-08-13 | Cd | 51517 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000458 | 328000 | 5.7914 | 96.5313 | P.241 | MJW/TWM
2010-08-13 | Pb | 51517 - Analytical | water | ng/mL | 100 | 5 | 20 | 6.1e-05 | 328000 | 19.2781 | 96.3901 | P.241 | MJW/TWM
2010-08-13 | Cu | 51544 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.028668 | 697.642 | 19.1693 | 95.7033 | P.241 | MJW/TWM
2010-08-13 | Zn | 51544 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.19239 | 779.6541 | 149.5559 | 99.5757 | P.241 | MJW/TWM
2010-08-13 | Cd | 51544 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.28115 | 21.3408 | 6.1572 | 97.9347 | P.241 | MJW/TWM
2010-08-13 | Pb | 51544 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002353 | 8499.7875 | 18.8195 | 94.0859 | P.241 | MJW/TWM
2010-08-13 | Cu | 51552 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.022603 | 884.8383 | 19.5826 | 97.7998 | P.241 | MJW/TWM
2010-08-13 | Zn | 51552 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.20669 | 725.707 | 153.2595 | 102.0352 | P.241 | MJW/TWM
2010-08-13 | Cd | 51552 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000638 | 9400 | 5.9284 | 98.796 | P.241 | MJW/TWM
2010-08-13 | Pb | 51552 - Analytical | water | ng/mL | 100 | 5 | 20 | 1.3439 | 14.8822 | 20.9436 | 97.9984 | P.241 | MJW/TWM
2010-08-17 | Cu | 51645 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.28367 | 70.5045 | 19.7917 | 97.5399 | P.241 | MJW/TWM
2010-08-17 | Zn | 51645 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.28557 | 525.2745 | 153.3306 | 102.03 | P.241 | MJW/TWM
2010-08-17 | Cd | 51645 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000158 | 38000 | 5.9729 | 99.5513 | P.241 | MJW/TWM
2010-08-17 | Pb | 51645 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002313 | 8650 | 19.6488 | 98.2326 | P.241 | MJW/TWM
2010-08-17 | Cu | 51654 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.026619 | 751.343 | 19.2119 | 95.9265 | P.241 | MJW/TWM
2010-08-17 | Zn | 51654 - Analytical | water | ng/mL | 750 | 5 | 150 | 13.3863 | 11.2055 | 164.6273 | 100.8273 | P.241 | MJW/TWM
2010-08-17 | Cd | 51654 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000606 | 9900.9901 | 5.9203 | 98.6813 | P.241 | MJW/TWM
2010-08-17 | Pb | 51654 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002197 | 9103.3227 | 19.4927 | 97.4523 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8e QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-08-18 | Cu | 51675 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.044469 | 449.7515 | 19.5914 | 97.7349 | P.241 | MJW/TWM
2010-08-18 | Zn | 51675 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.16842 | 890.6253 | 154.7704 | 103.068 | P.241 | MJW/TWM
2010-08-18 | Cd | 51675 - Analytical | water | ng/mL | 30 | 5 | 6 | 3.3e-05 | 182000 | 6.0311 | 100.5175 | P.241 | MJW/TWM
2010-08-18 | Pb | 51675 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000727 | 27500 | 19.8407 | 99.2001 | P.241 | MJW/TWM
2010-08-18 | Cu | 51693 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.023109 | 865.4637 | 19.2945 | 96.3571 | P.241 | MJW/TWM
2010-08-18 | Zn | 51693 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.091815 | 1630 | 151.112 | 100.6801 | P.241 | MJW/TWM
2010-08-18 | Cd | 51693 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000207 | 29000 | 5.8224 | 97.0439 | P.241 | MJW/TWM
2010-08-18 | Pb | 51693 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.26092 | 76.653 | 19.3994 | 95.6925 | P.241 | MJW/TWM
2010-08-26 | Cu | 51853 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.18541 | 107.869 | 19.2508 | 95.3267 | P.241 | MJW/TWM
2010-08-26 | Zn | 51853 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.43932 | 341.4383 | 152.9128 | 101.649 | P.241 | MJW/TWM
2010-08-26 | Cd | 51853 - Analytical | water | ng/mL | 30 | 5 | 6 | -0.000315 | 19000 | 5.8033 | 96.7274 | P.241 | MJW/TWM
2010-08-26 | Pb | 51853 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000478 | 41800 | 19.0941 | 95.468 | P.241 | MJW/TWM
2010-08-26 | Cu | 51860 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.024313 | 822.6052 | 19.2795 | 96.2761 | P.241 | MJW/TWM
2010-08-26 | Zn | 51860 - Analytical | water | ng/mL | 750 | 5 | 150 | 6.4195 | 23.3664 | 149.7969 | 95.5849 | P.241 | MJW/TWM
2010-08-26 | Cd | 51860 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000346 | 17300 | 5.8559 | 97.5932 | P.241 | MJW/TWM
2010-08-26 | Pb | 51860 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001752 | 11400 | 19.2291 | 96.1368 | P.241 | MJW/TWM
2010-08-26 | Cu | 51890 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.024136 | 828.6377 | 19.7632 | 98.6954 | P.241 | MJW/TWM
2010-08-26 | Zn | 51890 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.09704 | 1545.7543 | 155.65 | 103.702 | P.241 | MJW/TWM
2010-08-26 | Cd | 51890 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.066257 | 90.5565 | 5.8459 | 96.3277 | P.241 | MJW/TWM
2010-08-26 | Pb | 51890 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000243 | 82304.5267 | 19.1538 | 95.7676 | P.241 | MJW/TWM
2010-08-26 | Cu | 51897 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.026995 | 740.8779 | 19.9306 | 99.5183 | P.241 | MJW/TWM
2010-08-26 | Zn | 51897 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.10507 | 1430 | 156.706 | 104.4006 | P.241 | MJW/TWM
2010-08-26 | Cd | 51897 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000969 | 6190 | 5.7972 | 96.6037 | P.241 | MJW/TWM
2010-08-26 | Pb | 51897 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.6079 | 32.9004 | 19.6509 | 95.2149 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8f QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-09-07 | Cu | 51953 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.28972 | 69.0331 | 20.121 | 99.1566 | P.241 | MJW/TWM
2010-09-07 | Zn | 51953 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.21857 | 686.2884 | 155.1501 | 103.2877 | P.241 | MJW/TWM
2010-09-07 | Cd | 51953 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001068 | 5620 | 6.1064 | 101.7564 | P.241 | MJW/TWM
2010-09-07 | Pb | 51953 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.003852 | 5190 | 20.0682 | 100.3219 | P.241 | MJW/TWM
2010-09-07 | Cu | 51962 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.019034 | 1050 | 19.9677 | 99.7433 | P.241 | MJW/TWM
2010-09-07 | Zn | 51962 - Analytical | water | ng/mL | 750 | 5 | 150 | 13.1446 | 11.4115 | 169.2466 | 104.068 | P.241 | MJW/TWM
2010-09-07 | Cd | 51962 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001073 | 5590 | 6.0481 | 100.7838 | P.241 | MJW/TWM
2010-09-07 | Pb | 51962 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000972 | 20600 | 19.7939 | 98.9646 | P.241 | MJW/TWM
2010-09-07 | Cu | 52025 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.033346 | 599.7721 | 19.6158 | 97.9121 | P.241 | MJW/TWM
2010-09-07 | Zn | 52025 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.22613 | 663.3411 | 154.4534 | 102.8182 | P.241 | MJW/TWM
2010-09-07 | Cd | 52025 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.069452 | 86.3906 | 5.9375 | 97.801 | P.241 | MJW/TWM
2010-09-07 | Pb | 52025 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001888 | 10600 | 19.3948 | 96.9646 | P.241 | MJW/TWM
2010-09-07 | Cu | 52031 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.032365 | 617.9515 | 19.3309 | 96.4929 | P.241 | MJW/TWM
2010-09-07 | Zn | 52031 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.28278 | 530.4439 | 151.9739 | 101.1274 | P.241 | MJW/TWM
2010-09-07 | Cd | 52031 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.002301 | 2610 | 5.8384 | 97.2684 | P.241 | MJW/TWM
2010-09-07 | Pb | 52031 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.33252 | 60.1462 | 19.7542 | 97.1084 | P.241 | MJW/TWM
2010-09-10 | Cu | 51992 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.068439 | 292.231 | 19.5604 | 97.4599 | P.241 | MJW/TWM
2010-09-10 | Zn | 51992 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.13354 | 1120 | 150.9835 | 100.5666 | P.241 | MJW/TWM
2010-09-10 | Cd | 51992 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.070301 | 85.3473 | 5.8818 | 96.8585 | P.241 | MJW/TWM
2010-09-10 | Pb | 51992 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001293 | 15500 | 19.5038 | 97.5124 | P.241 | MJW/TWM
2010-09-10 | Cu | 51998 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.051929 | 385.1413 | 19.2846 | 96.1634 | P.241 | MJW/TWM
2010-09-10 | Zn | 51998 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.19641 | 763.7086 | 147.9193 | 98.4819 | P.241 | MJW/TWM
2010-09-10 | Cd | 51998 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001467 | 4089.9796 | 5.8656 | 97.7357 | P.241 | MJW/TWM
2010-09-10 | Pb | 51998 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.34401 | 58.1373 | 19.5559 | 96.0596 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8g QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-09-10 | Cu | 52046 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.74505 | 26.844 | 20.0477 | 96.5133 | P.241 | MJW/TWM
2010-09-10 | Zn | 52046 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.12652 | 1190 | 151.2434 | 100.7446 | P.241 | MJW/TWM
2010-09-10 | Cd | 52046 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000311 | 19300 | 6.0133 | 100.2161 | P.241 | MJW/TWM
2010-09-10 | Pb | 52046 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002284 | 8760 | 19.7087 | 98.532 | P.241 | MJW/TWM
2010-09-10 | Cu | 52053 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.023352 | 856.4577 | 19.4976 | 97.3712 | P.241 | MJW/TWM
2010-09-10 | Zn | 52053 - Analytical | water | ng/mL | 750 | 5 | 150 | 25.3409 | 5.9193 | 174.305 | 99.3094 | P.241 | MJW/TWM
2010-09-10 | Cd | 52053 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000894 | 6710 | 5.8715 | 97.8442 | P.241 | MJW/TWM
2010-09-10 | Pb | 52053 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002023 | 9890 | 19.6233 | 98.1062 | P.241 | MJW/TWM
2010-09-29 | Cu | 52117 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.73049 | 27.3788 | 19.9132 | 95.9134 | P.241 | MJW/TWM
2010-09-29 | Zn | 52117 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.20525 | 730.8232 | 150.8885 | 100.4555 | P.241 | MJW/TWM
2010-09-29 | Cd | 52117 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000732 | 8196.7213 | 5.822 | 97.0212 | P.241 | MJW/TWM
2010-09-29 | Pb | 52117 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001181 | 16934.801 | 19.3925 | 96.9565 | P.241 | MJW/TWM
2010-09-29 | Cu | 52125 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.020621 | 969.8851 | 18.9817 | 94.8054 | P.241 | MJW/TWM
2010-09-29 | Zn | 52125 - Analytical | water | ng/mL | 750 | 5 | 150 | 24.1916 | 6.2005 | 173.7512 | 99.7064 | P.241 | MJW/TWM
2010-09-29 | Cd | 52125 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000799 | 7509.3867 | 5.779 | 96.3027 | P.241 | MJW/TWM
2010-09-29 | Pb | 52125 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001037 | 19286.4031 | 19.1263 | 95.6262 | P.241 | MJW/TWM
2010-09-29 | Cu | 52269 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.75739 | 26.4065 | 20.1561 | 96.9934 | P.241 | MJW/TWM
2010-09-29 | Zn | 52269 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.17817 | 841.8973 | 151.7346 | 101.0376 | P.241 | MJW/TWM
2010-09-29 | Cd | 52269 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001008 | 5952.381 | 5.8321 | 97.1842 | P.241 | MJW/TWM
2010-09-29 | Pb | 52269 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.000796 | 25125.6281 | 19.1691 | 95.8414 | P.241 | MJW/TWM
2010-09-29 | Cu | 52283 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.016457 | 1215.2883 | 19.6165 | 98.0002 | P.241 | MJW/TWM
2010-09-29 | Zn | 52283 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.38803 | 386.5671 | 153.4499 | 102.0412 | P.241 | MJW/TWM
2010-09-29 | Cd | 52283 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.88588 | 6.773 | 6.7074 | 97.0254 | P.241 | MJW/TWM
2010-09-29 | Pb | 52283 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.00636 | 3144.6541 | 19.3005 | 96.4707 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-8h QC
Table 3-8 (continued). Percent recovery of elements spiked in sturgeon exposure water and analyzed by ICP-MS.
|  |  |  | Analysis | Spk Amt.b | Vol. | Effectivec | Bkgd.d | Spk/ | Totale | %
BIDa | Ele. | Spk Type | Matrix | Units | µg | (mL) | Conc. | Conc. | Bkgd | Conc. | Rec.f | ISOP | Oper. Init.
2010-10-05 | Cu | 52354 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.88025 | 22.7209 | 20.2749 | 96.9731 | P.241 | MJW/TWM
2010-10-05 | Zn | 52354 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.19423 | 772.2763 | 152.6435 | 101.6329 | P.241 | MJW/TWM
2010-10-05 | Cd | 52354 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000696 | 8620 | 5.9947 | 99.9003 | P.241 | MJW/TWM
2010-10-05 | Pb | 52354 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.002023 | 9890 | 19.6301 | 98.1405 | P.241 | MJW/TWM
2010-10-05 | Cu | 52371 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.019246 | 1040 | 18.6925 | 93.3664 | P.241 | MJW/TWM
2010-10-05 | Zn | 52371 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.20396 | 735.4311 | 147.1356 | 97.9544 | P.241 | MJW/TWM
2010-10-05 | Cd | 52371 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.88173 | 6.8048 | 6.705 | 97.0552 | P.241 | MJW/TWM
2010-10-05 | Pb | 52371 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001941 | 10300 | 19.2438 | 96.2092 | P.241 | MJW/TWM
2010-10-14 | Cu | 52473 - Analytical | water | ng/mL | 100 | 5 | 20 | 1.2446 | 16.0693 | 20.5155 | 96.3544 | P.241 | MJW/TWM
2010-10-14 | Zn | 52473 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.6177 | 242.8348 | 152.7686 | 101.4339 | P.241 | MJW/TWM
2010-10-14 | Cd | 52473 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.001251 | 4800 | 5.8667 | 97.7569 | P.241 | MJW/TWM
2010-10-14 | Pb | 52473 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001288 | 15500 | 19.2141 | 96.064 | P.241 | MJW/TWM
2010-10-14 | Cu | 52480 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.02651 | 754.4323 | 19.6081 | 97.9079 | P.241 | MJW/TWM
2010-10-14 | Zn | 52480 - Analytical | water | ng/mL | 750 | 5 | 150 | 56.8222 | 2.6398 | 209.6493 | 101.8847 | P.241 | MJW/TWM
2010-10-14 | Cd | 52480 - Analytical | water | ng/mL | 30 | 5 | 6 | 0.000772 | 7770 | 5.8646 | 97.7313 | P.241 | MJW/TWM
2010-10-14 | Pb | 52480 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001416 | 14100 | 19.0927 | 95.4566 | P.241 | MJW/TWM
2010-10-14 | Cu | 52515 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.02648 | 755.287 | 19.5868 | 97.8015 | P.241 | MJW/TWM
2010-10-14 | Zn | 52515 - Analytical | water | ng/mL | 750 | 5 | 150 | 0.10072 | 1490 | 154.2075 | 102.7379 | P.241 | MJW/TWM
2010-10-14 | Cd | 52515 - Analytical | water | ng/mL | 30 | 5 | 6 | 1.8012 | 3.3311 | 7.6831 | 98.031 | P.241 | MJW/TWM
2010-10-14 | Pb | 52515 - Analytical | water | ng/mL | 100 | 5 | 20 | 0.001255 | 15900 | 18.9529 | 94.7581 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
as a member of the group or "block."
bSpike Amt. µg = the absolute microgram (µg) amount of the spike which was added to a sample.
cEffective Conc. = the Spike Amt (ng) divided by the sample volume (mL), units ng/mL.
dUnspiked Conc. = the measured concentration of the sample prior to spiking, units ng/mL.
eSpiked Conc. = the measured concentration of the spiked sample (spike + unspiked, units ng/mL).
f% Rec. = percent recovery: [(Spiked Conc. - Unspiked Conc.)/Effective Conc. * 100]

## Table 3-9 QC
Table 3-9. Interference check using dilution percent difference during
ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-07-22 | 51247 - Analytical | water | Cu | 19.3969 | 3.8426 | 19.2128 | 0.94936
2010-07-22 | 51247 - Analytical | water | Zn | 151.9658 | 30.4278 | 152.1389 | 0.11391
2010-07-22 | 51247 - Analytical | water | Cd | 5.9757 | 1.1813 | 5.9065 | 1.1579
2010-07-22 | 51247 - Analytical | water | Pb | 19.8349 | 3.9433 | 19.7163 | 0.59794
2010-07-22 | 51260 - Analytical | water | Cu | 19.2927 | 3.8562 | 19.281 | 0.060665
2010-07-22 | 51260 - Analytical | water | Zn | 151.8488 | 30.5795 | 152.8975 | 0.69062
2010-07-22 | 51260 - Analytical | water | Cd | 6.0283 | 1.202 | 6.0102 | 0.30062
2010-07-22 | 51260 - Analytical | water | Pb | 19.6911 | 3.9783 | 19.8915 | 1.0177
2010-07-22 | 51265 - Analytical | water | Cu | 19.6388 | 3.9225 | 19.6125 | 0.13384
2010-07-22 | 51265 - Analytical | water | Zn | 155.7726 | 31.4058 | 157.0289 | 0.80649
2010-07-22 | 51265 - Analytical | water | Cd | 6.0228 | 1.184 | 5.92 | 1.7067
2010-07-22 | 51265 - Analytical | water | Pb | 19.9792 | 3.9765 | 19.8824 | 0.48461
2010-07-27 | 51301 - Analytical | water | Cu | 19.4185 | 3.884 | 19.4199 | 0.0076577
2010-07-27 | 51301 - Analytical | water | Zn | 151.3375 | 30.2259 | 151.1293 | 0.13752
2010-07-27 | 51301 - Analytical | water | Cd | 5.9136 | 1.1584 | 5.7921 | 2.0549
2010-07-27 | 51301 - Analytical | water | Pb | 19.7502 | 3.959 | 19.7948 | 0.2256
2010-07-27 | 51310 - Analytical | water | Cu | 19.1353 | 3.8547 | 19.2734 | 0.72175
2010-07-27 | 51310 - Analytical | water | Zn | 153.1193 | 30.8827 | 154.4135 | 0.84523
2010-07-27 | 51310 - Analytical | water | Cd | 5.9341 | 1.1706 | 5.853 | 1.3672
2010-07-27 | 51310 - Analytical | water | Pb | 19.5139 | 3.8679 | 19.3395 | 0.89368
2010-07-27 | 51319 - Analytical | water | Cu | 19.2213 | 3.8316 | 19.1578 | 0.33038
2010-07-27 | 51319 - Analytical | water | Zn | 152.1268 | 30.5993 | 152.9963 | 0.57156
2010-07-27 | 51319 - Analytical | water | Cd | 6.2224 | 1.2518 | 6.2588 | 0.58516
2010-07-27 | 51319 - Analytical | water | Pb | 19.6402 | 3.8828 | 19.414 | 1.1518
2010-08-17 | 51645 - Analytical | water | Cu | 19.6653 | 3.9477 | 19.7384 | 0.37159
2010-08-17 | 51645 - Analytical | water | Zn | 153.1474 | 30.7097 | 153.5486 | 0.26196
2010-08-17 | 51645 - Analytical | water | Cd | 5.9147 | 1.1644 | 5.822 | 1.5676
2010-08-17 | 51645 - Analytical | water | Pb | 19.5358 | 3.9221 | 19.6104 | 0.38192
2010-08-17 | 51654 - Analytical | water | Cu | 19.189 | 3.8439 | 19.2196 | 0.15971
2010-08-17 | 51654 - Analytical | water | Zn | 162.8725 | 32.9202 | 164.601 | 1.0612
2010-08-17 | 51654 - Analytical | water | Cd | 5.9173 | 1.1815 | 5.9077 | 0.1622
2010-08-17 | 51654 - Analytical | water | Pb | 19.4094 | 3.8962 | 19.4812 | 0.36999
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9b QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-07-29 | 51219 - Analytical | water | Cu | 19.6138 | 4.0939 | 20.4697 | 4.3638
2010-07-29 | 51219 - Analytical | water | Zn | 153.2288 | 31.9721 | 159.8603 | 4.3278
2010-07-29 | 51219 - Analytical | water | Cd | 5.9699 | 1.2262 | 6.1309 | 2.6967
2010-07-29 | 51226 - Analytical | water | Cu | 19.3293 | 3.9469 | 19.7343 | 2.0948
2010-07-29 | 51226 - Analytical | water | Zn | 153.8387 | 31.6197 | 158.0987 | 2.7692
2010-07-29 | 51226 - Analytical | water | Cd | 5.9152 | 1.1846 | 5.9231 | 0.13374
2010-07-29 | 51387 - Analytical | water | Cu | 18.8522 | 3.7644 | 18.8222 | 0.15922
2010-07-29 | 51387 - Analytical | water | Zn | 156.5226 | 31.5626 | 157.8132 | 0.82458
2010-07-29 | 51387 - Analytical | water | Cd | 5.8209 | 1.1525 | 5.7625 | 1.004
2010-07-29 | 51396 - Analytical | water | Cu | 19.3316 | 3.8368 | 19.184 | 0.76369
2010-07-29 | 51396 - Analytical | water | Zn | 153.1729 | 30.3208 | 151.6042 | 1.0241
2010-07-29 | 51396 - Analytical | water | Cd | 6.5621 | 1.2758 | 6.379 | 2.7901
2010-08-03 | 51337 - Analytical | water | Cu | 20.3409 | 4.0597 | 20.2986 | 0.20796
2010-08-03 | 51337 - Analytical | water | Zn | 157.9518 | 31.9147 | 159.5733 | 1.0266
2010-08-03 | 51337 - Analytical | water | Cd | 5.8395 | 1.1824 | 5.9122 | 1.2454
2010-08-03 | 51337 - Analytical | water | Pb | 19.8545 | 3.9529 | 19.7647 | 0.45234
2010-08-03 | 51343 - Analytical | water | Cu | 18.7728 | 3.742 | 18.7099 | 0.33527
2010-08-03 | 51343 - Analytical | water | Zn | 148.4042 | 29.9173 | 149.5865 | 0.79668
2010-08-03 | 51343 - Analytical | water | Cd | 5.9518 | 1.1768 | 5.8841 | 1.1382
2010-08-03 | 51343 - Analytical | water | Pb | 19.1648 | 3.838 | 19.1899 | 0.13094
2010-08-03 | 51420 - Analytical | water | Cu | 19.7341 | 3.9364 | 19.6822 | 0.26332
2010-08-03 | 51420 - Analytical | water | Zn | 153.8604 | 30.8687 | 154.3435 | 0.31395
2010-08-03 | 51420 - Analytical | water | Cd | 6.0359 | 1.1908 | 5.954 | 1.3573
2010-08-03 | 51420 - Analytical | water | Pb | 19.8564 | 3.9246 | 19.6232 | 1.1746
2010-08-03 | 51426 - Analytical | water | Cu | 20.0957 | 3.9911 | 19.9556 | 0.69718
2010-08-03 | 51426 - Analytical | water | Zn | 157.7298 | 31.4884 | 157.4422 | 0.18238
2010-08-03 | 51426 - Analytical | water | Cd | 5.8521 | 1.1526 | 5.7631 | 1.5203
2010-08-03 | 51426 - Analytical | water | Pb | 19.9974 | 4.0141 | 20.0704 | 0.36529
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9c QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-08-09 | 51454 - Analytical | water | Cu | 20.5056 | 4.0593 | 20.2965 | 1.0193
2010-08-09 | 51454 - Analytical | water | Zn | 157.6873 | 31.3457 | 156.7287 | 0.6079
2010-08-09 | 51454 - Analytical | water | Cd | 6.0651 | 1.2037 | 6.0183 | 0.77046
2010-08-09 | 51454 - Analytical | water | Pb | 19.8346 | 3.9102 | 19.5509 | 1.43
2010-08-09 | 51463 - Analytical | water | Cu | 19.2886 | 3.8726 | 19.3629 | 0.38507
2010-08-09 | 51463 - Analytical | water | Zn | 163.3089 | 32.5408 | 162.7041 | 0.37033
2010-08-09 | 51463 - Analytical | water | Cd | 5.8708 | 1.1702 | 5.8508 | 0.34047
2010-08-09 | 51463 - Analytical | water | Pb | 19.3934 | 3.8966 | 19.4832 | 0.46293
2010-08-09 | 51493 - Analytical | water | Cu | 18.85 | 3.7683 | 18.8416 | 0.044721
2010-08-09 | 51493 - Analytical | water | Zn | 151.2465 | 30.2994 | 151.4968 | 0.1655
2010-08-09 | 51493 - Analytical | water | Cd | 5.8947 | 1.1816 | 5.908 | 0.22544
2010-08-09 | 51493 - Analytical | water | Pb | 19.614 | 3.8689 | 19.3447 | 1.3727
2010-08-09 | 51499 - Analytical | water | Cu | 18.7187 | 3.7509 | 18.7546 | 0.19149
2010-08-09 | 51499 - Analytical | water | Zn | 149.9002 | 30.1007 | 150.5035 | 0.40245
2010-08-09 | 51499 - Analytical | water | Cd | 5.8495 | 1.1641 | 5.8204 | 0.49782
2010-08-09 | 51499 - Analytical | water | Pb | 19.6446 | 3.8736 | 19.3681 | 1.4072
2010-08-13 | 51510 - Analytical | water | Cu | 19.7168 | 3.9614 | 19.8068 | 0.45669
2010-08-13 | 51510 - Analytical | water | Zn | 152.3313 | 30.6802 | 153.4008 | 0.70206
2010-08-13 | 51510 - Analytical | water | Cd | 6.0228 | 1.18 | 5.9 | 2.0375
2010-08-13 | 51510 - Analytical | water | Pb | 19.66 | 3.9344 | 19.6722 | 0.062116
2010-08-13 | 51517 - Analytical | water | Cu | 19.4488 | 3.9075 | 19.5375 | 0.45594
2010-08-13 | 51517 - Analytical | water | Zn | 164.5775 | 33.0735 | 165.3677 | 0.48016
2010-08-13 | 51517 - Analytical | water | Cd | 5.8595 | 1.2401 | 6.2006 | 5.822
2010-08-13 | 51517 - Analytical | water | Pb | 19.1613 | 3.8563 | 19.2813 | 0.62602
2010-08-13 | 51544 - Analytical | water | Cu | 19.2993 | 3.9269 | 19.6344 | 1.7362
2010-08-13 | 51544 - Analytical | water | Zn | 150.4563 | 30.7571 | 153.7853 | 2.2126
2010-08-13 | 51544 - Analytical | water | Cd | 5.92 | 1.1771 | 5.8855 | 0.58277
2010-08-13 | 51544 - Analytical | water | Pb | 19.5576 | 3.9369 | 19.6844 | 0.64865
2010-08-13 | 51552 - Analytical | water | Cu | 19.6844 | 3.925 | 19.6248 | 0.30277
2010-08-13 | 51552 - Analytical | water | Zn | 153.1099 | 30.5361 | 152.6807 | 0.28033
2010-08-13 | 51552 - Analytical | water | Cd | 5.95 | 1.1438 | 5.7192 | 3.8791
2010-08-13 | 51552 - Analytical | water | Pb | 20.84 | 4.1782 | 20.891 | 0.24472
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9d QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-08-18 | 51675 - Analytical | water | Cu | 19.3671 | 3.9167 | 19.5836 | 1.1178
2010-08-18 | 51675 - Analytical | water | Zn | 152.3206 | 30.9118 | 154.5589 | 1.4695
2010-08-18 | 51675 - Analytical | water | Cd | 5.9543 | 1.1996 | 5.9979 | 0.73256
2010-08-18 | 51675 - Analytical | water | Pb | 19.6874 | 3.9407 | 19.7035 | 0.081616
2010-08-18 | 51693 - Analytical | water | Cu | 19.1332 | 3.8433 | 19.2163 | 0.43417
2010-08-18 | 51693 - Analytical | water | Zn | 150.9724 | 30.2362 | 151.1812 | 0.13832
2010-08-18 | 51693 - Analytical | water | Cd | 5.8786 | 1.1428 | 5.714 | 2.8007
2010-08-18 | 51693 - Analytical | water | Pb | 19.5871 | 3.8905 | 19.4525 | 0.68734
2010-08-26 | 51853 - Analytical | water | Cu | 19.1927 | 3.9456 | 19.7281 | 2.7894
2010-08-26 | 51853 - Analytical | water | Zn | 152.9834 | 30.7068 | 153.534 | 0.35996
2010-08-26 | 51853 - Analytical | water | Cd | 5.7875 | 1.1464 | 5.732 | 0.95864
2010-08-26 | 51853 - Analytical | water | Pb | 18.9158 | 3.7892 | 18.9461 | 0.16001
2010-08-26 | 51860 - Analytical | water | Cu | 19.143 | 3.8816 | 19.4082 | 1.3853
2010-08-26 | 51860 - Analytical | water | Zn | 149.2366 | 29.9075 | 149.5376 | 0.20174
2010-08-26 | 51860 - Analytical | water | Cd | 5.8775 | 1.1594 | 5.7968 | 1.3727
2010-08-26 | 51860 - Analytical | water | Pb | 19.6131 | 3.8538 | 19.2689 | 1.7549
2010-08-26 | 51890 - Analytical | water | Cu | 19.5731 | 3.9535 | 19.7676 | 0.99389
2010-08-26 | 51890 - Analytical | water | Zn | 155.9862 | 31.2184 | 156.092 | 0.067808
2010-08-26 | 51890 - Analytical | water | Cd | 5.8189 | 1.1545 | 5.7724 | 0.79941
2010-08-26 | 51890 - Analytical | water | Pb | 19.1883 | 3.8819 | 19.4095 | 1.1528
2010-08-26 | 51897 - Analytical | water | Cu | 19.5054 | 3.878 | 19.3901 | 0.5912
2010-08-26 | 51897 - Analytical | water | Zn | 154.5637 | 30.803 | 154.0148 | 0.35512
2010-08-26 | 51897 - Analytical | water | Cd | 5.7473 | 1.1395 | 5.6975 | 0.86731
2010-08-26 | 51897 - Analytical | water | Pb | 19.5521 | 3.888 | 19.4401 | 0.57263
2010-09-07 | 51953 - Analytical | water | Cu | 19.706 | 3.9452 | 19.7258 | 0.10067
2010-09-07 | 51953 - Analytical | water | Zn | 152.0309 | 31.6687 | 158.3436 | 4.1523
2010-09-07 | 51953 - Analytical | water | Cd | 5.9853 | 1.1855 | 5.9276 | 0.96429
2010-09-07 | 51953 - Analytical | water | Pb | 19.7028 | 3.9721 | 19.8607 | 0.80155
2010-09-07 | 51962 - Analytical | water | Cu | 19.8821 | 3.9409 | 19.7043 | 0.89413
2010-09-07 | 51962 - Analytical | water | Zn | 169.2355 | 33.4936 | 167.4679 | 1.0444
2010-09-07 | 51962 - Analytical | water | Cd | 6.0395 | 1.1724 | 5.8618 | 2.9422
2010-09-07 | 51962 - Analytical | water | Pb | 19.728 | 3.8912 | 19.4559 | 1.3793
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9e QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-09-07 | 52025 - Analytical | water | Cu | 19.9178 | 3.8671 | 19.3354 | 2.924
2010-09-07 | 52025 - Analytical | water | Zn | 155.6669 | 30.8765 | 154.3826 | 0.82501
2010-09-07 | 52025 - Analytical | water | Cd | 6.0645 | 1.1517 | 5.7585 | 5.045
2010-09-07 | 52025 - Analytical | water | Pb | 19.6377 | 3.8424 | 19.212 | 2.1677
2010-09-07 | 52031 - Analytical | water | Cu | 19.2249 | 3.8503 | 19.2515 | 0.13795
2010-09-07 | 52031 - Analytical | water | Zn | 152.5182 | 30.778 | 153.8901 | 0.89952
2010-09-07 | 52031 - Analytical | water | Cd | 5.7772 | 1.1426 | 5.7132 | 1.1069
2010-09-07 | 52031 - Analytical | water | Pb | 19.6802 | 3.9095 | 19.5477 | 0.67287
2010-09-10 | 51992 - Analytical | water | Cu | 19.735 | 3.9172 | 19.5859 | 0.75563
2010-09-10 | 51992 - Analytical | water | Zn | 150.3482 | 29.961 | 149.8048 | 0.36145
2010-09-10 | 51992 - Analytical | water | Cd | 5.8858 | 1.1444 | 5.7221 | 2.78
2010-09-10 | 51992 - Analytical | water | Pb | 19.4369 | 3.8644 | 19.3221 | 0.59076
2010-09-10 | 51998 - Analytical | water | Cu | 19.4897 | 3.8958 | 19.4792 | 0.054044
2010-09-10 | 51998 - Analytical | water | Zn | 149.7634 | 29.939 | 149.6952 | 0.04557
2010-09-10 | 51998 - Analytical | water | Cd | 5.8862 | 1.1578 | 5.7891 | 1.6499
2010-09-10 | 51998 - Analytical | water | Pb | 19.7402 | 3.9322 | 19.6611 | 0.40067
2010-09-10 | 52046 - Analytical | water | Cu | 20.1768 | 3.9897 | 19.9487 | 1.1306
2010-09-10 | 52046 - Analytical | water | Zn | 151.7647 | 30.11 | 150.5498 | 0.80048
2010-09-10 | 52046 - Analytical | water | Cd | 6.0194 | 1.1851 | 5.9254 | 1.5619
2010-09-10 | 52046 - Analytical | water | Pb | 19.6787 | 3.9398 | 19.6988 | 0.10193
2010-09-10 | 52053 - Analytical | water | Cu | 19.6919 | 3.954 | 19.7702 | 0.39753
2010-09-10 | 52053 - Analytical | water | Zn | 176.026 | 35.5057 | 177.5286 | 0.85359
2010-09-10 | 52053 - Analytical | water | Cd | 5.9065 | 1.1752 | 5.8761 | 0.5145
2010-09-10 | 52053 - Analytical | water | Pb | 19.6886 | 3.9433 | 19.7163 | 0.14051
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9f QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-09-29 | 52117 - Analytical | water | Cu | 19.8701 | 4.0325 | 20.1625 | 1.4718
2010-09-29 | 52117 - Analytical | water | Zn | 151.7554 | 30.8021 | 154.0106 | 1.486
2010-09-29 | 52117 - Analytical | water | Cd | 5.8329 | 1.173 | 5.8652 | 0.55363
2010-09-29 | 52117 - Analytical | water | Pb | 19.2919 | 3.8692 | 19.3458 | 0.27933
2010-09-29 | 52125 - Analytical | water | Cu | 19.1273 | 3.7592 | 18.7959 | 1.7324
2010-09-29 | 52125 - Analytical | water | Zn | 175.0712 | 34.8165 | 174.0823 | 0.56485
2010-09-29 | 52125 - Analytical | water | Cd | 5.8033 | 1.1367 | 5.6834 | 2.0664
2010-09-29 | 52125 - Analytical | water | Pb | 19.1196 | 3.7307 | 18.6533 | 2.4387
2010-09-29 | 52269 - Analytical | water | Cu | 20.2797 | 4.0507 | 20.2536 | 0.12903
2010-09-29 | 52269 - Analytical | water | Zn | 152.977 | 30.8109 | 154.0546 | 0.70447
2010-09-29 | 52269 - Analytical | water | Cd | 5.8437 | 1.1582 | 5.7908 | 0.90469
2010-09-29 | 52269 - Analytical | water | Pb | 19.4747 | 3.8929 | 19.4646 | 0.051852
2010-09-29 | 52283 - Analytical | water | Cu | 19.6702 | 3.9758 | 19.8791 | 1.0622
2010-09-29 | 52283 - Analytical | water | Zn | 154.3938 | 31.4119 | 157.0597 | 1.7267
2010-09-29 | 52283 - Analytical | water | Cd | 6.8143 | 1.3446 | 6.7228 | 1.3423
2010-09-29 | 52283 - Analytical | water | Pb | 19.3816 | 3.9175 | 19.5874 | 1.0616
2010-10-05 | 52354 - Analytical | water | Cu | 20.3454 | 4.0688 | 20.344 | 0.0066108
2010-10-05 | 52354 - Analytical | water | Zn | 153.4109 | 31.0209 | 155.1043 | 1.1038
2010-10-05 | 52354 - Analytical | water | Cd | 6.0702 | 1.1948 | 5.974 | 1.5846
2010-10-05 | 52354 - Analytical | water | Pb | 19.7735 | 3.9397 | 19.6983 | 0.38035
2010-10-05 | 52371 - Analytical | water | Cu | 18.9213 | 3.8013 | 19.0066 | 0.45058
2010-10-05 | 52371 - Analytical | water | Zn | 148.4872 | 30.1311 | 150.6556 | 1.4604
2010-10-05 | 52371 - Analytical | water | Cd | 6.717 | 1.3325 | 6.6626 | 0.80913
2010-10-05 | 52371 - Analytical | water | Pb | 19.2058 | 3.8399 | 19.1993 | 0.034099
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-9h QC
Table 3-9 (continued). Interference check using dilution percent difference
during ICP-MS analysis of sturgeon exposure water.
|  |  |  | Undiluted | Diluted | Dil Conc | Dil
BIDa | Dilution Type | Matrix | Element | Sample | Sampleb | X 5 | % Diffc
2010-10-14 | 52473 - Analytical | water | Cu | 20.5244 | 4.1565 | 20.7825 | 1.2574
2010-10-14 | 52473 - Analytical | water | Zn | 151.9668 | 30.5785 | 152.8923 | 0.60902
2010-10-14 | 52473 - Analytical | water | Cd | 5.8926 | 1.16 | 5.7998 | 1.5754
2010-10-14 | 52473 - Analytical | water | Pb | 19.1902 | 3.8052 | 19.0261 | 0.85521
2010-10-14 | 52480 - Analytical | water | Cu | 19.9908 | 3.9755 | 19.8776 | 0.56637
2010-10-14 | 52480 - Analytical | water | Zn | 213.6627 | 42.5329 | 212.6644 | 0.4672
2010-10-14 | 52480 - Analytical | water | Cd | 6.0322 | 1.1718 | 5.8591 | 2.8696
2010-10-14 | 52480 - Analytical | water | Pb | 19.6968 | 3.8673 | 19.3365 | 1.8294
2010-10-14 | 52515 - Analytical | water | Cu | 19.7123 | 3.9524 | 19.762 | 0.25225
2010-10-14 | 52515 - Analytical | water | Zn | 153.7896 | 30.9871 | 154.9357 | 0.7452
2010-10-14 | 52515 - Analytical | water | Cd | 7.6853 | 1.5323 | 7.6613 | 0.31132
2010-10-14 | 52515 - Analytical | water | Pb | 19.0773 | 3.819 | 19.0952 | 0.093692
aBID = Block Initiation Date: a date assigned to each member of a group
of samples that will identify the sample as a member of the
group or "block."
bdilution factor = 5 (1+4); digestates spiked with mid-range standard prior to analysis.
cdilution % difference acceptance criteria = +/- 10%; concentrations exceeding +/- 10%.
indicative of suspect interferent.

## Table 3-10 QC
Table 3-10. Recovery of elements from an interference checka
determined during ICP-MS analysis of sturgeon exposure water.
|  |  | Conc (ppb) | Conc (ppb) | Dilution | Mean
BID | Run Date | Element | actual | measured | Factor | % Rec.b
2010-07-22 | 2010-07-22 | Cu | 100 | 10.9437 | 10 | 109.4369
2010-07-22 | 2010-07-22 | Znc | 100 | 99.6877 | --- | 99.6877
2010-07-22 | 2010-07-22 | Cd | 50 | 5.0257 | 10 | 100.5144
2010-07-22 | 2010-07-22 | Pbc | 10 | 10.8654 | --- | 108.6543
2010-07-27 | 2010-07-27 | Cu | 100 | 11.1352 | 10 | 111.3519
2010-07-27 | 2010-07-27 | Znc | 100 | 102.6394 | --- | 102.6394
2010-07-27 | 2010-07-27 | Cd | 50 | 4.8515 | 10 | 97.0303
2010-07-27 | 2010-07-27 | Pbc | 10 | 10.6966 | --- | 106.966
2010-07-29 | 2010-07-29 | Cu | 100 | 10.9751 | 10 | 109.7507
2010-07-29 | 2010-07-29 | Znc | 100 | 101.0078 | --- | 101.0078
2010-07-29 | 2010-07-29 | Cd | 50 | 4.9649 | 10 | 99.2984
2010-08-03 | 2010-08-03 | Cu | 100 | 11.3207 | 10 | 113.2074
2010-08-03 | 2010-08-03 | Znc | 100 | 104.2771 | --- | 104.2771
2010-08-03 | 2010-08-03 | Cd | 50 | 5.1747 | 10 | 103.4948
2010-08-03 | 2010-08-03 | Pbc | 10 | 10.8313 | --- | 108.3128
2010-08-09 | 2010-08-09 | Cu | 100 | 10.7198 | 10 | 107.1981
2010-08-09 | 2010-08-09 | Znc | 100 | 98.4698 | --- | 98.4698
2010-08-09 | 2010-08-09 | Cd | 50 | 4.975 | 10 | 99.5002
2010-08-09 | 2010-08-09 | Pbc | 10 | 10.8636 | --- | 108.6356
2010-08-13 | 2010-08-13 | Cu | 100 | 11.5162 | 10 | 115.1616
2010-08-13 | 2010-08-13 | Znc | 100 | 105.7463 | --- | 105.7463
2010-08-13 | 2010-08-13 | Cd | 50 | 5.0187 | 10 | 100.3742
2010-08-13 | 2010-08-13 | Pbc | 10 | 10.8552 | --- | 108.5525
2010-08-17 | 2010-08-17 | Cu | 100 | 10.8774 | 10 | 108.7738
2010-08-17 | 2010-08-17 | Znc | 100 | 98.0544 | --- | 98.0544
2010-08-17 | 2010-08-17 | Cd | 50 | 4.7698 | 10 | 95.397
2010-08-17 | 2010-08-17 | Pbc | 10 | 10.4563 | --- | 104.5627
aHigh Purity ICP-MS Solution AB in 2% nitric acid, Charleston, SC.; CAT # ICP-MS-ICS.
bsuggested acceptance tolerance 80% - 120%.
cZn conc adjusted to 100ppb and Pb 10ppb in 10X diluted solution; Indigenous Zn
and Pb concentrations too low in 10X diluted AB solution for an accurate
interference check.

## Table 3-10b QC
Table 3-10 (continued). Recovery of elements in an interference checka
determined during ICP-MS analysis of sturgeon exposure water.
|  |  | Conc (ppb) | Conc (ppb) | Dilution | Mean
BID | Run Date | Element | actual | measured | Factor | % Rec.b
2010-08-18 | 2010-08-18 | Cu | 100 | 10.9677 | 10 | 109.6767
2010-08-18 | 2010-08-18 | Znc | 100 | 100.6327 | --- | 100.6327
2010-08-18 | 2010-08-18 | Cd | 50 | 4.8583 | 10 | 97.1669
2010-08-18 | 2010-08-18 | Pbc | 10 | 10.896 | --- | 108.9597
2010-08-26 | 2010-08-24 | Cu | 100 | 11.5333 | 10 | 115.3329
2010-08-26 | 2010-08-24 | Znc | 100 | 102.6838 | --- | 102.6838
2010-08-26 | 2010-08-24 | Cd | 50 | 4.8251 | 10 | 96.5029
2010-08-26 | 2010-08-24 | Pbc | 10 | 10.5889 | --- | 105.8887
2010-09-07 | 2010-09-07 | Cu | 100 | 10.9651 | 10 | 109.6514
2010-09-07 | 2010-09-07 | Znc | 100 | 100.2286 | --- | 100.2286
2010-09-07 | 2010-09-07 | Cd | 50 | 4.8322 | 10 | 96.6435
2010-09-07 | 2010-09-07 | Pbc | 10 | 10.7346 | --- | 107.3461
2010-09-10 | 2010-09-10 | Cu | 100 | 10.9907 | 10 | 109.9068
2010-09-10 | 2010-09-10 | Znc | 100 | 100.5438 | --- | 100.5438
2010-09-10 | 2010-09-10 | Cd | 50 | 4.9452 | 10 | 98.904
2010-09-10 | 2010-09-10 | Pbc | 10 | 10.9315 | --- | 109.3154
2010-09-29 | 2010-09-29 | Cu | 100 | 10.1882 | 10 | 101.8816
2010-09-29 | 2010-09-29 | Znc | 100 | 93.1678 | --- | 93.1678
2010-09-29 | 2010-09-29 | Cd | 50 | 4.7682 | 10 | 95.3647
2010-09-29 | 2010-09-29 | Pbc | 10 | 10.7829 | --- | 107.8292
2010-10-05 | 2010-10-06 | Cu | 100 | 9.6622 | 10 | 96.6217
2010-10-05 | 2010-10-06 | Znc | 100 | 89.5042 | --- | 89.5042
2010-10-05 | 2010-10-06 | Cd | 50 | 4.7026 | 10 | 94.0519
2010-10-05 | 2010-10-06 | Pbc | 10 | 10.7429 | --- | 107.4288
2010-10-14 | 2010-10-15 | Cu | 100 | 9.7855 | 10 | 97.8554
2010-10-14 | 2010-10-15 | Znc | 100 | 91.2844 | --- | 91.2844
2010-10-14 | 2010-10-15 | Cd | 50 | 4.776 | 10 | 95.52
2010-10-14 | 2010-10-15 | Pbc | 10 | 10.5279 | --- | 105.2792
aHigh Purity ICP-MS Solution AB in 2% nitric acid, Charleston, SC.; CAT # ICP-MS-ICS.
bsuggested acceptance tolerance 80% - 120%.
cZn conc adjusted to 100ppb and Pb 10ppb in 10X diluted solution; Indigenous Zn
and Pb concentrations too low in 10X diluted AB solution for an accurate
interference check.

## Table 3-11 QC
Table 3-11. Mean blank equivalent concentrations (BEC) of Cu, Zn, Cd, and Pb for reagent blanks analyzed with sturgeon exposure water.
Mean
|  |  | Blk 1 | Blk 2 | Blk 3 | Mean | BEC | BEC SD |  | Oper.
BIDa | Ele. | Matrix | Conc. | Conc. | Conc. | Conc.b | ng/mL | ng/mL | ISOP | Init.
2010-07-22 | Cu | water | 0.016291 | -0.009995 | 0.006121 | 0.004139 | 0.004139 | 0.013255 | P.241 | MJW/TWM
2010-07-22 | Zn | water | 0.75479 | 0.65994 | 0.62818 | 0.68097 | 0.68097 | 0.065877 | P.241 | MJW/TWM
2010-07-22 | Cd | water | -0.003334 | 0.00372 | 0.008729 | 0.0030383 | 0.0030383 | 0.0060603 | P.241 | MJW/TWM
2010-07-22 | Pb | water | -0.007296 | -0.014639 | -0.016754 | -0.012896 | -0.012896 | 0.004964 | P.241 | MJW/TWM
2010-07-27 | Cu | water | 0.064499 | 0.064225 | 0.022579 | 0.050434 | 0.050434 | 0.024124 | P.241 | MJW/TWM
2010-07-27 | Zn | water | 1.0071 | 0.71775 | 0.53655 | 0.75381 | 0.75381 | 0.23735 | P.241 | MJW/TWM
2010-07-27 | Cd | water | -0.003448 | 0.003186 | -0.0032 | -0.001154 | -0.001154 | 0.0037606 | P.241 | MJW/TWM
2010-07-27 | Pb | water | -0.016761 | -0.018947 | -0.022656 | -0.019455 | -0.019455 | 0.0029801 | P.241 | MJW/TWM
2010-07-29 | Cu | water | 0.052253 | 0.038349 | 0.01214 | 0.034247 | 0.034247 | 0.020369 | P.241 | MJW/TWM
2010-07-29 | Zn | water | 0.58063 | 1.4851 | 1.4105 | 1.1587 | 1.1587 | 0.50205 | P.241 | MJW/TWM
2010-07-29 | Cd | water | -0.00427 | 0.000863 | 0.002499 | -0.00030267 | -0.00030267 | 0.0035318 | P.241 | MJW/TWM
2010-08-03 | Cu | water | 0.021785 | 0.006663 | -0.029119 | -0.00022367 | -0.00022367 | 0.026141 | P.241 | MJW/TWM
2010-08-03 | Zn | water | 0.79618 | 0.80727 | 0.82843 | 0.81063 | 0.81063 | 0.016384 | P.241 | MJW/TWM
2010-08-03 | Cd | water | 0.005912 | 0.004703 | 0.004167 | 0.0049273 | 0.0049273 | 0.00089387 | P.241 | MJW/TWM
2010-08-03 | Pb | water | 0.016896 | 0.020581 | 0.0288 | 0.022092 | 0.022092 | 0.0060942 | P.241 | MJW/TWM
2010-08-09 | Cu | water | 0.036319 | 0.024198 | 0.01314 | 0.024552 | 0.024552 | 0.011594 | P.241 | MJW/TWM
2010-08-09 | Zn | water | -0.044209 | 0.12723 | 0.060448 | 0.047822 | 0.047822 | 0.086413 | P.241 | MJW/TWM
2010-08-09 | Cd | water | -0.005025 | -0.003475 | 0.000421 | -0.002693 | -0.002693 | 0.002806 | P.241 | MJW/TWM
2010-08-09 | Pb | water | 0.001023 | 0.015508 | -0.006568 | 0.003321 | 0.003321 | 0.011216 | P.241 | MJW/TWM
2010-08-13 | Cu | water | 0.044297 | 0.032356 | 0.007119 | 0.027924 | 0.027924 | 0.018981 | P.241 | MJW/TWM
2010-08-13 | Zn | water | 0.61688 | 0.42627 | 0.38544 | 0.4762 | 0.4762 | 0.12353 | P.241 | MJW/TWM
2010-08-13 | Cd | water | 0.008684 | 0.008971 | 0.005661 | 0.007772 | 0.007772 | 0.0018338 | P.241 | MJW/TWM
2010-08-13 | Pb | water | 0.002187 | -0.012535 | -0.015308 | -0.008552 | -0.008552 | 0.009403 | P.241 | MJW/TWM
2010-08-17 | Cu | water | 0.054009 | 0.04224 | 0.036735 | 0.044328 | 0.044328 | 0.0088243 | P.241 | MJW/TWM
2010-08-17 | Zn | water | 0.6109 | 1.1248 | 1.0327 | 0.92282 | 0.92282 | 0.27403 | P.241 | MJW/TWM
2010-08-17 | Cd | water | 0.004651 | 0.008095 | 0.010322 | 0.0076893 | 0.0076893 | 0.0028572 | P.241 | MJW/TWM
2010-08-17 | Pb | water | 0.004271 | 0.005054 | 0.00507 | 0.0047983 | 0.0047983 | 0.00045675 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bMean Conc. = the mean solution concentration of the procedural blanks for a block, n = 3 and individual reagent blanks; units ng/mL.

## Table 3-11b QC
Table 3-11 (continued). Mean blank equivalent concentrations (BEC) of Cu, Zn, Cd, and Pb for reagent blanks analyzed with sturgeon exposure water.
Mean
|  |  | Blk 1 | Blk 2 | Blk 3 | Mean | BEC | BEC SD |  | Oper.
BIDa | Ele. | Matrix | Conc. | Conc. | Conc. | Conc.b | ng/mL | ng/mL | ISOP | Init.
2010-08-18 | Cu | water | 0.038487 | 0.04972 | 0.045488 | 0.044565 | 0.044565 | 0.0056731 | P.241 | MJW/TWM
2010-08-18 | Zn | water | 0.89872 | 0.74267 | 0.78646 | 0.80929 | 0.80929 | 0.080487 | P.241 | MJW/TWM
2010-08-18 | Cd | water | -0.00257 | -0.002153 | -0.005376 | -0.0033663 | -0.0033663 | 0.0017529 | P.241 | MJW/TWM
2010-08-18 | Pb | water | 0.007423 | 0.013878 | 0.018355 | 0.013219 | 0.013219 | 0.0054957 | P.241 | MJW/TWM
2010-08-26 | Cu | water | 0.1147 | 0.085164 | 0.060983 | 0.086949 | 0.086949 | 0.026903 | P.241 | MJW/TWM
2010-08-26 | Zn | water | 0.97497 | 1.5198 | 1.7117 | 1.4022 | 1.4022 | 0.38221 | P.241 | MJW/TWM
2010-08-26 | Cd | water | 0.004511 | 0.019859 | 0.023612 | 0.015994 | 0.015994 | 0.01012 | P.241 | MJW/TWM
2010-08-26 | Pb | water | 0.004315 | 0.002642 | 0.006669 | 0.004542 | 0.004542 | 0.0020231 | P.241 | MJW/TWM
2010-09-07 | Cu | water | 0.059719 | 0.018416 | 0.038922 | 0.039019 | 0.039019 | 0.020652 | P.241 | MJW/TWM
2010-09-07 | Zn | water | 1.4946 | 1.9747 | 1.999 | 1.8228 | 1.8228 | 0.28444 | P.241 | MJW/TWM
2010-09-07 | Cd | water | 0.002371 | 0.024747 | 0.030785 | 0.019301 | 0.019301 | 0.014969 | P.241 | MJW/TWM
2010-09-07 | Pb | water | 0.017565 | 0.016945 | 0.019498 | 0.018003 | 0.018003 | 0.0013316 | P.241 | MJW/TWM
2010-09-10 | Cu | water | 0.02908 | 0.041719 | -0.002068 | 0.02291 | 0.02291 | 0.022536 | P.241 | MJW/TWM
2010-09-10 | Zn | water | 0.48432 | 0.6276 | 0.83785 | 0.64993 | 0.64993 | 0.17782 | P.241 | MJW/TWM
2010-09-10 | Cd | water | 0.010669 | 0.017553 | 0.022383 | 0.016868 | 0.016868 | 0.0058869 | P.241 | MJW/TWM
2010-09-10 | Pb | water | 0.004761 | 0.006753 | 0.004095 | 0.005203 | 0.005203 | 0.001383 | P.241 | MJW/TWM
2010-09-29 | Cu | water | 0.010852 | 0.047364 | -0.023478 | 0.011579 | 0.011579 | 0.035427 | P.241 | MJW/TWM
2010-09-29 | Zn | water | -0.074776 | -0.035213 | 0.072263 | -0.012575 | -0.012575 | 0.076089 | P.241 | MJW/TWM
2010-09-29 | Cd | water | 0.005649 | -0.011534 | -0.007184 | -0.0043563 | -0.0043563 | 0.0089337 | P.241 | MJW/TWM
2010-09-29 | Pb | water | -0.010274 | -0.003124 | -0.009225 | -0.007541 | -0.007541 | 0.003861 | P.241 | MJW/TWM
2010-10-05 | Cu | water | 0.031165 | 0.009466 | 1.5e-05 | 0.013549 | 0.013549 | 0.015971 | P.241 | MJW/TWM
2010-10-05 | Zn | water | 0.49442 | 0.35242 | 0.4546 | 0.43381 | 0.43381 | 0.073248 | P.241 | MJW/TWM
2010-10-05 | Cd | water | 0.001802 | 0.000857 | 0.007182 | 0.0032803 | 0.0032803 | 0.0034118 | P.241 | MJW/TWM
2010-10-05 | Pb | water | -0.001673 | -0.005234 | -0.006387 | -0.0044313 | -0.0044313 | 0.0024574 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bMean Conc. = the mean solution concentration of the procedural blanks for a block, n = 3 and individual reagent blanks; units ng/mL.

## Table 3-11c QC
Table 3-11 (continued). Mean blank equivalent concentrations (BEC) of Cu, Zn, Cd, and Pb for reagent blanks analyzed with sturgeon exposure water.
Mean
|  |  | Blk 1 | Blk 2 | Blk 3 | Mean | BEC | BEC SD |  | Oper.
BIDa | Ele. | Matrix | Conc. | Conc. | Conc. | Conc.b | ng/mL | ng/mL | ISOP | Init.
2010-10-14 | Cu | water | 0.038408 | 0.049399 | 0.016746 | 0.034851 | 0.034851 | 0.016615 | P.241 | MJW/TWM
2010-10-14 | Zn | water | -0.020535 | -0.012499 | -0.081797 | -0.038277 | -0.038277 | 0.037903 | P.241 | MJW/TWM
2010-10-14 | Cd | water | 0.005379 | -0.004406 | 0.004471 | 0.0018147 | 0.0018147 | 0.0054064 | P.241 | MJW/TWM
2010-10-14 | Pb | water | -0.001138 | 0.00182 | 0.00456 | 0.0017473 | 0.0017473 | 0.0028497 | P.241 | MJW/TWM
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bMean Conc. = the mean solution concentration of the procedural blanks for a block, n = 3 and individual reagent blanks; units ng/mL.

## Table 3-12 QC
Table 3-12. Method detection and quantitation limits for Cu, Zn, Cd, and Pb analyzed with sturgeon exposure water.
Standard
BIDa | Ele. | Matrix | W/D/Lb | Blk SD | SD | IDLc | MDLd | MQLe | PSOP | Prep. Init. | ISOP | Inst. Init. | Units
2010-07-22 | Cu | water | L | 0.013255 | 0.0016221 | 0.048 | 0.04 | 0.132 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-22 | Zn | water | L | 0.065877 | 0.047867 | 0.87 | 0.24 | 0.792 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-22 | Cd | water | L | 0.0060603 | 0.0053859 | 0.002 | 0.024 | 0.0792 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-22 | Pb | water | L | 0.004964 | 0.0067503 | 0.004 | 0.025 | 0.0825 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-27 | Cu | water | L | 0.024124 | 0.0086959 | 0.048 | 0.077 | 0.2541 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-27 | Zn | water | L | 0.23735 | 0.077524 | 0.87 | 0.75 | 2.475 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-27 | Cd | water | L | 0.0037606 | 0.001951 | 0.002 | 0.013 | 0.0429 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-27 | Pb | water | L | 0.0029801 | 0.0060023 | 0.004 | 0.02 | 0.066 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-29 | Cu | water | L | 0.020369 | 0.008608 | 0.048 | 0.067 | 0.2211 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-29 | Zn | water | L | 0.50205 | 0.036295 | 0.87 | 1.5 | 4.95 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-07-29 | Cd | water | L | 0.0035318 | 0.0028553 | 0.002 | 0.014 | 0.0462 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-03 | Cu | water | L | 0.026141 | 0.012782 | 0.048 | 0.087 | 0.2871 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-03 | Zn | water | L | 0.016384 | 0.12498 | 0.87 | 0.38 | 1.254 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-03 | Cd | water | L | 0.00089387 | 0.0028329 | 0.002 | 0.009 | 0.0297 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-03 | Pb | water | L | 0.0060942 | 0.0013776 | 0.004 | 0.019 | 0.0627 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-09 | Cu | water | L | 0.011594 | 0.0089585 | 0.048 | 0.044 | 0.1452 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-09 | Zn | water | L | 0.086413 | 0.028898 | 0.87 | 0.27 | 0.891 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-09 | Cd | water | L | 0.002806 | 0.0052798 | 0.002 | 0.018 | 0.0594 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-09 | Pb | water | L | 0.011216 | 0.0046224 | 0.004 | 0.036 | 0.1188 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-13 | Cu | water | L | 0.018981 | 0.0077129 | 0.048 | 0.061 | 0.2013 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-13 | Zn | water | L | 0.12353 | 0.054303 | 0.87 | 0.41 | 1.353 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-13 | Cd | water | L | 0.0018338 | 0.0022193 | 0.002 | 0.009 | 0.0297 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-13 | Pb | water | L | 0.009403 | 0.0044003 | 0.004 | 0.031 | 0.1023 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bW/D/L = state of starting sample: wet (W), dry (D), or liquid (L).
cIDL = instrument detection limit, unit ng/mL.
dMDL = method limit of detection, computed as 3 X (SDb2 + SDs2)1/2 where SDb = standard deviation of field blanks (n = 8) and
SDs = standard deviation of a low level standard diluted 100X (n = 3).
eMQL = 3.3 x MDL.

## Table 3-12b QC
Table 3-12 (continued). Method detection and quantitation limits for Cu, Zn, Cd, and Pb analyzed with sturgeon exposure water.
Standard
BIDa | Ele. | Matrix | W/D/Lb | Blk SD | SD | IDLc | MDLd | MQLe | PSOP | Prep. Init. | ISOP | Inst. Init. | Units
2010-08-17 | Cu | water | L | 0.0088243 | 0.006938 | 0.048 | 0.034 | 0.1122 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-17 | Zn | water | L | 0.27403 | 0.043462 | 0.87 | 0.83 | 2.739 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-17 | Cd | water | L | 0.0028572 | 0.0016456 | 0.002 | 0.01 | 0.033 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-17 | Pb | water | L | 0.00045675 | 0.0039384 | 0.004 | 0.012 | 0.0396 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-18 | Cu | water | L | 0.0056731 | 0.0026728 | 0.048 | 0.019 | 0.0627 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-18 | Zn | water | L | 0.080487 | 0.0028365 | 0.87 | 0.24 | 0.792 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-18 | Cd | water | L | 0.0017529 | 0.0054845 | 0.002 | 0.017 | 0.0561 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-18 | Pb | water | L | 0.0054957 | 0.008693 | 0.004 | 0.031 | 0.1023 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-26 | Cu | water | L | 0.026903 | 0.0077484 | 0.048 | 0.084 | 0.2772 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-26 | Zn | water | L | 0.38221 | 0.045954 | 0.87 | 1.2 | 3.96 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-26 | Cd | water | L | 0.01012 | 0.0040618 | 0.002 | 0.033 | 0.1089 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-08-26 | Pb | water | L | 0.0020231 | 0.0017881 | 0.004 | 0.008 | 0.0264 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-07 | Cu | water | L | 0.020652 | 0.0022247 | 0.048 | 0.062 | 0.2046 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-07 | Zn | water | L | 0.28444 | 0.0055325 | 0.87 | 0.85 | 2.805 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-07 | Cd | water | L | 0.014969 | 0.0021112 | 0.002 | 0.045 | 0.1485 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-07 | Pb | water | L | 0.0013316 | 0.0021543 | 0.004 | 0.008 | 0.0264 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-10 | Cu | water | L | 0.022536 | 0.0016982 | 0.048 | 0.068 | 0.2244 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-10 | Zn | water | L | 0.17782 | 0.051655 | 0.87 | 0.56 | 1.848 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-10 | Cd | water | L | 0.0058869 | 5.0239e-05 | 0.002 | 0.018 | 0.0594 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-10 | Pb | water | L | 0.001383 | 0.004171 | 0.004 | 0.013 | 0.0429 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-29 | Cu | water | L | 0.035427 | 0.0059311 | 0.048 | 0.11 | 0.363 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-29 | Zn | water | L | 0.076089 | 0.055771 | 0.87 | 0.28 | 0.924 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-29 | Cd | water | L | 0.0089337 | 0.0019508 | 0.002 | 0.027 | 0.0891 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-09-29 | Pb | water | L | 0.003861 | 0.0050839 | 0.004 | 0.019 | 0.0627 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bW/D/L = state of starting sample: wet (W), dry (D), or liquid (L).
cIDL = instrument detection limit, unit ng/mL.
dMDL = method limit of detection, computed as 3 X (SDb2 + SDs2)1/2 where SDb = standard deviation of field blanks (n = 8) and
SDs = standard deviation of a low level standard diluted 100X (n = 3).
eMQL = 3.3 x MDL.

## Table 3-12c QC
Table 3-12 (continued). Method detection and quantitation limits for Cu, Zn, Cd, and Pb analyzed with sturgeon exposure water.
Standard
BIDa | Ele. | Matrix | W/D/Lb | Blk SD | SD | IDLc | MDLd | MQLe | PSOP | Prep. Init. | ISOP | Inst. Init. | Units
2010-10-05 | Cu | water | L | 0.015971 | 0.0080675 | 0.048 | 0.054 | 0.1782 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-05 | Zn | water | L | 0.073248 | 0.040498 | 0.87 | 0.25 | 0.825 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-05 | Cd | water | L | 0.0034118 | 0.00019653 | 0.002 | 0.01 | 0.033 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-05 | Pb | water | L | 0.0024574 | 0.0055942 | 0.004 | 0.018 | 0.0594 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-14 | Cu | water | L | 0.016615 | 0.002272 | 0.048 | 0.05 | 0.165 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-14 | Zn | water | L | 0.037903 | 0.040567 | 0.87 | 0.17 | 0.561 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-14 | Cd | water | L | 0.0054064 | 0.001544 | 0.002 | 0.017 | 0.0561 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
2010-10-14 | Pb | water | L | 0.0028497 | 0.002073 | 0.004 | 0.011 | 0.0363 | P.214 | MJW | P.241 | MJW/TWM | ng/mL
aBID = Block Initiation Date: a date assigned to each member of a group of samples that will identify the sample as a member of the group or "block."
bW/D/L = state of starting sample: wet (W), dry (D), or liquid (L).
cIDL = instrument detection limit, unit ng/mL.
dMDL = method limit of detection, computed as 3 X (SDb2 + SDs2)1/2 where SDb = standard deviation of field blanks (n = 8) and
SDs = standard deviation of a low level standard diluted 100X (n = 3).
eMQL = 3.3 x MDL.