## T24_lims_tpc_9030
Table 24. Recovery of analytes in third-party check standards analyzed by direct aqueous-injection liquid chromatography-tandem mass spectrometry in positive electrospray ionization mode from September 2012 to March 2013 [Analytes in both 2033 and 2060 spike mixes are shown in bold. Recovery <70 or >130 percent or relative standard deviation (RSD) >30 percent are shown in bold, pink highlight. <, less than; >, greater than; Validation quality code: A, no qualification; E, estimated; na, not applicable; TPC, third-party check standard; USGS, U.S. Geological Survey; ng/L, nanograms per liter; N, number of samples; %, percent; Min., minimum; Max., maximum]
Analytical use | Validation quality code | TPC spike mix | Analyte | USGS parameter code | TPC spike amount (ng/L) | N | Mean (%) | RSD (%) | Min. (%) | Max. (%)
Analyte | A | 2033 | 2-Chloro-2,6-diethylacetanilide | 68525 | 250 | 11 | 96.1476 | 11.2012 | 86.0493 | 121.2943
Analyte | A | 2060 | 2-Hydroxyatrazine | 68660 | 250 | 11 | 90.8192 | 12.1414 | 77.0046 | 107.9296
Analyte | E | 2060 | 3-Hydroxycarbofuran | 68508 | 250 | 11 | 93.6568 | 10.5884 | 78.8099 | 109.5546
Analyte | A | 2033 | Acetochlor | 68520 | 250 | 11 | 84.6603 | 11.0825 | 70.1526 | 95.4015
Analyte | A | 2033 | Alachlor | 65064 | 250 | 11 | 100.7873 | 10.2292 | 79.0514 | 117.5814
Analyte | A | 2060 | Aldicarb | 68528 | 250 | 11 | 96.863 | 14.9279 | 79.5051 | 129.637
Analyte | A | 2060 | Aldicarb sulfone | 68529 | 250 | 11 | 104.7331 | 8.9079 | 87.4353 | 117.573
Analyte | A | 2060 | Aldicarb sulfoxide | 68530 | 250 | 11 | 67.7769 | 10.0613 | 55.8455 | 78.1466
Analyte | A | 2033 | Ametryn | 68533 | 250 | 11 | 89.2243 | 10.0184 | 80.7355 | 104.9643
Analyte | A | 2033 | Atrazine | 65065 | 250 | 11 | 98.7805 | 7.8742 | 85.0136 | 110.8524
Analyte | A | 2060 | Atrazine | 65065 | 250 | 11 | 99.6704 | 13.8456 | 86.763 | 127.8168
Analyte | A | 2033 | Azinphos-methyl | 65066 | 250 | 11 | 87.5469 | 14.929 | 73.7069 | 121.2723
Analyte | A | 2033 | Azinphos-methyl oxon | 68211 | 250 | 11 | 102.1006 | 12.1903 | 81.2136 | 123.0843
Analyte | A | 2033 | Bromacil | 68542 | 250 | 11 | 97.1341 | 12.6337 | 77.2847 | 121.2855
Analyte | A | 2060 | Bromacil | 68542 | 250 | 11 | 97.896 | 12.316 | 82.3214 | 120.0343
Analyte | A | 2033 | Butylate | 65068 | 250 | 11 | 84.5743 | 10.7873 | 71.8772 | 105.8233
Analyte | A | 2033 | Carbaryl | 65069 | 250 | 11 | 105.4466 | 9.3877 | 87.1084 | 124.346
Analyte | A | 2060 | Carbaryl | 65069 | 250 | 11 | 106.8324 | 14.5708 | 88.3325 | 134.4398
Analyte | E | 2060 | Carbendazim | 68548 | 250 | 11 | 111.224 | 32.8119 | 66.3183 | 150.1891
Analyte | A | 2033 | Carbofuran | 65070 | 250 | 11 | 109.6572 | 7.7648 | 99.8151 | 129.9666
Analyte | A | 2060 | Carbofuran | 65070 | 250 | 11 | 110.8301 | 11.3912 | 93.7175 | 133.6162
Analyte | A | 2060 | Chlorimuron-ethyl | 68872 | 250 | 11 | 86.7904 | 10.8266 | 74.502 | 105.8436
Analyte | A | 2033 | Chlorpyrifos | 65072 | 250 | 11 | 101.1466 | 10.857 | 85.356 | 115.1634
Analyte | A | 2033 | Chlorpyrifos oxon | 68216 | 250 | 11 | 112.2662 | 13.3261 | 84.5324 | 136.014
Analyte | A | 2033 | cis-Permethrin | 68769 | 250 | 11 | 79.4858 | 20.9504 | 49.5612 | 108.8263
Analyte | A | 2033 | Cyanazine | 66592 | 250 | 11 | 69.5549 | 23.7924 | 48.9274 | 102.0805
Analyte | A | 2033 | Deethylatrazine | 68552 | 250 | 11 | 113.4142 | 13.0194 | 87.9598 | 137.0035
Analyte | A | 2060 | Deethylatrazine | 68552 | 250 | 11 | 115.9515 | 13.1574 | 95.117 | 150.8822
Analyte | A | 2033 | Deisopropylatrazine | 68550 | 250 | 11 | 99.7277 | 14.2153 | 81.767 | 127.2862
Analyte | A | 2060 | Deisopropylatrazine | 68550 | 250 | 11 | 102.6242 | 12.0715 | 85.8825 | 129.6458
Analyte | A | 2033 | Diazinon | 65078 | 250 | 11 | 129.728 | 12.3783 | 113.2027 | 155.5861
Analyte | E | 2033 | Diazinon oxon | 68236 | 250 | 11 | 109.3839 | 18.4247 | 89.8173 | 154.614
Analyte | A | 2033 | Dichlorvos | 68572 | 250 | 11 | 150.9565 | 11.2958 | 129.3933 | 178.9133
Analyte | A | 2033 | Dicrotophos | 68573 | 250 | 11 | 131.0121 | 13.4297 | 111.3405 | 169.3671
Analyte | A | 2033 | Dimethoate | 66596 | 250 | 11 | 97.0744 | 14.8081 | 83.835 | 132.7653
Analyte | A | 2033 | Disulfoton | 67595 | 250 | 11 | 92.792 | 11.573 | 75.7989 | 109.1106
Analyte | A | 2033 | Disulfoton sulfone | 68589 | 250 | 11 | 101.5519 | 19.3485 | 72.9799 | 141.0063
Analyte | A | 2060 | Diuron | 66598 | 250 | 11 | 101.5193 | 8.5552 | 91.4678 | 119.9624
Analyte | A | 2033 | EPTC | 65080 | 250 | 11 | 88.6063 | 20.0456 | 57.9462 | 118.4836
Analyte | A | 2033 | Ethoprop | 68596 | 250 | 11 | 91.6714 | 13.6498 | 78.8118 | 119.0661
Analyte | A | 2033 | Fenamiphos | 68599 | 250 | 11 | 66.3738 | 12.0192 | 54.8461 | 80.0072
Analyte | A | 2033 | Fenamiphos sulfone | 68600 | 250 | 11 | 95.2404 | 15.0184 | 80.265 | 125.1304
Analyte | A | 2033 | Fenamiphos sulfoxide | 68601 | 250 | 11 | 94.6295 | 14.5268 | 77.043 | 127.1431
Analyte | E | 2060 | Flumetsulam | 61679 | 250 | 11 | 112.2585 | 13.6651 | 91.7023 | 133.048
Analyte | A | 2060 | Fluometuron | 68608 | 250 | 11 | 115.385 | 9.0735 | 99.3885 | 136.9964
Analyte | A | 2033 | Fonofos | 65084 | 250 | 11 | 99.2211 | 8.1792 | 84.5584 | 109.0756
Analyte | A | 2033 | Hexazinone | 65085 | 250 | 11 | 101.1904 | 8.2124 | 86.3888 | 114.5845
Analyte | A | 2060 | Imazaquin | 61682 | 250 | 11 | 108.0985 | 10.5881 | 89.9221 | 124.5497
Analyte | A | 2060 | Imazethapyr | 61683 | 250 | 11 | 126.9953 | 8.4222 | 115.1739 | 151.9489
Analyte | A | 2060 | Imidacloprid | 68426 | 250 | 11 | 97.4246 | 6.9919 | 87.2674 | 108.5233
Analyte | A | 2033 | Linuron | 68639 | 250 | 11 | 96.1743 | 9.8255 | 75.7078 | 108.3221
Analyte | A | 2060 | Linuron | 68639 | 250 | 11 | 97.2676 | 10.1673 | 83.4327 | 117.2075
Analyte | A | 2033 | Malaoxon | 68240 | 250 | 11 | 99.8839 | 14.763 | 74.4921 | 127.1693
Analyte | A | 2033 | Malathion | 65087 | 250 | 11 | 110.1765 | 10.3674 | 91.5678 | 132.7639
Analyte | A | 2033 | Metalaxyl | 68437 | 250 | 11 | 105.5652 | 9.1079 | 95.8766 | 128.2081
Analyte | A | 2060 | Metalaxyl | 68437 | 250 | 11 | 103.4294 | 9.0299 | 92.2139 | 117.9476
Analyte | A | 2033 | Methidathion | 65088 | 250 | 11 | 92.6727 | 15.2257 | 81.1879 | 128.4626
Analyte | A | 2060 | Methomyl | 68645 | 250 | 11 | 89.7471 | 13.4702 | 73.2506 | 114.7118
Analyte | A | 2033 | Metolachlor | 65090 | 250 | 11 | 89.0846 | 8.3713 | 76.0884 | 102.3416
Analyte | A | 2033 | Metribuzin | 68652 | 250 | 11 | 108.9454 | 17.6882 | 84.5657 | 158.1477
Analyte | A | 2033 | Molinate | 65091 | 250 | 11 | 127.5479 | 7.254 | 113.525 | 143.3268
Analyte | A | 2033 | Myclobutanil | 66632 | 250 | 11 | 99.7087 | 9.9744 | 80.1753 | 116.7966
Analyte | A | 2060 | Nicosulfuron | 61685 | 250 | 11 | 115.2455 | 12.9483 | 101.4796 | 152.1852
Analyte | A | 2060 | Norflurazon | 67685 | 250 | 11 | 107.4527 | 8.4006 | 93.2968 | 120.051
Analyte | A | 2060 | Oxamyl | 68664 | 250 | 11 | 96.2566 | 10.1238 | 74.676 | 106.6235
Analyte | E | 2033 | Oxyfluorfen | 65093 | 250 | 11 | 80.5015 | 65.1204 | 34.4581 | 210.3011
Analyte | A | 2033 | Paraoxon-methyl | 68648 | 250 | 11 | 101.1873 | 15.1905 | 80.1869 | 126.3773
Analyte | E | 2033 | Parathion-methyl | 65089 | 250 | 11 | 98.9285 | 41.8342 | 36.1608 | 159.8944
Analyte | A | 2033 | Pendimethalin | 65098 | 250 | 11 | 99.8593 | 10.0447 | 87.0451 | 113.2566
Analyte | A | 2033 | Phorate | 68668 | 250 | 11 | 108.8945 | 13.1635 | 87.9709 | 129.0277
Analyte | A | 2033 | Phorate oxon | 68669 | 250 | 11 | 108.3384 | 12.436 | 91.0115 | 135.8677
Analyte | E | 2033 | Phosmet | 65101 | 250 | 11 | 55.4254 | 51.6404 | 23.6206 | 123.2906
Analyte | A | 2033 | Prometon | 67702 | 250 | 11 | 102.0773 | 8.5401 | 87.157 | 118.9154
Analyte | A | 2033 | Prometryn | 65103 | 250 | 11 | 91.0824 | 7.6916 | 77.5412 | 99.9022
Analyte | A | 2033 | Pronamide | 67706 | 500 | 11 | 109.1415 | 11.5245 | 86.3683 | 125.2096
Analyte | A | 2033 | Propanil | 66641 | 250 | 11 | 115.9892 | 10.8639 | 94.6088 | 129.8997
Analyte | A | 2033 | Propargite | 68677 | 83 | 11 | 140.2649 | 13.0587 | 115.6113 | 170.9917
Analyte | A | 2033 | Propazine | 68678 | 250 | 11 | 99.4657 | 7.8009 | 88.2666 | 112.171
Analyte | A | 2033 | Propiconazole | 66643 | 250 | 11 | 90.9042 | 7.217 | 79.802 | 100.3544
Analyte | A | 2060 | Propiconazole | 66643 | 250 | 11 | 91.5433 | 9.9139 | 76.6754 | 108.2233
Analyte | A | 2060 | Propoxur | 68679 | 250 | 11 | 106.176 | 13.7897 | 88.7104 | 129.5345
Analyte | A | 2060 | Siduron | 68686 | 250 | 11 | 99.0191 | 8.1072 | 89.1617 | 119.9843
Analyte | A | 2033 | Simazine | 65105 | 250 | 11 | 100.8933 | 15.3261 | 75.4909 | 128.7736
Analyte | A | 2060 | Sulfometuron-methyl | 68688 | 250 | 11 | 102.4526 | 9.9169 | 89.596 | 127.4865
Analyte | A | 2033 | Tebuconazole | 66649 | 250 | 11 | 98.1015 | 7.2333 | 84.8229 | 107.4394
Analyte | A | 2033 | Tebuthiuron | 68695 | 250 | 11 | 97.4323 | 9.5553 | 81.4564 | 112.4017
Analyte | A | 2060 | Tebuthiuron | 68695 | 250 | 11 | 99.5062 | 9.4718 | 90.6254 | 119.7923
Analyte | A | 2033 | Terbacil | 68698 | 250 | 11 | 98.5315 | 13.4345 | 72.8677 | 122.6805
Analyte | A | 2060 | Terbacil | 68698 | 250 | 11 | 105.1257 | 15.1238 | 86.4943 | 140.4266
Analyte | A | 2033 | Terbufos | 68699 | 250 | 11 | 89.3373 | 13.2523 | 72.8471 | 106.5963
Analyte | A | 2033 | Terbufos oxon sulfone | 68701 | 250 | 11 | 124.2568 | 12.5683 | 96.8254 | 146.3644
Analyte | A | 2033 | Terbuthylazine | 66651 | 250 | 11 | 97.4671 | 9.4178 | 85.2304 | 114.7446
Analyte | A | 2033 | Thiobencarb | 65107 | 250 | 11 | 79.8867 | 8.2325 | 70.4242 | 92.7431
Analyte | A | 2033 | trans-Permethrin | 68708 | 250 | 11 | 73.3295 | 21.465 | 44.015 | 95.3574
Analyte | A | 2033 | Triallate | 68710 | 250 | 11 | 88.5214 | 7.1695 | 79.9005 | 95.7495
Analyte | A | 2033 | Tribufos | 68711 | 250 | 11 | 101.4114 | 17.5807 | 70.7232 | 125.9995
Surrogate | na | 2033 | Acetochlor-d11 (surrogate) | 90517 | 100 | 11 | 96.8925 | 5.414 | 89.151 | 107.0896
Surrogate | na | 2033 | Alachlor-d13 (surrogate) | 90518 | 100 | 11 | 97.323 | 8.5799 | 85.0547 | 109.845
Surrogate | na | 2033 | Carbaryl-d7 (surrogate) | 90519 | 100 | 11 | 99.3567 | 7.846 | 86.6927 | 113.456
Surrogate | na | 2033 | Carbendazim-d4 (surrogate) | 90520 | 100 | 11 | 94.1681 | 4.009 | 87.2684 | 99.1806
Surrogate | na | 2033 | Carbofuran-d3 (surrogate) | 90521 | 100 | 11 | 100.3433 | 4.1802 | 93.2632 | 106.2824
Surrogate | na | 2033 | cis-Permethrin-13C6 (surrogate) | 90558 | 100 | 11 | 81.1287 | 22.1522 | 54.9245 | 119.4837
Surrogate | na | 2033 | Deethylatrazine-d6 (surrogate) | 90522 | 100 | 11 | 102.5447 | 6.7544 | 94.2067 | 120.8376
Surrogate | na | 2033 | Diazinon-d10 (surrogate) | 90523 | 100 | 11 | 96.5364 | 5.3467 | 86.615 | 105.5202
Surrogate | na | 2033 | Diflubenzuron-d4 (surrogate) | 90524 | 100 | 11 | 95.7648 | 4.6659 | 89.8964 | 103.1116
Surrogate | na | 2033 | Diuron-d6 (surrogate) | 90808 | 100 | 11 | 96.3193 | 4.5348 | 89.5722 | 104.501
Surrogate | na | 2033 | Hexazinone-d6 (surrogate) | 90527 | 100 | 11 | 99.7851 | 5.4815 | 93.2425 | 111.1229
Surrogate | na | 2033 | Linuron-d6 (surrogate) | 90529 | 100 | 11 | 97.8555 | 3.9812 | 92.1179 | 103.3415
Surrogate | na | 2033 | Malathion-d10 (surrogate) | 90552 | 100 | 11 | 90.3221 | 7.4755 | 74.3704 | 98.1577
Surrogate | na | 2033 | Metolachlor-d6 (surrogate) | 90553 | 100 | 11 | 99.4856 | 4.2918 | 93.7972 | 109.694
Surrogate | na | 2033 | Nicosulfuron-d6 (surrogate) | L903L | 100 | 11 | 97.0973 | 4.3771 | 91.7089 | 104.525
Surrogate | na | 2033 | Tebuconazole-d6 (surrogate) | 90555 | 100 | 11 | na | na | na | na
Surrogate | na | 2033 | Thiobencarb-d10 (surrogate) | 90556 | 100 | 11 | 96.158 | 5.9954 | 86.2139 | 102.5376
Surrogate | na | 2060 | Acetochlor-d11 (surrogate) | 90517 | 100 | 11 | 99.6518 | 9.4801 | 86.9614 | 119.8657
Surrogate | na | 2060 | Alachlor-d13 (surrogate) | 90518 | 100 | 11 | 103.0673 | 8.9783 | 91.6877 | 124.4445
Surrogate | na | 2060 | Carbaryl-d7 (surrogate) | 90519 | 100 | 11 | 102.0421 | 8.8411 | 81.4532 | 112.5498
Surrogate | na | 2060 | Carbendazim-d4 (surrogate) | 90520 | 100 | 11 | 99.9092 | 5.5662 | 86.7447 | 107.3605
Surrogate | na | 2060 | Carbofuran-d3 (surrogate) | 90521 | 100 | 11 | 103.6158 | 6.4213 | 93.3891 | 112.6098
Surrogate | na | 2060 | cis-Permethrin-13C6 (surrogate) | 90558 | 100 | 11 | 83.8217 | 20.8156 | 60.7937 | 116.9619
Surrogate | na | 2060 | Deethylatrazine-d6 (surrogate) | 90522 | 100 | 11 | 102.7963 | 9.8803 | 87.1118 | 121.7666
Surrogate | na | 2060 | Diazinon-d10 (surrogate) | 90523 | 100 | 11 | 99.9062 | 9.5925 | 89.5522 | 118.4311
Surrogate | na | 2060 | Diflubenzuron-d4 (surrogate) | 90524 | 100 | 11 | 96.6785 | 7.6321 | 87.5481 | 111.4994
Surrogate | na | 2060 | Diuron-d6 (surrogate) | 90808 | 100 | 11 | 99.4936 | 4.6237 | 92.3065 | 105.9258
Surrogate | na | 2060 | Hexazinone-d6 (surrogate) | 90527 | 100 | 11 | 101.1866 | 6.5675 | 93.32 | 115.4226
Surrogate | na | 2060 | Linuron-d6 (surrogate) | 90529 | 100 | 11 | 101.3747 | 6.3424 | 94.653 | 113.8794
Surrogate | na | 2060 | Malathion-d10 (surrogate) | 90552 | 100 | 11 | 97.6981 | 10.1355 | 86.0089 | 115.6385
Surrogate | na | 2060 | Metolachlor-d6 (surrogate) | 90553 | 100 | 11 | 101.7893 | 6.0845 | 93.1734 | 110.371
Surrogate | na | 2060 | Nicosulfuron-d6 (surrogate) | L903L | 100 | 11 | 98.7227 | 7.7167 | 88.9758 | 111.3665
Surrogate | na | 2060 | Tebuconazole-d6 (surrogate) | 90555 | 100 | 11 | na | na | na | na
Surrogate | na | 2060 | Thiobencarb-d10 (surrogate) | 90556 | 100 | 11 | 100.6236 | 5.968 | 92.6032 | 112.9101
Note: the TPC spike amount of carbendazim is the nominal concnetration of benomyl in the schedule 2060 TPC mixture; benomyl rapidly degrades to carbendzim in water.
Note: the nominal concentration of propargite for this lot of the schedule 2033 standard provided by the vendor did not match previous lots
|  |  | Analytes in schedule 2033 | 65
|  |  | Analytes in schedule 2060 | 31
|  |  | Qualified E analytes in 2033 | 4
|  |  | Qualified E analytes in 2060 | 3
Legend:
|  |  | Recovery <70 or >130 percent |  |  |  | 60
|  |  | Recovery <80 or >120 percent |  |  |  | 126.9953
|  |  | RSD > 30 pecent |  |  |  | 32.8119

## T25_lims_tpc_9031
Table 25. Recovery of analytes in third-party check standards analyzed by direct aqueous-injection liquid chromatography-tandem mass spectrometry in negative electrospray ionization mode from September 2012 to March 2013 [Positive electrospray ionization (ESI+) method is the preferred analysis mode. Analytes in both 2033 and 2060 spike mixes are shown in bold. Recovery <70 or >130 percent or relative standard deviation (RSD) >30 percent are shown in bold, pink highlight. <, less than; >, greater than; Validation quality code: A, no qualification; E, estimated; na, not applicable; TPC, third-party check standard; USGS, U.S. Geological Survey; ng/L, nanograms per liter; N, number of samples; %, percent; Min., minimum; Max., maximum]
Analytical use | Validation quality code | TPC spike mix | Analyte | USGS parameter code | TPC spike amount (ng/L) | N | Mean (%) | RSD (%) | Min. (%) | Max. (%)
Analyte | A | 2033 | Desulfinylfipronil | 66607 | 250 | 11 | 93.4634 | 11.2708 | 73.0945 | 106.6496
Analyte | A | 2033 | Desulfinylfipronil amide | 68570 | 250 | 11 | 97.6985 | 8.4563 | 82.2999 | 106.5726
Analyte | A | 2033 | Fipronil | 66604 | 250 | 11 | 89.1912 | 14.0173 | 70.3436 | 104.1272
Analyte | A | 2033 | Fipronil sulfide | 66610 | 500 | 11 | 91.8408 | 16.7743 | 62.1353 | 106.7154
Analyte | A | 2033 | Fipronil sulfone | 66613 | 500 | 11 | 84.4258 | 20.9381 | 51.743 | 103.0284
Analyte | A | 2060 | 2,4-D | 68500 | 250 | 11 | 90.1743 | 9.9337 | 78.8753 | 104.9597
Analyte | A | 2060 | Bentazon | 68538 | 250 | 11 | 101.956 | 11.3032 | 83.9696 | 119.7688
Analyte | A | 2060 | Bromoxynil | 68543 | 250 | 11 | 101.2019 | 11.1884 | 86.0101 | 121.4365
Analyte | A | 2060 | MCPA | 68641 | 250 | 11 | 87.7535 | 21.5567 | 53.7988 | 114.2866
Analyte | A | 2060 | Oryzalin | 68663 | 250 | 11 | 95.6604 | 6.3222 | 85.2861 | 103.8482
Analyte | A | 2060 | Triclopyr | 68712 | 250 | 11 | 86.9202 | 15.9735 | 66.9113 | 119.3441
Analyte | E | 2060 | Chlorthal-monomethyl | 68560 | 250 | 11 | 102.8088 | 35.1941 | 65.6018 | 169.9685
Analyte | E | 2060 | Dicamba | 68571 | 250 | 11 | 95.2577 | 27.928 | 51.8094 | 157.3801
Analyte | 9030 | 2033 | Propanil | L903F | 250 | 11 | 103.8543 | 6.1542 | 94.0513 | 114.5938
Analyte | 9030 | 2060 | Chlorimuron-ethyl | L903C | 250 | 11 | 98.5701 | 27.3197 | 59.1274 | 134.1061
Surrogate | na | 2033 | 2,4-D-d3 (surrogate) | 91986 | 100 | 11 | 88.989 | 10.3937 | 72.9206 | 102.1843
Surrogate | na | 2033 | 3-Phenoxybenzoic acid-13C6 (surrogate) | 90516 | 100 | 11 | 102.142 | 4.1735 | 97.0561 | 109.6477
Surrogate | na | 2033 | Butachlor SA (surrogate) | 90624 | 100 | 11 | 101.8296 | 6.4151 | 90.257 | 113.1554
Surrogate | na | 2033 | Dimethachlor SA (surrogate) | 90625 | 100 | 11 | 102.393 | 11.689 | 86.431 | 118.6181
Surrogate | na | 2033 | Diuron-d6 (surrogate) | L903M | 100 | 11 | 99.4784 | 7.6347 | 90.0724 | 112.1842
Surrogate | na | 2033 | Linuron-d6 (surrogate) | L903K | 100 | 11 | 103.8719 | 9.4598 | 86.7254 | 124.2793
Surrogate | na | 2033 | Nicosulfuron-d6 (surrogate) | 90554 | 100 | 11 | 92.7867 | 12.4086 | 76.4964 | 116.0285
Surrogate | na | 2060 | 2,4-D-d3 (surrogate) | 91986 | 100 | 11 | 90.5758 | 11.1657 | 75.3516 | 106.0822
Surrogate | na | 2060 | 3-Phenoxybenzoic acid-13C6 (surrogate) | 90516 | 100 | 11 | 101.9897 | 3.7491 | 97.1543 | 109.3727
Surrogate | na | 2060 | Butachlor SA (surrogate) | 90624 | 100 | 11 | 101.8271 | 4.3236 | 93.6953 | 107.5299
Surrogate | na | 2060 | Dimethachlor SA (surrogate) | 90625 | 100 | 11 | 103.1312 | 9.1588 | 91.5349 | 122.9509
Surrogate | na | 2060 | Diuron-d6 (surrogate) | L903M | 100 | 11 | 100.0551 | 11.193 | 85.9684 | 119.3931
Surrogate | na | 2060 | Linuron-d6 (surrogate) | L903K | 100 | 11 | 101.2161 | 8.8945 | 88.0983 | 118.4639
Surrogate | na | 2060 | Nicosulfuron-d6 (surrogate) | 90554 | 100 | 11 | 94.0811 | 22.2585 | 66.5178 | 129.4039
|  |  | Analytes in schedule 2033 | 5
|  |  | Analytes in schedule 2060 | 6
Qualified E analytes in 2033
|  |  | Qualified E analytes in 2033 | 2

## T26_lims_2437_pblnk
Table 26. Occurrence and concentrations of pesticide compounds in laboratory reagent-blank samples analyzed by direct aqueous-injection liquid chromatography-tandem mass spectrometry between June 19, 2012, and March 3, 2013 [Pesticide compounds with detection frequencies greater than 15 percent are highlighted in bold. Validation quality code: A, acceptable; E, estimated; na, not applicable; D, deleted; 9030, moved to positive electrospray ionization (ESI+) method; 9031,moved to ESI- method. Lab code 9030, ESI+ mode; lab code 9031, ESI-mode; USGS, U.S. Geological Survey; LRL, laboratory reporting level; ng/L, nanograms per liter; N, number of samples; Max., maximum]
Validation quality code | Analytical use | Lab code | Pesticide compound | USGS parameter code | LRL (ng/L) | N | Median (ng/L) | 95th percentile (ng/L) | Max. (ng/L) | Number of detections | Detection frequency (percent)
A | Analyte | 9030 | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Amino-N-isopropylbenzamide | 68503 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Aminobenzimidazole | 68502 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Chloro-2,6-diethylacetanilide | 68525 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 2-Hydroxyatrazine | 68660 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 3,4-Dichlorophenylurea | 68226 | 144 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 4-Chlorobenzylmethyl sulfoxide | 68514 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 4-Hydroxy molinate | 68515 | 7 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | 4-Hydroxyhexazinone A | 68517 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Acephate | 68519 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Acetochlor | 68520 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Alachlor | 65064 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Aldicarb | 68528 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Aldicarb sulfone | 68529 | 20 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Aldicarb sulfoxide | 68530 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Ametryn | 68533 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Atrazine | 65065 | 5 | 17 | 0 | 1.7877 | 1.7877 | 1 | 5.8824
A | Analyte | 9030 | Azinphos-methyl | 65066 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Azinphos-methyl oxon | 68211 | 15 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Azoxystrobin | 66589 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Bromacil | 68542 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Butylate | 65068 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Carbaryl | 65069 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Carbofuran | 65070 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Carboxy molinate | 68549 | 22 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Chlorimuron-ethyl | 68872 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Chlorpyrifos | 65072 | 4 | 17 | 0 | 1.0999 | 1.0999 | 1 | 5.8824
A | Analyte | 9030 | Chlorpyrifos oxon | 68216 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Chlorsulfuron | 61678 | 50 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | cis-Permethrin | 68769 | 2 | 17 | 0 | 0.3892 | 0.3892 | 1 | 5.8824
A | Analyte | 9030 | Cyanazine | 66592 | 16 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dechlorometolachlor | 68562 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Deethylatrazine (CIAT) | 68552 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Desiodo flubendiamide | 68563 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Deisopropyl prometryn | 68564 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Deisopropylatrazine (CEAT) | 68550 | 20 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Demethyl fluometuron | 68591 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Demethyl hexazinone B | 68566 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Demethyl norflurazon | 68567 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Desamino metribuzin | 68568 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Diazinon | 65078 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dichlorvos | 68572 | 52 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dicrotophos | 68573 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dideaklatrazine (CAAT) | 68547 | 24 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Didemethyl hexazinone F (Hexazione TP F) | 68574 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Diflubenzuron | 68576 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dimethenamid | 68580 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Dimethoate | 66596 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton | 67595 | 13 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton oxon | 68586 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton oxon sulfone | 68588 | 7 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton oxon sulfoxide | 68587 | 7 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton sulfone | 68589 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Disulfoton sulfoxide | 68590 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Diuron | 66598 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | EPTC | 65080 | 206 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | EPTC R248722 | 68594 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Ethoprop | 68596 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Etoxazole | 68598 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fenamiphos | 68599 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fenamiphos sulfone | 68600 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fenamiphos sulfoxide | 68601 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fentin | 68603 | 16 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fluometuron | 68608 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Fonofos | 65084 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Halosulfuron-methyl | 61680 | 22 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hexazinone | 65085 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hexazinone TP C | 68612 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hexazinone TP G | 68713 | 22 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxy mono demethyl fluometuron | 68617 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxyacetochlor | 68615 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxyalachlor | 68616 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxydiazinon | 68618 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxydidemethyl fluometuron | 68619 | 20 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxyfluometuron | 68620 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxymetolachlor | 68622 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxyphthalazinone | 68623 | 46 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxysimazine | 68624 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Hydroxytebuthurion | 68621 | 11 | 16 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Imazamox | 68625 | 28 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Imazaquin | 61682 | 19 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Imazethapyr | 61683 | 20 | 12 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Imidacloprid | 68426 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Indoxacarb | 68627 | 8 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Isoxaflutole | 68632 | 13 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Kresoxim-methyl | 67670 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Lactofen | 68638 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Linuron | 68639 | 6 | 17 | 0 | 1.492 | 1.492 | 1 | 5.8824
A | Analyte | 9030 | Malaoxon | 68240 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Malathion | 65087 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Metalaxyl | 68437 | 6 | 17 | 0 | 0.8838 | 0.8838 | 1 | 5.8824
A | Analyte | 9030 | Metconazole | 66620 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Methamidophos | 68644 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Methidathion | 65088 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Methomyl | 68645 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Methomyl oxime | 68646 | 500 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Methoxyfenozide | 68647 | 2 | 17 | 0 | 1.1138 | 1.1138 | 1 | 5.8824
A | Analyte | 9030 | Methyl paraoxon | 68648 | 19 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Metolachlor | 65090 | 9 | 17 | 0 | 5.744 | 5.744 | 2 | 11.7647
A | Analyte | 9030 | Metolachlor hydroxy morpholinone | 68649 | 7 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Metribuzin | 68652 | 20 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Molinate | 65091 | 18 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Myclobutanil | 66632 | 7 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Nicosulfuron | 61685 | 14 | 17 | 0 | 12.1881 | 12.1881 | 2 | 11.7647
A | Analyte | 9030 | Norflurazon | 67685 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | O-Ethyl-O-methyl-S-propylphosphorothioate | 68597 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | O-Ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | O-Ethyl-S-propyl phosphorothioate | 68658 | 64 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Omethoate | 68661 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Orthosulfamuron | 68662 | 6 | 17 | 0 | 1.6171 | 1.6171 | 1 | 5.8824
A | Analyte | 9030 | Oxamyl | 68664 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Oxamyl oxime | 68665 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Paraoxon | 68666 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Pendimethalin | 65098 | 3 | 17 | 0 | 3.2873 | 3.2873 | 2 | 11.7647
A | Analyte | 9030 | Phorate | 68668 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Phorate oxon | 68669 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Phorate oxon sulfone | 68670 | 20 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Phorate oxon sulfoxide | 68671 | 7 | 17 | 0 | 1.7234 | 1.7234 | 1 | 5.8824
A | Analyte | 9030 | Phorate sulfone | 68672 | 9 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Phorate sulfoxide | 68673 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Phthalazinone | 68675 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Piperonyl butoxide | 65102 | 4 | 17 | 0 | 1.5054 | 1.5054 | 3 | 17.6471
A | Analyte | 9030 | Profenofos | 68676 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Prometon | 67702 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Prometryn | 65103 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propanil | 66641 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propargite | 68677 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propazine | 68678 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propiconazole | 66643 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propoxur | 68679 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Propyzamide (Pronamide) | 67706 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Prosulfuron | 61687 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Pymetrozine | 68680 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Pyraclostrobin | 66646 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Pyridaben | 68682 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Pyrimidinol | 68505 | 20 | 17 | 0 | 6.4662 | 6.4662 | 6 | 35.2941
A | Analyte | 9030 | Pyriproxyfen | 68683 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Siduron | 68686 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Simazine | 65105 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Sulfometuron-methyl | 68688 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Sulfosulfuron | 68689 | 9 | 17 | 0 | 0.8619 | 0.8619 | 1 | 5.8824
A | Analyte | 9030 | Sulfosulfuron ethyl sulfone | 68690 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuconazole | 66649 | 5 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebufenozide | 68692 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebupirimfos | 68693 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebupirimfos oxon | 68694 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuthiuron | 68695 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuthiuron TP 104 | 68575 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuthiuron TP 106 | 68714 | 76 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuthiuron TP 108 | 68696 | 10 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tebuthiuron TP 109 (OH) | 68697 | 17 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbacil | 68698 | 21 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos | 68699 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos oxon | 68700 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos oxon sulfone | 68701 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos oxon sulfoxide | 68702 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos sulfone | 68703 | 11 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbufos sulfoxide | 68704 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Terbuthylazine | 66651 | 3 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tetraconazole | 66654 | 6 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Thiobencarb | 65107 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | trans-Permethrin | 68708 | 4 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Triallate | 68710 | 16 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Tribufos | 68711 | 2 | 17 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9030 | Trifloxystrobin | 66660 | 2 | 17 | 0 | 0.0827 | 0.0827 | 1 | 5.8824
A | Analyte | 9031 | 2,4-D | 68500 | 62 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | 3-Phenoxybenzoic acid | 68873 | 61 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | 4-(Hydroxymethyl)pendimethalin | 68511 | 213 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | 4-Hydroxychlorothalonil | 68336 | 98 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Acetochlor OA | 68522 | 90 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Acetochlor SA | 68523 | 100 | 10 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Acetochlor SAA | 68524 | 176 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Alachlor OA | 68526 | 84 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Alachlor SA | 68871 | 200 | 10 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Alachlor SAA | 68527 | 169 | 18 | 0 | 4.945 | 4.945 | 1 | 5.5556
A | Analyte | 9031 | Bentazon | 68538 | 10 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Bromoxynil | 68543 | 79 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | cis-Cyhalothric acid | 68553 | 1000 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Dechlorofipronil | 68561 | 4 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Desamino-diketo Metribuzin | 68569 | 200 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Desulfinylfipronil | 66607 | 4 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Desulfinylfipronil amide | 68570 | 9 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Diflufenzopyr | 68577 | 88 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Diketonitrile-isoxaflutole | 68578 | 62 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Dimethenamid OA | 68581 | 85 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Dimethenamid SA | 68582 | 79 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Dimethenamid SAA | 68583 | 189 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Fipronil | 66604 | 4 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Fipronil amide | 68604 | 9 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Fipronil sulfide | 66610 | 4 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Fipronil sulfonate | 68605 | 96 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Fipronil sulfone | 66613 | 4 | 18 | 0 | 0.94783 | 0.94783 | 6 | 33.3333
A | Analyte | 9031 | Flubendiamide | 68606 | 5 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Hexazinone TP E | 68614 | 76 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Isoxaflutole Acid RPA 203328 | 68633 | 9 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | MCPA | 68641 | 95 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Metolachlor OA | 68650 | 149 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Metolachlor SA | 68651 | 68 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Metribuzin DK | 68653 | 200 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Oryzalin | 68663 | 11 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | sec-Acetochlor OA | 68684 | 52 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | sec-Alachlor OA | 68685 | 135 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Sulfentrazone | 68687 | 11 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | TCPSA (Triallate-SA) | 68691 | 54 | 18 | 0 | 0 | 0 | 0 | 0
A | Analyte | 9031 | Triclopyr | 68712 | 88 | 18 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | 1H-1,2,4-Triazole | 68498 | 22 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | 2-(1-Hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 94 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | 3-Hydroxycarbofuran | 68508 | 16 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | 7-Hydroxycarbofuran | 68518 | 500 | 11 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Asulam | 68536 | 28 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Bifenthrin | 65067 | 19 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Butralin | 68545 | 5 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Carbendazim | 68548 | 3 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Diazinon oxon | 68236 | 4 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Famoxadone | 67609 | 60 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Fenbutatin oxide | 68602 | 20 | 17 | 0 | 38.0215 | 38.0215 | 6 | 35.2941
E | Analyte | 9030 | Flumetsulam | 61679 | 17 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Methyl parathion | 65089 | 500 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Naled | 68654 | 30 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Novaluron | 68655 | 9 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Oxyfluorfen | 65093 | 500 | 16 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9030 | Phosmet | 65101 | 20 | 17 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9031 | Chlorosulfonamide acid | 68551 | 75 | 18 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9031 | DCPA monoacid (Chlorthal-monomethyl) | 68560 | 500 | 18 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9031 | Dicamba | 68571 | 500 | 18 | 0 | 0 | 0 | 0 | 0
E | Analyte | 9031 | Hexazinone TP D | 68613 | 200 | 18 | 0 | 0 | 0 | 0 | 0
na | Surrogate | 9030 | Acetochlor-d11 | 90517 | na | 17 | 95.5726 | 105.3921 | 105.3921 | 17 | 100
na | Surrogate | 9030 | Alachlor-d13 | 90518 | na | 17 | 98.3942 | 108.6993 | 108.6993 | 17 | 100
na | Surrogate | 9030 | Carbaryl-d7 | 90519 | na | 17 | 96.1936 | 108.6916 | 108.6916 | 17 | 100
na | Surrogate | 9030 | Carbendazim-d4 | 90520 | na | 17 | 96.6281 | 117.5756 | 117.5756 | 17 | 100
na | Surrogate | 9030 | Carbofuran-d3 | 90521 | na | 17 | 98.8857 | 111.6164 | 111.6164 | 17 | 100
na | Surrogate | 9030 | cis-Permethrin-13C6 | 90558 | na | 17 | 95.2905 | 155.5265 | 155.5265 | 17 | 100
na | Surrogate | 9030 | Deethylatrazine-d6 | 90522 | na | 17 | 98.8274 | 116.1419 | 116.1419 | 17 | 100
na | Surrogate | 9030 | Diazinon-d10 (diethyl-d10) | 90523 | na | 17 | 93.5905 | 105.1118 | 105.1118 | 17 | 100
na | Surrogate | 9030 | Diflubenzuron-d4 (4-chlorophenyl-d4) | 90524 | na | 17 | 95.4652 | 113.78 | 113.78 | 17 | 100
na | Surrogate | 9030 | Diuron-d6 | 90808 | na | 17 | 97.2581 | 111.173 | 111.173 | 17 | 100
na | Surrogate | 9030 | Hexazinone-d6 (N,N-dimethyl-d6) | 90527 | na | 17 | 96.1515 | 112.4299 | 112.4299 | 17 | 100
na | Surrogate | 9030 | Linuron-d6 (dimethyl-d6) | 90529 | na | 17 | 97.894 | 108.2713 | 108.2713 | 17 | 100
na | Surrogate | 9030 | Malathion-d10 (diethyl-d10) | 90552 | na | 17 | 90.3599 | 106.3535 | 106.3535 | 17 | 100
na | Surrogate | 9030 | Metolachlor-d6 | 90553 | na | 17 | 97.5961 | 104.023 | 104.023 | 17 | 100
na | Surrogate | 9030 | Nicosulfuron-d6 (dimethoxy-d6) | L903L | na | 17 | 96.0814 | 108.711 | 108.711 | 17 | 100
na | Surrogate | 9030 | Tebuconazole-d6 | 90555 | na | 1 | 194.5389 | 194.5389 | 194.5389 | 1 | 100
na | Surrogate | 9030 | Thiobencarb-d10 (diethyl-d10) | 90556 | na | 17 | 93.7002 | 101.6804 | 101.6804 | 17 | 100
na | Surrogate | 9031 | 2,4-D-d3 | 91986 | na | 18 | 90.8908 | 137.4597 | 137.4597 | 18 | 100
na | Surrogate | 9031 | 3-Phenoxybenzoic acid-13C6 | 90516 | na | 18 | 102.0559 | 117.5298 | 117.5298 | 18 | 100
na | Surrogate | 9031 | Butachlor SA, sur | 90624 | na | 18 | 105.3445 | 119.9464 | 119.9464 | 18 | 100
na | Surrogate | 9031 | Dimethachlor SA, sur | 90625 | na | 18 | 104.9344 | 120.7413 | 120.7413 | 18 | 100
na | Surrogate | 9031 | Diuron-d6 | L903M | na | 18 | 101.6936 | 137.3847 | 137.3847 | 18 | 100
na | Surrogate | 9031 | Linuron-d6 (dimethyl-d6) | L903K | na | 18 | 101.0307 | 128.5335 | 128.5335 | 18 | 100
na | Surrogate | 9031 | Nicosulfuron-d6 (dimethoxy-d6) | 90554 | na | 18 | 96.9891 | 128.0268 | 128.0268 | 18 | 100
9030 | Analyte | 9031 | Chlorimuron-ethyl | L903C | 1000 | 18 | 0 | 0 | 0 | 0 | 0
9030 | Analyte | 9031 | Chlorsulfuron | L903R | 1000 | 13 | 0 | 0 | 0 | 0 | 0
9030 | Analyte | 9031 | Novaluron | L903J | 1000 | 17 | 0 | 0 | 0 | 0 | 0
9030 | Analyte | 9031 | Propanil | L903F | 1000 | 18 | 0 | 0 | 0 | 0 | 0
9030 | Analyte | 9031 | Prosulfuron | 61687 | 8 | 18 | 0 | 0 | 0 | 0 | 0
9031 | Analyte | 9030 | Dimethenamid SAA | L903G | 1000 | 17 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | 3-Ketocarbofuran | 68509 | 20000 | 12 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | Aminopyralid | 68534 | 20000 | 16 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | Bifenazate | 68540 | 20000 | 16 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | Bifenazate diazene | 68541 | 20000 | 17 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | flumiclorac-pentyl | 68607 | 20000 | 13 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | Formetanate | 68609 | 20000 | 16 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9030 | Phosmet oxon | 68674 | 20000 | 5 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9031 | Ammelide | 68535 | 20000 | 18 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9031 | Deisopropyliprodione | 68565 | 20000 | 18 | 0 | 0 | 0 | 0 | 0
D | Analyte | 9031 | Iprodione | 66617 | 20000 | 1 | 0 | 0 | 0 | 0 | 0
Note: Blank results for ESI+ method blank sample 20121708011 analyzed June 19, 2012 were excluded based on analysts comments that blank was run first without any pre-wash sample (injection number 1 in table 8).
[5 empty] | 211 |  |  |  | Count of 9030 detections: | 16
[5 empty] | 0 |  |  |  | Count of 9031 detections: | 2

## T27_lims_s2437_pspk
Table 27. Recovery of pesticides determined by direct aqueous-injection liquid chromatography-tandem mass spectrometry in laboratory reagent-spike samples collected during implementation of the method from May 2012 to June 2013 [Results are highlighted in bold if recovery outside study objectives of 100 ± 30 percent or relative standard deviation (RSD) >30 percent. Validation quality code: A, no qualification; na, not applicable; E, estimated; D, deleted; 9030, moved to positive electrospray ionization (ESI+) method; 9031,moved to ESI- method. USGS, U.S. Geological Survey; LRL, laboratory reporting level; ng/L, nanograms per liter; N, number of spike samples; recov., recovery; %, percent; conc. concentration; Min., minimum; Max., maximum; F-pseu, F-pseudoigma, interquartile range/1.349]
Validation quality code | Lab code | Analytical use | Analytical method group | Pesticide group | Pesticide compound | USGS parameter code | LRL (ng/L) | N | Mean recov. (%) | RSD (%) | Min. (ng/L) | Max. (ng/L) | Median recov. (%) | F-pseu (%) | Recovery 25th percentile (%) | Recovery 75th percentile (%)
A | 9030 | Analyte | acetanilide & amide | Acetochlor | Acetochlor | 68520 | 10 | 32 | 94.3071 | 16.9231 | 65.5805 | 131.5039 | 94.7443 | 18.7364 | 79.1917 | 104.4672
A | 9030 | Analyte |  |  | Hydroxyacetochlor | 68615 | 8 | 32 | 89.4696 | 18.4383 | 59.2387 | 120.583 | 90.953 | 18.5435 | 78.3958 | 103.4109
A | 9030 | Analyte |  | Acetochlor/Metolachlor | 2-Chloro-N-(2-ethyl-6-methylphenyl)acetamide | 68521 | 4 | 32 | 90.3657 | 19.4779 | 60.9131 | 127.6729 | 89.7238 | 19.7437 | 78.0587 | 104.693
A | 9030 | Analyte |  | Alachlor | 2-Chloro-2,6-diethylacetanilide | 68525 | 4 | 32 | 91.4214 | 17.8152 | 59.6756 | 121.5203 | 90.567 | 17.5394 | 78.8005 | 102.4612
A | 9030 | Analyte |  |  | Alachlor | 65064 | 8 | 32 | 95.3189 | 12.297 | 78.3468 | 123.1114 | 96.8811 | 12.907 | 84.0666 | 101.4782
A | 9030 | Analyte |  |  | Hydroxyalachlor | 68616 | 8 | 32 | 91.089 | 17.738 | 62.9971 | 119.8493 | 89.6025 | 20.3256 | 78.8991 | 106.3182
A | 9030 | Analyte |  | Dimethenamid | Dimethenamid | 68580 | 2 | 32 | 97.3479 | 14.7905 | 69.7952 | 129.7582 | 97.9423 | 14.0681 | 88.0646 | 107.0425
A | 9030 | Analyte |  | Metolachlor | 2-(2-ethyl-6-methyl-phenyl)aminopropan-1-ol | 68595 | 4 | 32 | 89.9626 | 16.4604 | 62.7086 | 119.4907 | 88.5641 | 12.7154 | 81.6088 | 98.7618
A | 9030 | Analyte |  |  | Dechlorometolachlor | 68562 | 2 | 32 | 89.6737 | 10.6548 | 72.8131 | 109.3682 | 90.4729 | 11.604 | 81.6004 | 97.2541
A | 9030 | Analyte |  |  | Hydroxymetolachlor | 68622 | 2 | 32 | 91.5132 | 11.7568 | 72.0428 | 112.4461 | 92.3432 | 12.4985 | 81.796 | 98.6565
A | 9030 | Analyte |  |  | Metolachlor | 65090 | 8 | 32 | 93.7041 | 9.628 | 80.3589 | 119.6047 | 92.6075 | 7.5396 | 87.2927 | 97.4636
A | 9030 | Analyte |  |  | Metolachlor hydroxy morpholinone | 68649 | 6 | 32 | 89.765 | 18.59 | 53.4945 | 120.2614 | 87.6642 | 16.6853 | 78.5355 | 101.044
A | 9030 | Analyte |  | Pronamide | Pronamide | 67706 | 4 | 32 | 93.9135 | 17.2291 | 64.421 | 129.0775 | 94.9485 | 16.7115 | 82.8943 | 105.4381
A | 9030 | Analyte |  | Propanil | Propanil | 66641 | 10 | 32 | 98.5551 | 17.4769 | 65.0145 | 131.968 | 97.5258 | 17.3134 | 86.3063 | 109.662
A | 9030 | Analyte | acid | Bentazone | 2-Amino-N-isopropylbenzamide | 68503 | 4 | 32 | 86.0737 | 12.0048 | 68.5481 | 111.6236 | 84.7959 | 8.8314 | 79.6045 | 91.518
A | 9030 | Analyte | carbamate & thiocarbamate | Aldicarb | Aldicarb | 68528 | 8 | 32 | 84.4421 | 15.3972 | 62.1496 | 114.0678 | 81.9618 | 13.647 | 76.2425 | 94.6523
A | 9030 | Analyte |  |  | Aldicarb sulfone | 68529 | 20 | 32 | 94.5263 | 18.6504 | 68.6342 | 132.7029 | 92.0862 | 19.2065 | 79.5054 | 105.415
A | 9030 | Analyte |  |  | Aldicarb sulfoxide | 68530 | 2 | 32 | 90.8966 | 16.0002 | 67.0671 | 132.2721 | 89.1323 | 9.4689 | 82.601 | 95.3745
A | 9030 | Analyte |  | Butylate | Butylate | 65068 | 6 | 32 | 90.5047 | 13.9942 | 69.7326 | 116.1862 | 89.7324 | 13.8659 | 80.4478 | 99.1529
A | 9030 | Analyte |  | Carbaryl | Carbaryl | 65069 | 6 | 32 | 90.3848 | 12.6028 | 66.9258 | 112.7088 | 89.1523 | 12.3259 | 81.7644 | 98.392
A | 9030 | Analyte |  | Carbofuran | Carbofuran | 65070 | 4 | 32 | 91.8224 | 14.8025 | 72.3429 | 135.9243 | 89.3185 | 14.6849 | 81.2706 | 101.0806
A | 9030 | Analyte |  | EPTC | EPTC | 65080 | 206 | 32 | 93.0699 | 17.6535 | 60.492 | 119.6439 | 95.5148 | 18.4641 | 81.8878 | 106.7959
A | 9030 | Analyte |  |  | EPTC R248722 | 68594 | 4 | 32 | 95.9442 | 14.3673 | 73.3321 | 121.4354 | 95.3564 | 16.796 | 84.7057 | 107.3635
A | 9030 | Analyte |  | Methomyl | Methomyl | 68645 | 2 | 32 | 90.5229 | 19.9995 | 59.2451 | 133.2834 | 89.6813 | 15.2523 | 80.8465 | 101.4218
A | 9030 | Analyte |  |  | Methomyl oxime | 68646 | 78 | 32 | 91.5315 | 23.4607 | 54.7088 | 159.5112 | 90.8244 | 20.6378 | 73.9698 | 101.8101
A | 9030 | Analyte |  | Molinate | 4-Hydroxy molinate | 68515 | 6 | 32 | 92.4528 | 15.0892 | 68.3719 | 117.8711 | 90.9723 | 14.9059 | 82.802 | 102.9101
A | 9030 | Analyte |  |  | Carboxy molinate | 68549 | 22 | 32 | 92.4286 | 16.5616 | 72.6954 | 128.8844 | 91.7057 | 19.3858 | 79.818 | 105.9694
A | 9030 | Analyte |  |  | Molinate | 65091 | 18 | 32 | 93.9705 | 17.4324 | 64.8646 | 124.8987 | 95.9387 | 18.5597 | 80.7888 | 105.8258
A | 9030 | Analyte |  | Oxamyl | Oxamyl | 68664 | 2 | 32 | 84.3109 | 16.1409 | 58.3186 | 116.5375 | 84.6865 | 12.3249 | 75.2091 | 91.8354
A | 9030 | Analyte |  |  | Oxamyl oxime | 68665 | 2 | 32 | 94.8402 | 18.0652 | 66.0922 | 147.0495 | 93.6177 | 13.2517 | 84.0372 | 101.9137
A | 9030 | Analyte |  | Propoxur | Propoxur | 68679 | 4 | 32 | 100.119 | 15.1521 | 76.4768 | 143.5253 | 99.8195 | 14.2413 | 89.1593 | 108.3708
A | 9030 | Analyte |  | thiobencarb | 4-Chlorobenzylmethyl sulfoxide | 68514 | 4 | 32 | 96.4287 | 17.1467 | 51.4854 | 134.1197 | 97.5452 | 12.2777 | 89.964 | 106.5266
A | 9030 | Analyte |  | Thiobencarb | Thiobencarb | 65107 | 2 | 32 | 93.3473 | 13.2305 | 74.1493 | 114.762 | 92.6496 | 15.2266 | 81.8306 | 102.3713
A | 9030 | Analyte |  | Triallate | Triallate | 68710 | 16 | 32 | 89.9545 | 12.5566 | 66.4329 | 110.7452 | 91.3207 | 11.6632 | 81.5217 | 97.2554
A | 9030 | Analyte | fungicide | Azoxystrobin | Azoxystrobin | 66589 | 2 | 32 | 96.5105 | 13.9765 | 74.7312 | 116.7162 | 94.7969 | 16.9827 | 86.0239 | 108.9336
A | 9030 | Analyte |  | Benomyl | 2-Aminobenzimidazole | 68502 | 8 | 32 | 92.8873 | 12.9705 | 69.0279 | 127.4349 | 91.2853 | 13.0571 | 84.7428 | 102.3568
A | 9030 | Analyte |  | Fentin | Fentin | 68603 | 16 | 32 | 93.3666 | 15.3699 | 67.191 | 123.9089 | 88.3948 | 15.2309 | 83.1744 | 103.7209
A | 9030 | Analyte |  | Kresoxim-methyl | Kresoxim-methyl | 67670 | 4 | 32 | 90.5374 | 15.6733 | 62.8517 | 121.227 | 90.0553 | 18.2632 | 78.7591 | 103.3962
A | 9030 | Analyte |  | Metalaxyl | Metalaxyl | 68437 | 6 | 32 | 94.3628 | 13.4328 | 70.8665 | 119.5357 | 92.5418 | 10.943 | 85.9624 | 100.7245
A | 9030 | Analyte |  | Metconazole | Metconazole | 66620 | 4 | 32 | 95.7524 | 13.0861 | 71.0223 | 117.447 | 95.5509 | 12.3881 | 88.1924 | 104.9039
A | 9030 | Analyte |  | Myclobutanil | Myclobutanil | 66632 | 6 | 32 | 95.7924 | 16.1817 | 70.1696 | 122.6537 | 96.7712 | 18.3637 | 81.4772 | 106.2499
A | 9030 | Analyte |  | Propiconazole | Propiconazole | 66643 | 6 | 32 | 94.2582 | 14.4261 | 72.7409 | 119.611 | 93.2559 | 13.7059 | 83.5665 | 102.0557
A | 9030 | Analyte |  | Pyraclostrobin | Pyraclostrobin | 66646 | 2 | 32 | 91.8027 | 14.4012 | 71.7659 | 120.2311 | 91.722 | 16.7897 | 79.2575 | 101.9069
A | 9030 | Analyte |  | Tebuconazole | Tebuconazole | 66649 | 4 | 32 | 97.4656 | 14.9839 | 72.0083 | 124.0288 | 96.4644 | 15.2813 | 87.7027 | 108.3171
A | 9030 | Analyte |  | Tetraconazole | Tetraconazole | 66654 | 6 | 32 | 94.1352 | 16.1664 | 72.0222 | 121.3591 | 93.1937 | 19.2746 | 80.3456 | 106.347
A | 9030 | Analyte |  | Trifloxystrobin | Trifloxystrobin | 66660 | 2 | 32 | 87.4205 | 13.3461 | 64.729 | 112.7278 | 87.1448 | 12.7351 | 79.0242 | 96.2038
A | 9030 | Analyte | miscellaneous | Bromacil | Bromacil | 68542 | 6 | 32 | 98.7262 | 16.483 | 63.9355 | 144.9023 | 98.4227 | 15.4244 | 88.7032 | 109.5107
A | 9030 | Analyte |  | Etoxazole | Etoxazole | 68598 | 2 | 31 | 80.5409 | 16.4718 | 62.4857 | 121.1385 | 77.1736 | 13.4329 | 69.7117 | 87.8327
A | 9030 | Analyte |  | Flubendiamide | Desiodo flubendiamide | 68563 | 4 | 32 | 84.1328 | 17.6816 | 63.546 | 119.3386 | 83.0254 | 15.4955 | 70.9286 | 91.832
A | 9030 | Analyte |  | Imazamox | Imazamox | 68625 | 28 | 32 | 92.039 | 10.5032 | 73.3628 | 114.2386 | 91.7357 | 10.3563 | 83.1511 | 97.1218
A | 9030 | Analyte |  | Imazaquin | Imazaquin | 61682 | 18 | 32 | 90.6048 | 10.969 | 73.6705 | 113.5809 | 88.9471 | 10.8614 | 82.3594 | 97.0114
A | 9030 | Analyte |  | Imazethapyr | Imazethapyr | 61683 | 20 | 32 | 90.4772 | 14.6696 | 76.158 | 119.5385 | 85.7276 | 14.6397 | 81.1615 | 100.9104
A | 9030 | Analyte |  | Imidacloprid | Imidacloprid | 68426 | 10 | 32 | 88.011 | 9.3973 | 71.6005 | 104.2017 | 85.9796 | 8.4686 | 81.1067 | 92.5308
A | 9030 | Analyte |  | Indoxacarb | Indoxacarb | 68627 | 8 | 32 | 78.7471 | 18.3821 | 51.0616 | 104.7658 | 78.7954 | 16.0528 | 67.3821 | 89.0373
A | 9030 | Analyte |  | Isoxaflutole | Isoxaflutole | 68632 | 12 | 32 | 84.4908 | 16.7338 | 56.2433 | 112.2178 | 85.4741 | 16.1175 | 73.7251 | 95.4676
A | 9030 | Analyte |  | Lactofen | Lactofen | 68638 | 2 | 32 | 72.5952 | 30.9434 | 28.8158 | 140.1869 | 72.3687 | 20.0742 | 58.2683 | 85.3484
A | 9030 | Analyte |  | Methoxyfenozide | Methoxyfenozide | 68647 | 2 | 32 | 99.7379 | 16.6063 | 66.5047 | 130.6298 | 101.1127 | 16.3511 | 88.9828 | 111.0405
A | 9030 | Analyte |  | Norflurazon | Demethyl norflurazon | 68567 | 4 | 32 | 93.5896 | 11.2906 | 70.0298 | 109.8919 | 96.8398 | 10.7057 | 85.4843 | 99.9263
A | 9030 | Analyte |  |  | Norflurazon | 67685 | 4 | 32 | 100.0705 | 13.2771 | 74.1097 | 130.8699 | 99.4455 | 11.1178 | 92.1891 | 107.1869
A | 9030 | Analyte |  | Pendimethalin | Pendimethalin | 65098 | 2 | 32 | 94.257 | 14.508 | 69.3775 | 119.008 | 95.4429 | 14.3935 | 83.7205 | 103.1374
A | 9030 | Analyte |  | Propargite | Propargite | 68677 | 2 | 31 | 80.8328 | 15.766 | 50.9079 | 102.6237 | 81.3268 | 15.4237 | 69.4902 | 90.2968
A | 9030 | Analyte |  | Pymetrozine | Pymetrozine | 68680 | 2 | 32 | 101.7049 | 23.7534 | 76.4382 | 190.0944 | 93.7734 | 19.7191 | 86.4158 | 113.0168
A | 9030 | Analyte |  | Pyridaben | Pyridaben | 68682 | 2 | 31 | 82.161 | 16.453 | 62.0998 | 123.7422 | 81.0058 | 10.3542 | 73.4229 | 87.3907
A | 9030 | Analyte |  | Pyriproxyfen | Pyriproxyfen | 68683 | 2 | 32 | 94.0513 | 13.4817 | 71.5782 | 125.418 | 91.666 | 11.7096 | 86.8991 | 102.6954
A | 9030 | Analyte |  | Tebufenozide | Tebufenozide | 68692 | 2 | 32 | 99.9216 | 13.8712 | 77.4767 | 122.6177 | 100.1655 | 16.1548 | 89.345 | 111.1379
A | 9030 | Analyte |  | Terbacil | Terbacil | 68698 | 20 | 32 | 100.9083 | 20.969 | 58.8789 | 148.0096 | 104.4261 | 22.0717 | 83.7655 | 113.5402
A | 9030 | Analyte | organophosphate | Acephate | Acephate | 68519 | 10 | 32 | 98.9797 | 18.1937 | 73.8632 | 171.9469 | 97.1726 | 10.2613 | 89.6176 | 103.4602
A | 9030 | Analyte |  |  | Methamidophos | 68644 | 2 | 32 | 101.7158 | 13.0554 | 76.712 | 136.9672 | 101.3099 | 12.3563 | 92.1425 | 108.8113
A | 9030 | Analyte |  | Azinphos-methyl | Azinphos-methyl | 65066 | 8 | 32 | 100.9898 | 14.4046 | 79.1973 | 142.9752 | 101.0045 | 12.2433 | 89.6748 | 106.1909
A | 9030 | Analyte |  |  | Azinphos-methyl oxon | 68211 | 14 | 32 | 102.104 | 14.1302 | 73.7339 | 137.0261 | 101.8292 | 15.7321 | 91.7679 | 112.9905
A | 9030 | Analyte |  | Chlorpyrifos | Chlorpyrifos | 65072 | 4 | 32 | 86.1979 | 13.4733 | 58.852 | 108.6342 | 85.2582 | 10.8029 | 79.4382 | 94.0113
A | 9030 | Analyte |  |  | Chlorpyrifos oxon | 68216 | 4 | 32 | 93.6538 | 17.8161 | 60.0218 | 127.8627 | 92.6335 | 16.9142 | 79.8772 | 102.6944
A | 9030 | Analyte |  | Diazinon | Diazinon | 65078 | 2 | 32 | 87.9292 | 13.1576 | 64.2078 | 110.2489 | 88.0087 | 12.6575 | 79.3933 | 96.4682
A | 9030 | Analyte |  |  | Hydroxydiazinon | 68618 | 10 | 32 | 84.2987 | 9.1596 | 67.4848 | 98.2438 | 84.5426 | 9.8735 | 78.0303 | 91.3497
A | 9030 | Analyte |  |  | Pyrimidinol | 68505 | 20 | 32 | 113.8624 | 16.1202 | 84.3161 | 158.4986 | 108.3198 | 20.4048 | 98.6485 | 126.1746
A | 9030 | Analyte |  | Dicrotophos | Dicrotophos | 68573 | 4 | 32 | 102.5279 | 13.208 | 82.0793 | 134.7552 | 100.6094 | 15.2028 | 92.3583 | 112.8669
A | 9030 | Analyte |  | Dimethoate | Dimethoate | 66596 | 2 | 32 | 102.1195 | 13.9192 | 81.8718 | 142.6711 | 98.886 | 15.4369 | 91.5011 | 112.3256
A | 9030 | Analyte |  |  | Dimethoate oxon | 68661 | 2 | 32 | 93.2028 | 13.8276 | 66.8976 | 124.202 | 92.7147 | 11.0259 | 86.0808 | 100.9547
A | 9030 | Analyte |  | Disulfoton | Disulfoton | 67595 | 12 | 32 | 85.1382 | 14.0477 | 58.9738 | 101.9053 | 84.4611 | 11.674 | 79.1468 | 94.895
A | 9030 | Analyte |  |  | Disulfoton oxon | 68586 | 2 | 32 | 103.6611 | 12.0646 | 81.6254 | 132.4076 | 103.9025 | 13.3451 | 92.3053 | 110.3078
A | 9030 | Analyte |  |  | Disulfoton oxon sulfone | 68588 | 6 | 32 | 95.4074 | 11.0894 | 75.4713 | 123.9901 | 93.608 | 7.8187 | 90.8051 | 101.3525
A | 9030 | Analyte |  |  | Disulfoton oxon sulfoxide | 68587 | 6 | 32 | 103.5967 | 14.7593 | 80.7461 | 142.3763 | 101.1372 | 11.8961 | 93.1894 | 109.2373
A | 9030 | Analyte |  |  | Disulfoton sulfone | 68589 | 8 | 32 | 96.3273 | 13.4422 | 76.6834 | 132.3479 | 96.3693 | 10.623 | 86.9519 | 101.2824
A | 9030 | Analyte |  |  | Disulfoton sulfoxide | 68590 | 4 | 32 | 107.6747 | 16.5455 | 83.4793 | 150.5519 | 102.7053 | 18.4377 | 95.2886 | 120.1611
A | 9030 | Analyte |  | Ethoprop | Ethoprop | 68596 | 4 | 32 | 91.8558 | 10.0798 | 74.3981 | 107.1651 | 93.3885 | 12.1857 | 83.219 | 99.6576
A | 9030 | Analyte |  |  | O-Ethyl-O-methyl-S-propyl phosphorothioate | 68597 | 4 | 32 | 103.6959 | 13.6149 | 83.3448 | 137.6129 | 103.0763 | 14.1015 | 93.5003 | 112.5233
A | 9030 | Analyte |  |  | O-Ethyl-S-methyl-S-propyl phosphorodithioate | 68657 | 4 | 32 | 102.9923 | 13.3216 | 78.6866 | 136.4367 | 102.6782 | 11.5612 | 95.0264 | 110.6225
A | 9030 | Analyte |  |  | O-Ethyl-S-propyl phosphorothioate | 68658 | 64 | 32 | 102.9661 | 16.4969 | 75.5865 | 134.6609 | 102.1384 | 18.0396 | 89.5063 | 113.8418
A | 9030 | Analyte |  | Fenamiphos | Fenamiphos | 68599 | 2 | 32 | 89.9552 | 9.9362 | 71.8927 | 108.9926 | 91.0535 | 10.2895 | 82.645 | 96.5255
A | 9030 | Analyte |  |  | Fenamiphos sulfone | 68600 | 4 | 32 | 101.6523 | 15.8538 | 81.858 | 154.0483 | 100.1881 | 16.15 | 87.1976 | 108.9839
A | 9030 | Analyte |  |  | Fenamiphos sulfoxide | 68601 | 4 | 32 | 102.3231 | 14.7849 | 80.6094 | 136.9782 | 99.4908 | 12.8754 | 91.0889 | 108.4578
A | 9030 | Analyte |  | Fonofos | Fonofos | 65084 | 10 | 32 | 91.0726 | 12.606 | 66.3808 | 114.6523 | 89.8046 | 11.9758 | 83.7016 | 99.8569
A | 9030 | Analyte |  | Malathion | Malaoxon | 68240 | 2 | 32 | 84.8334 | 15.8277 | 55.7506 | 117.6737 | 82.7906 | 11.9527 | 77.0855 | 93.2097
A | 9030 | Analyte |  |  | Malathion | 65087 | 4 | 32 | 88.3878 | 11.4605 | 72.5824 | 113.7955 | 87.3477 | 9.4673 | 80.9775 | 93.7488
A | 9030 | Analyte |  | Methidathion | Methidathion | 65088 | 10 | 32 | 97.4265 | 12.7113 | 71.7332 | 122.8583 | 97.6826 | 14.9353 | 87.4205 | 107.5682
A | 9030 | Analyte |  | Naled | Dichlorvos | 68572 | 52 | 32 | 104.4251 | 19.144 | 67.1896 | 145.3166 | 103.9937 | 19.6465 | 89.9132 | 116.4163
A | 9030 | Analyte |  | Parathion-ethyl | Paraoxon | 68666 | 4 | 32 | 106.1473 | 14.0166 | 77.6958 | 150.1192 | 104.8485 | 13.5994 | 96.788 | 115.1336
A | 9030 | Analyte |  | Parathion-methyl | Paraoxon-methyl | 68648 | 18 | 32 | 97.3067 | 18.9831 | 59.6837 | 142.0699 | 95.68 | 16.7347 | 84.7374 | 107.3124
A | 9030 | Analyte |  | Phorate | Phorate | 68668 | 6 | 32 | 84.6071 | 13.772 | 61.2366 | 99.3965 | 85.4458 | 12.4836 | 77.4486 | 94.289
A | 9030 | Analyte |  |  | Phorate oxon | 68669 | 4 | 32 | 99.9595 | 13.3512 | 74.6466 | 131.4283 | 98.5846 | 11.7144 | 90.684 | 106.4866
A | 9030 | Analyte |  |  | Phorate oxon sulfone | 68670 | 20 | 32 | 93.2868 | 12.2696 | 72.382 | 116.4638 | 93.7351 | 12.3676 | 84.5151 | 101.199
A | 9030 | Analyte |  |  | Phorate oxon sulfoxide | 68671 | 6 | 32 | 96.5486 | 15.0833 | 70.5443 | 124.7764 | 96.1255 | 13.9617 | 86.7748 | 105.6092
A | 9030 | Analyte |  |  | Phorate sulfone | 68672 | 8 | 32 | 104.6335 | 15.1695 | 62.851 | 145.5238 | 104.3416 | 17.2547 | 92.8565 | 116.1332
A | 9030 | Analyte |  |  | Phorate sulfoxide | 68673 | 4 | 32 | 106.0409 | 15.9813 | 82.9696 | 143.9349 | 104.6728 | 17.3492 | 92.0726 | 115.4768
A | 9030 | Analyte |  | Profenofos | Profenofos | 68676 | 2 | 32 | 81.4471 | 14.4241 | 54.1323 | 97.6125 | 83.751 | 12.8532 | 72.4541 | 89.7931
A | 9030 | Analyte |  | Tebupirimfos | Tebupirimfos | 68693 | 2 | 32 | 84.7784 | 10.8901 | 64.3935 | 101.1357 | 86.8602 | 9.6188 | 78.185 | 91.1608
A | 9030 | Analyte |  |  | Tebupirimfos oxon | 68694 | 2 | 32 | 90.7023 | 10.2461 | 71.0902 | 109.7197 | 90.8179 | 10.2935 | 82.9269 | 96.8128
A | 9030 | Analyte |  | Terbufos | Terbufos | 68699 | 4 | 32 | 80.2194 | 12.1521 | 55.8606 | 103.2635 | 81.0182 | 9.8891 | 74.1181 | 87.4585
A | 9030 | Analyte |  |  | Terbufos oxon | 68700 | 4 | 32 | 80.6332 | 12.5817 | 60.6931 | 99.0835 | 79.4947 | 11.3655 | 74.2165 | 89.5485
A | 9030 | Analyte |  |  | Terbufos oxon sulfone | 68701 | 10 | 32 | 98.4436 | 10.1401 | 71.9794 | 117.2876 | 97.7078 | 9.6753 | 91.9326 | 104.9846
A | 9030 | Analyte |  |  | Terbufos oxon sulfoxide | 68702 | 4 | 32 | 99.8031 | 12.2812 | 79.1239 | 126.029 | 98.2117 | 12.4512 | 89.386 | 106.1827
A | 9030 | Analyte |  |  | Terbufos sulfone | 68703 | 10 | 32 | 103.5398 | 17.6217 | 73.9769 | 153.0512 | 104.5348 | 14.5112 | 90.9521 | 110.5276
A | 9030 | Analyte |  |  | Terbufos sulfoxide | 68704 | 2 | 32 | 105.6914 | 14.3768 | 83.3154 | 149.5926 | 103.1783 | 16.0408 | 93.3138 | 114.9528
A | 9030 | Analyte |  | Tribufos | Tribufos | 68711 | 2 | 31 | 72.5976 | 22.7621 | 41.2254 | 113.0013 | 72.8097 | 17.1398 | 62.8493 | 85.9709
A | 9030 | Analyte | pyrethroid & phenylpyrazine | Permethrin | cis-Permethrin | 68769 | 2 | 31 | 70.0606 | 20.9708 | 42.4048 | 117.4195 | 70.3105 | 14.7522 | 59.5607 | 79.4614
A | 9030 | Analyte |  |  | trans-Permethrin | 68708 | 4 | 31 | 69.8565 | 20.0146 | 39.2117 | 107.9793 | 71.0319 | 14.1987 | 59.3539 | 78.5079
A | 9030 | Analyte |  | Piperonyl butoxide | Piperonyl butoxide | 65102 | 4 | 32 | 85.9009 | 17.3476 | 51.4114 | 121.0384 | 84.9351 | 16.8653 | 73.7942 | 96.5455
A | 9030 | Analyte | sulfonylurea & urea | Chlorimuron-ethyl | Chlorimuron-ethyl | 68872 | 8 | 32 | 78.4725 | 16.9817 | 53.4568 | 109.2079 | 78.4647 | 17.3249 | 66.5319 | 89.9032
A | 9030 | Analyte |  | Chlorsulfuron | Chlorsulfuron | 61678 | 50 | 32 | 85.4957 | 14.995 | 57.1623 | 122.9034 | 84.9164 | 9.6686 | 77.2003 | 90.2433
A | 9030 | Analyte |  | Diflubenzuron | Diflubenzuron | 68576 | 6 | 32 | 90.3014 | 10.9921 | 70.5585 | 111.474 | 89.3265 | 9.2586 | 83.7532 | 96.243
A | 9030 | Analyte |  | Diflufenzopyr | Hydroxyphthalazinone | 68623 | 46 | 32 | 93.7845 | 10.7919 | 73.9405 | 112.6224 | 95.0129 | 12.8053 | 84.9426 | 102.2171
A | 9030 | Analyte |  |  | Phthalazinone | 68675 | 6 | 32 | 92.0815 | 16.2859 | 71.0915 | 134.9829 | 90.4745 | 15.2608 | 81.0445 | 101.6313
A | 9030 | Analyte |  | Diuron | 3,4-Dichlorophenylurea | 68226 | 144 | 32 | 90.217 | 14.0768 | 69.52 | 116.972 | 88.2089 | 12.3854 | 81.3496 | 98.0574
A | 9030 | Analyte |  |  | Diuron | 66598 | 4 | 32 | 92.6735 | 10.905 | 74.3936 | 121.6341 | 93.3163 | 7.9963 | 87.8285 | 98.6155
A | 9030 | Analyte |  |  | N-(3,4-Dichlorophenyl)-N-methylurea | 68231 | 4 | 32 | 91.4909 | 10.7183 | 69.3382 | 109.2133 | 93.676 | 10.48 | 82.7717 | 96.9092
A | 9030 | Analyte |  | Fluometuron | 4-Hydroxy-tert-fluometuron | 68619 | 20 | 32 | 88.1022 | 16.3643 | 64.6761 | 121.5341 | 87.0813 | 16.6219 | 75.6022 | 98.0251
A | 9030 | Analyte |  |  | Demethyl fluometuron | 68591 | 2 | 32 | 89.8948 | 8.8728 | 76.4177 | 104.1707 | 89.5007 | 9.0203 | 84.7047 | 96.873
A | 9030 | Analyte |  |  | Fluometuron | 68608 | 2 | 32 | 93.8585 | 7.4714 | 79.2086 | 104.02 | 95.35 | 9.1186 | 87.6478 | 99.9488
A | 9030 | Analyte |  |  | Hydroxy mono demethyl fluometuron | 68617 | 10 | 32 | 86.0951 | 14.5193 | 49.3983 | 109.8229 | 87.4085 | 9.1747 | 82.0288 | 94.4054
A | 9030 | Analyte |  |  | Hydroxyfluometuron | 68620 | 8 | 32 | 82.754 | 19.5273 | 37.6053 | 109.9969 | 85.2746 | 11.2825 | 75.6516 | 90.8718
A | 9030 | Analyte |  | Halosulfuron-methyl | Halosulfuron-methyl | 61680 | 22 | 32 | 83.9202 | 11.5576 | 65.3497 | 106.3568 | 83.441 | 9.2487 | 76.3247 | 88.8013
A | 9030 | Analyte |  | Linuron | Linuron | 68639 | 6 | 32 | 93.4187 | 11.4737 | 71.4629 | 118.3184 | 92.2306 | 9.4824 | 86.7503 | 99.5422
A | 9030 | Analyte |  | Nicosulfuron | Nicosulfuron | 61685 | 14 | 32 | 81.7335 | 15.8888 | 64.8268 | 117.3301 | 77.9642 | 13.049 | 72.3556 | 89.9587
A | 9030 | Analyte |  | Orthosulfamuron | Orthosulfamuron | 68662 | 6 | 32 | 65.8612 | 26.5719 | 45.0095 | 105.687 | 58.2358 | 21.5578 | 51.629 | 80.7106
A | 9030 | Analyte |  | Prosulfuron | Prosulfuron | 61687 | 8 | 32 | 78.0186 | 19.2834 | 57.5841 | 110.486 | 73.784 | 12.5038 | 67.0866 | 83.9542
A | 9030 | Analyte |  | Siduron | Siduron | 68686 | 4 | 32 | 94.4752 | 11.1342 | 74.4965 | 115.022 | 93.4244 | 12.1448 | 86.6517 | 103.035
A | 9030 | Analyte |  | Sulfometuron-methyl | Sulfometuron-methyl | 68688 | 4 | 32 | 89.101 | 11.0292 | 74.9029 | 122.3916 | 87.354 | 6.8296 | 82.6047 | 91.8178
A | 9030 | Analyte |  | Sulfosulfuron | Sulfosulfuron | 68689 | 8 | 32 | 89.615 | 11.0678 | 69.5087 | 120.752 | 89.8285 | 8.1362 | 84.0734 | 95.0491
A | 9030 | Analyte |  |  | Sulfosulfuron ethyl sulfone | 68690 | 2 | 32 | 95.4145 | 11.0772 | 72.6779 | 121.5061 | 96.5233 | 9.6546 | 87.2351 | 100.2591
A | 9030 | Analyte |  | Tebuthiuron | Hydroxytebuthurion | 68621 | 10 | 32 | 81.1266 | 30.0663 | 0 | 118.5747 | 87.3748 | 13.5221 | 75.2032 | 93.4446
A | 9030 | Analyte |  |  | Tebuthiuron | 68695 | 2 | 32 | 96.0769 | 8.5562 | 80.0827 | 115.9435 | 96.82 | 9.6688 | 89.2673 | 102.3106
A | 9030 | Analyte |  |  | Tebuthiuron TP 104 | 68575 | 6 | 35 | 93.8113 | 7.9032 | 82.701 | 108.344 | 93.6766 | 9.3249 | 86.6314 | 99.2107
A | 9030 | Analyte |  |  | Tebuthiuron TP 106 | 68714 | 76 | 32 | 89.5373 | 15.1426 | 65.5628 | 119.3318 | 89.1102 | 16.8118 | 79.8273 | 102.5063
A | 9030 | Analyte |  |  | Tebuthiuron TP 108 | 68696 | 10 | 32 | 91.8261 | 10.1967 | 74.2999 | 109.6211 | 91.0632 | 12.3493 | 83.5715 | 100.2306
A | 9030 | Analyte |  |  | Tebuthiuron TP 109 (OH) | 68697 | 16 | 32 | 87.3982 | 13.6875 | 67.8702 | 110.9149 | 86.2554 | 15.6678 | 76.8608 | 97.9967
A | 9030 | Analyte | triazine | Ametryn | Ametryn | 68533 | 2 | 32 | 92.9184 | 9.0088 | 78.173 | 109.4488 | 92.7335 | 10.3379 | 85.8578 | 99.8037
A | 9030 | Analyte |  | Atrazine | 2-Hydroxy-4-isopropylamino-6-amino-s-triazine | 68659 | 4 | 32 | 84.689 | 14.386 | 68.1162 | 119.5216 | 82.1322 | 9.9557 | 77.217 | 90.6472
A | 9030 | Analyte |  |  | 2-Hydroxy-6-ethylamino-4-amino-s-triazine | 68656 | 6 | 32 | 85.2186 | 11.9304 | 68.9428 | 108.9643 | 83.1569 | 8.7895 | 78.5125 | 90.3695
A | 9030 | Analyte |  |  | 2-Hydroxyatrazine | 68660 | 8 | 32 | 84.4008 | 12.0626 | 72.9352 | 112.3929 | 83.2665 | 8.9569 | 75.8027 | 87.8856
A | 9030 | Analyte |  |  | Atrazine | 65065 | 4 | 32 | 91.286 | 10.5994 | 69.4113 | 109.1769 | 90.6939 | 9.5862 | 83.8316 | 96.7634
A | 9030 | Analyte |  |  | Deethylatrazine | 68552 | 10 | 32 | 88.5747 | 13.5609 | 69.1372 | 120.3365 | 85.3899 | 10.0347 | 80.7334 | 94.2702
A | 9030 | Analyte |  |  | Deisopropylatrazine | 68550 | 20 | 32 | 88.2295 | 16.0344 | 68.0615 | 124.752 | 87.7746 | 12.6312 | 78.6703 | 95.7098
A | 9030 | Analyte |  |  | Didealkylatrazine | 68547 | 24 | 32 | 90.8826 | 14.7225 | 64.487 | 122.9555 | 90.2336 | 11.5623 | 81.6786 | 97.2762
A | 9030 | Analyte |  | Cyanazine | Cyanazine | 66592 | 16 | 32 | 91.781 | 15.6599 | 60.3905 | 124.055 | 94.2065 | 14.6944 | 79.1201 | 98.9428
A | 9030 | Analyte |  | Hexazinone | 4-Hydroxyhexazinone A | 68517 | 2 | 32 | 89.9259 | 12.6644 | 69.5948 | 110.0414 | 90.0446 | 12.1234 | 80.6714 | 97.0258
A | 9030 | Analyte |  |  | Demethyl hexazinone B | 68566 | 2 | 32 | 91.662 | 10.5194 | 70.5697 | 107.3648 | 92.4885 | 10.2832 | 85.958 | 99.83
A | 9030 | Analyte |  |  | Hexazinone | 65085 | 2 | 32 | 93.278 | 9.5893 | 74.865 | 110.1742 | 93.7514 | 8.1865 | 87.4414 | 98.485
A | 9030 | Analyte |  |  | Hexazinone TP C | 68612 | 2 | 32 | 86.9018 | 11.9344 | 68.4371 | 107.4963 | 87.4258 | 10.8045 | 77.9607 | 92.5359
A | 9030 | Analyte |  |  | Hexazinone TP F | 68574 | 10 | 32 | 89.3072 | 10.5192 | 65.3195 | 106.0089 | 89.7934 | 5.7358 | 85.8588 | 93.5965
A | 9030 | Analyte |  |  | Hexazinone TP G | 68713 | 22 | 32 | 87.1556 | 12.7153 | 69.3206 | 115.6124 | 85.237 | 12.5896 | 77.8721 | 94.8555
A | 9030 | Analyte |  | Metribuzin | Metribuzin | 68652 | 20 | 32 | 96.0803 | 11.6053 | 75.691 | 115.1877 | 94.3351 | 11.2911 | 87.7786 | 103.0103
A | 9030 | Analyte |  |  | Metribuzin DA | 68568 | 8 | 32 | 88.5183 | 11.5 | 70.2473 | 112.7371 | 87.3194 | 9.1218 | 83.255 | 95.5603
A | 9030 | Analyte |  | Prometon | Prometon | 67702 | 4 | 32 | 90.3685 | 11.3898 | 73.4997 | 112.9317 | 88.6158 | 10.8823 | 83.8541 | 98.5344
A | 9030 | Analyte |  | Prometryn | Deisopropyl prometryn | 68564 | 2 | 32 | 88.149 | 10.3091 | 69.8408 | 107.5806 | 88.2571 | 8.4624 | 81.9378 | 93.3536
A | 9030 | Analyte |  |  | Prometryn | 65103 | 2 | 32 | 92.5329 | 9.2787 | 77.0271 | 112.1064 | 92.5951 | 9.5232 | 85.6548 | 98.5016
A | 9030 | Analyte |  | Propazine | Propazine | 68678 | 4 | 32 | 91.4568 | 11.904 | 66.5412 | 111.9396 | 92.1517 | 11.7953 | 84.1346 | 100.0465
A | 9030 | Analyte |  | Simazine | Hydroxysimazine | 68624 | 2 | 32 | 98.1828 | 19.6258 | 74.0237 | 142.4124 | 89.5343 | 23.9359 | 82.3476 | 114.6371
A | 9030 | Analyte |  |  | Simazine | 65105 | 10 | 32 | 92.4693 | 11.4229 | 75.9119 | 123.2246 | 91.1405 | 11.9368 | 84.0996 | 100.2023
A | 9030 | Analyte |  | Terbuthylazine | Terbuthylazine | 66651 | 2 | 32 | 92.624 | 8.5532 | 78.3759 | 108.7209 | 91.6968 | 9.6225 | 86.258 | 99.2387
na | 9030 | Surrogate | acetanilide & amide | Acetochlor | Acetochlor-d11 (surrogate) | 90517 | 100 | 32 | 97.0767 | 8.4674 | 80.8546 | 110.7448 | 97.1164 | 10.1405 | 90.2519 | 103.9314
na | 9030 | Surrogate |  | Alachlor | Alachlor-d13 (surrogate) | 90518 | 100 | 32 | 98.5805 | 7.8422 | 87.0528 | 116.5746 | 98.8508 | 7.7477 | 92.597 | 103.0486
na | 9030 | Surrogate |  | Metolachlor | Metolachlor-d6 (surrogate) | 90553 | 100 | 32 | 99.1198 | 6.2448 | 79.5682 | 118.8056 | 98.9279 | 3.7954 | 96.5058 | 101.6258
na | 9030 | Surrogate | carbamate & thiocarbamate | Carbaryl | Carbaryl-d7 (surrogate) | 90519 | 100 | 32 | 92.6557 | 14.9562 | 56.8895 | 117.9757 | 95.2849 | 15.3757 | 83.3848 | 104.1267
na | 9030 | Surrogate |  | Carbofuran | Carbofuran-d3 (surrogate) | 90521 | 100 | 32 | 100.5547 | 8.2503 | 79.3425 | 117.43 | 100.7359 | 7.1425 | 96.8445 | 106.4797
na | 9030 | Surrogate |  | Thiobencarb | Thiobencarb-d10 (surrogate) | 90556 | 100 | 32 | 94.9949 | 7.8162 | 76.561 | 108.842 | 95.9485 | 7.6393 | 90.2888 | 100.5943
na | 9030 | Surrogate | fungicide | Benomyl | Carbendazim-d4 (surrogate) | 90520 | 100 | 32 | 98.642 | 6.6419 | 85.331 | 113.4691 | 99.1228 | 5.9257 | 95.0433 | 103.0371
na | 9030 | Surrogate | organophosphate | Diazinon | Diazinon-d10 (surrogate) | 90523 | 100 | 32 | 94.7933 | 9.1918 | 82.2796 | 120.5319 | 95.8001 | 9.3845 | 86.6034 | 99.2631
na | 9030 | Surrogate |  | Malathion | Malathion-d10 (surrogate) | 90552 | 100 | 32 | 85.9419 | 16.578 | 46.0092 | 101.6808 | 89.7235 | 12.0852 | 80.0511 | 96.354
na | 9030 | Surrogate | pyrethroid & organochlorine& phenylpyrazine | Permethrin | cis-Permethrin-13C6 (surrogate) | 90558 | 100 | 32 | 96.4941 | 30.8861 | 34.8526 | 206.8049 | 95.7747 | 18.1518 | 83.4275 | 107.9142
na | 9030 | Surrogate | sulfonylurea & urea | Diflubenzuron | Diflubenzuron-d4 (surrogate) | 90524 | 100 | 32 | 97.875 | 8.9546 | 77.0847 | 115.9446 | 98.801 | 8.2236 | 92.0349 | 103.1285
na | 9030 | Surrogate |  | Diuron | Diuron-d6 (surrogate) | 90808 | 100 | 32 | 99.2792 | 7.7403 | 82.7918 | 119.0083 | 97.6794 | 5.8889 | 95.0721 | 103.0162
na | 9030 | Surrogate |  | Linuron | Linuron-d6 (surrogate) | 90529 | 100 | 32 | 100.0183 | 5.3842 | 85.7116 | 113.5925 | 99.7081 | 3.3483 | 96.7505 | 101.2673
na | 9030 | Surrogate |  | Nicosulfuron | Nicosulfuron-d6 (surrogate) | L903L | 100 | 32 | 98.0353 | 6.5764 | 87.6161 | 115.347 | 96.9508 | 5.3919 | 93.8335 | 101.1072
na | 9030 | Surrogate | triazine | Atrazine | Deethylatrazine-d6 (surrogate) | 90522 | 100 | 32 | 103.2343 | 9.564 | 82.7995 | 118.8393 | 103.3574 | 10.7022 | 97.5748 | 112.0121
na | 9030 | Surrogate |  | Hexazinone | Hexazinone-d6 (surrogate) | 90527 | 100 | 32 | 100.2768 | 7.6914 | 82.4541 | 119.7047 | 100.3326 | 7.3584 | 94.7756 | 104.7021
A | 9031 | Analyte | acetanilide & amide | Acetochlor | Acetochlor OA | 68522 | 90 | 34 | 94.219 | 18.5708 | 64.4135 | 131.4368 | 95.1863 | 16.3066 | 83.0369 | 105.0346
A | 9031 | Analyte |  |  | Acetochlor SA | 68523 | 100 | 19 | 100.8218 | 37.1186 | 30.593 | 161.8417 | 94.8809 | 37.8617 | 81.0849 | 132.1603
A | 9031 | Analyte |  |  | Acetochlor SAA | 68524 | 176 | 34 | 92.9686 | 28.6563 | 44.0685 | 159.4918 | 89.5605 | 23.3917 | 76.2095 | 107.7649
A | 9031 | Analyte |  | Acetochlor/Metolachlor | sec-Acetochlor-OA | 68684 | 52 | 34 | 93.5775 | 13.9653 | 69.4591 | 125.7808 | 90.6277 | 13.6754 | 84.2532 | 102.7013
A | 9031 | Analyte |  | Alachlor | Alachlor OA | 68526 | 84 | 34 | 90.2214 | 16.4724 | 51.7718 | 122.9862 | 91.4566 | 13.026 | 82.2029 | 99.775
E | 9031 | Analyte |  |  | Alachlor SA | 68871 | 200 | 19 | 102.1508 | 59.3555 | 0 | 225.9958 | 93.9374 | 67.2048 | 59.0038 | 149.6631
A | 9031 | Analyte |  |  | Alachlor SAA | 68527 | 168 | 34 | 89.1716 | 21.35 | 65.5582 | 138.4933 | 86.7097 | 17.8206 | 73.7904 | 97.8304
A | 9031 | Analyte |  |  | sec-Alachlor-OA | 68685 | 136 | 34 | 94.2153 | 16.5912 | 58.4332 | 128.9063 | 93.7891 | 13.7973 | 84.0115 | 102.624
A | 9031 | Analyte |  | Dimethenamid | Dimethenamid OA | 68581 | 84 | 34 | 92.804 | 14.8332 | 70.0022 | 129.3581 | 89.0431 | 16.0051 | 82.6135 | 104.2043
A | 9031 | Analyte |  |  | Dimethenamid SA | 68582 | 78 | 34 | 94.3613 | 16.8357 | 62.7041 | 128.5875 | 93.2512 | 17.1975 | 82.1399 | 105.3393
A | 9031 | Analyte |  |  | Dimethenamid SAA | 68583 | 188 | 34 | 88.7899 | 25.403 | 53.2058 | 139.4228 | 85.6078 | 28.3881 | 68.9913 | 107.2869
A | 9031 | Analyte |  | Metolachlor | Metolachlor OA | 68650 | 148 | 34 | 90.8772 | 18.2134 | 59.4202 | 126.1148 | 94.2072 | 14.1537 | 80.188 | 99.2814
A | 9031 | Analyte |  |  | Metolachlor SA | 68651 | 68 | 34 | 92.4187 | 20.497 | 50.0885 | 136.8979 | 89.9029 | 14.1151 | 82.2883 | 101.3295
A | 9031 | Analyte | acid | 2,4-D | 2,4-D | 68500 | 62 | 34 | 91.2143 | 15.7975 | 63.7585 | 121.6399 | 91.8024 | 17.1348 | 79.1829 | 102.2977
A | 9031 | Analyte |  | Bentazon | Bentazon | 68538 | 10 | 34 | 97.8052 | 11.6061 | 76.5351 | 124.7852 | 97.1645 | 12.6376 | 89.2554 | 106.3036
A | 9031 | Analyte |  | Bromoxynil | Bromoxynil | 68543 | 78 | 34 | 89.3065 | 16.1772 | 65.3863 | 127.9767 | 88.7809 | 15.4747 | 76.3873 | 97.2627
A | 9031 | Analyte |  | MCPA | MCPA | 68641 | 96 | 34 | 86.1628 | 19.2999 | 51.1876 | 128.5606 | 85.42 | 15.944 | 75.4586 | 96.9671
A | 9031 | Analyte |  | Triclopyr | Triclopyr | 68712 | 88 | 34 | 91.7373 | 16.5872 | 52.5521 | 120.9384 | 92.3729 | 16.357 | 81.0851 | 103.1506
A | 9031 | Analyte | carbamate & thiocarbamate | Triallate | 2,3,3-trichloroprop-2-ene-1-SA | 68691 | 54 | 34 | 93.5725 | 14.9223 | 75.733 | 125.9863 | 90.0332 | 18.0712 | 81.9517 | 106.3297
A | 9031 | Analyte | fungicide | Chlorothalonil | 4-Hydroxychlorothalonil | 68336 | 98 | 34 | 96.4133 | 33.1052 | 24.3463 | 189.2418 | 98.1955 | 24.4934 | 78.5794 | 111.621
A | 9031 | Analyte | miscellaneous | Flubendiamide | Flubendiamide | 68606 | 4 | 34 | 102.8299 | 21.2495 | 64.8497 | 144.8173 | 108.2883 | 23.7978 | 85.8992 | 118.0024
A | 9031 | Analyte |  | Isoxaflutole | Diketonitrile isoxaflutole | 68578 | 62 | 34 | 118.7155 | 25.0046 | 68.2757 | 197.4115 | 116.1303 | 29.1918 | 96.7549 | 136.1346
A | 9031 | Analyte |  |  | Isoxaflutole acid RPA 203328 | 68633 | 8 | 34 | 98.7328 | 11.5577 | 76.6525 | 121.6982 | 96.415 | 14.0799 | 89.4089 | 108.4028
A | 9031 | Analyte |  | Oryzalin | Oryzalin | 68663 | 10 | 34 | 105.4081 | 14.4525 | 70.9094 | 132.609 | 104.8182 | 8.1742 | 101.4808 | 112.5078
A | 9031 | Analyte |  | Pendimethalin | 4-(Hydroxymethyl)pendimethalin | 68511 | 212 | 34 | 92.7862 | 27.2804 | 56.0485 | 165.2033 | 88.8173 | 27.5179 | 71.691 | 108.8126
A | 9031 | Analyte | pyrethroid & organochlorine& phenylpyrazine | Bifenthrin; lambda-cyhalothrin; Tefluthrin | cis-Cyhalothric acid | 68553 | 88 | 34 | 102.3512 | 28.512 | 55.7907 | 168.6529 | 100.4451 | 25.8642 | 83.1447 | 118.0355
A | 9031 | Analyte |  | Fipronil | Dechlorofipronil | 68561 | 4 | 34 | 83.9376 | 15.5211 | 64.2147 | 110.2779 | 84.9912 | 13.8802 | 72.5727 | 91.297
A | 9031 | Analyte |  |  | Desulfinylfipronil | 66607 | 4 | 34 | 86.3564 | 10.2635 | 71.2191 | 110.4743 | 86.1949 | 9.4963 | 80.5553 | 93.3657
A | 9031 | Analyte |  |  | Desulfinylfipronil amide | 68570 | 8 | 34 | 88.4748 | 15.093 | 60.8955 | 124.9076 | 89.2048 | 11.3031 | 80.2077 | 95.4556
A | 9031 | Analyte |  |  | Fipronil | 66604 | 4 | 34 | 86.2577 | 14.2732 | 63.8904 | 111.0591 | 87.2083 | 15.0166 | 75.9446 | 96.202
A | 9031 | Analyte |  |  | Fipronil amide | 68604 | 8 | 34 | 96.1152 | 17.7239 | 62.4787 | 126.7762 | 94.5346 | 20.5457 | 81.9042 | 109.6203
A | 9031 | Analyte |  |  | Fipronil sulfide | 66610 | 4 | 34 | 81.8788 | 15.492 | 55.7647 | 111.1172 | 82.0899 | 12.8826 | 73.3889 | 90.7675
A | 9031 | Analyte |  |  | Fipronil sulfonate | 68605 | 96 | 34 | 96.1212 | 24.5134 | 59.2375 | 145.431 | 95.3773 | 30.2551 | 74.2195 | 115.0336
A | 9031 | Analyte |  |  | Fipronil sulfone | 66613 | 4 | 34 | 73.7433 | 20.2708 | 42.857 | 102.6927 | 76.3242 | 11.9196 | 67.7699 | 83.8494
A | 9031 | Analyte |  | Permethrin | 3-Phenoxybenzoic acid | 68873 | 60 | 34 | 90.7089 | 10.79 | 74.54 | 113.1391 | 91.2031 | 10.8444 | 81.8859 | 96.515
A | 9031 | Analyte | sulfonylurea & urea | Diflufenzopyr | Diflufenzopyr | 68577 | 88 | 34 | 95.2543 | 24.6954 | 53.8097 | 166.053 | 98.7338 | 28.49 | 75.3149 | 113.7478
A | 9031 | Analyte |  | Sulfentrazone | Sulfentrazone | 68687 | 10 | 34 | 117.3724 | 17.6084 | 76.8596 | 156.4681 | 115.5525 | 16.9225 | 103.9736 | 126.8021
A | 9031 | Analyte | triazine | Hexazinone | Hexazinone TP E | 68614 | 76 | 34 | 92.3957 | 17.4625 | 51.6684 | 128.7 | 93.3086 | 11.528 | 83.8455 | 99.3968
A | 9031 | Analyte |  | Metribuzin | Metribuzin DADK | 68569 | 200 | 34 | 83.3362 | 31.537 | 19.5913 | 147.4481 | 83.4928 | 26.219 | 67.9657 | 103.3351
A | 9031 | Analyte |  |  | Metribuzin DK | 68653 | 200 | 34 | 85.387 | 21.0151 | 31.1492 | 116.2521 | 89.1942 | 13.5677 | 77.4409 | 95.7437
na | 9031 | Surrogate | acetanilide & amide | Butachlor sulfonic acid (surrogate) | Butachlor SA (surrogate) | 90624 | 100 | 34 | 105.2066 | 10.1913 | 85.8848 | 135.5028 | 104.1392 | 7.1971 | 98.7053 | 108.4142
na | 9031 | Surrogate |  | Dimethachlor sulfonic acid (surrogate) | Dimethachlor SA (surrogate) | 90625 | 100 | 34 | 103.7685 | 10.97 | 81.4757 | 142.5118 | 102.2171 | 7.5036 | 98.2029 | 108.3253
na | 9031 | Surrogate | acid | 2,4-D | 2,4-D-d3 (surrogate) | 91986 | 100 | 34 | 93.4483 | 9.5532 | 74.8764 | 121.1008 | 92.9823 | 8.4956 | 87.0145 | 98.4751
na | 9031 | Surrogate | pyrethroid & organochlorine& phenylpyrazine | Permethrin | 3-Phenoxybenzoic acid-13C6 (surrogate) | 90516 | 100 | 34 | 102.9185 | 6.768 | 94.5892 | 130.426 | 101.1172 | 4.7078 | 98.538 | 104.8888
na | 9031 | Surrogate | sulfonylurea & urea | Diuron | Diuron-d6 (surrogate) | L903M | 100 | 34 | 107.4387 | 17.5512 | 75.7292 | 149.5668 | 107.5062 | 19.3551 | 91.0517 | 117.1618
na | 9031 | Surrogate |  | Linuron | Linuron-d6 (surrogate) | L903K | 100 | 34 | 103.9017 | 9.1582 | 85.7652 | 127.8298 | 101.584 | 10.9464 | 96.8419 | 111.6085
E | 9030 | Analyte | acetanilide & amide | Acetochlor/Metolachlor | 2-(1-hydroxyethyl)-6-methylaniline (HEMA) | 68611 | 94 | 32 | 55.9994 | 38.5923 | 14.1535 | 110.6577 | 53.8428 | 14.2375 | 47.7663 | 66.9726
E | 9030 | Analyte | carbamate & thiocarbamate | Asulam | Asulam | 68536 | 28 | 32 | 98.684 | 22.2869 | 55.0767 | 163.5803 | 97.988 | 14.2311 | 88.2015 | 107.3994
E | 9030 | Analyte |  | Carbofuran | 3-Hydroxycarbofuran | 68508 | 16 | 32 | 87.8111 | 17.1017 | 62.7257 | 118.6194 | 88.1616 | 16.6344 | 78.2596 | 100.6994
E | 9030 | Analyte |  |  | 7-Hydroxycarbofuran | 68518 | 500 | 32 | 61.8076 | 91.5196 | 0 | 192.7484 | 69.3576 | 80.6618 | 0 | 108.8127
E | 9030 | Analyte | fungicide | Benomyl | Carbendazim | 68548 | 2 | 32 | 112.5492 | 33.668 | 73.6578 | 185.3328 | 93.001 | 51.7415 | 85.6692 | 155.4685
E | 9030 | Analyte |  | Famoxadone | Famoxadone | 67609 | 60 | 32 | 61.3945 | 55.2988 | 0 | 123.4896 | 67.8995 | 39.7337 | 33.4708 | 87.0716
E | 9030 | Analyte |  | Propiconazole | 1H-1,2,4-Triazole | 68498 | 22 | 32 | 79.993 | 25.5442 | 50.4602 | 138.2576 | 81.9009 | 22.2516 | 61.8021 | 91.8195
E | 9030 | Analyte | miscellaneous | Butralin | Butralin | 68545 | 4 | 31 | 89.9348 | 13.6471 | 68.4524 | 117.1412 | 87.6383 | 10.7081 | 83.5845 | 98.0297
E | 9030 | Analyte |  | Fenbutatin oxide | Fenbutatin oxide | 68602 | 20 | 31 | 73.5011 | 31.7303 | 33.9656 | 129.6815 | 69.2717 | 25.2434 | 56.9606 | 91.0139
E | 9030 | Analyte |  | Flumetsulam | Flumetsulam | 61679 | 16 | 32 | 98.9479 | 15.6687 | 69.0137 | 136.5476 | 95.0086 | 16.276 | 88.9185 | 110.8748
E | 9030 | Analyte |  | Oxyfluorfen | Oxyfluorfen | 65093 | 500 | 32 | 96.0019 | 53.0086 | 0 | 226.8669 | 86.3413 | 44.6121 | 63.9027 | 124.0845
E | 9030 | Analyte | organophosphate | Diazinon | Diazinon oxon | 68236 | 4 | 32 | 66.5596 | 17.698 | 45.7286 | 97.5384 | 66.0553 | 11.2615 | 56.7447 | 71.9365
E | 9030 | Analyte |  | Naled | Naled | 68654 | 30 | 32 | 67.5122 | 39.4363 | 8.5404 | 128.4594 | 68.9444 | 22.2321 | 50.0472 | 80.0384
E | 9030 | Analyte |  | Parathion-methyl | Parathion-methyl | 65089 | 500 | 32 | 99.4712 | 41.442 | 0 | 173.058 | 99.1832 | 51.9641 | 69.4516 | 139.5512
E | 9030 | Analyte |  | Phosmet | Phosmet | 65101 | 20 | 32 | 63.4397 | 55.7891 | 19.5263 | 149.7282 | 58.6129 | 31.6991 | 34.2048 | 76.9668
E | 9030 | Analyte | pyrethroid & organochlorine& phenylpyrazine | Bifenthrin | Bifenthrin | 65067 | 18 | 31 | 62.412 | 33.5921 | 20.0564 | 106.9634 | 64.2985 | 24.5503 | 44.3626 | 77.4809
E | 9030 | Analyte | sulfonylurea & urea | Novaluron | Novaluron | 68655 | 8 | 32 | 76.5626 | 42.9689 | 12.631 | 158.8726 | 77.5688 | 28.7091 | 54.0522 | 92.7807
E | 9031 | Analyte | acid | Dacthal | Chlorthal-monomethyl | 68560 | 500 | 34 | 107.8917 | 53.7392 | 17.9372 | 275.7378 | 99.0645 | 54.5715 | 65.652 | 139.269
E | 9031 | Analyte |  | Dicamba | Dicamba | 68571 | 500 | 34 | 88.7687 | 27.5925 | 47.824 | 139.3388 | 87.4745 | 23.7767 | 71.7193 | 103.794
E | 9031 | Analyte | sulfonylurea & urea | Halosulfuron-methyl | Chlorosulfonamide acid | 68551 | 74 | 34 | 92.8234 | 28.3469 | 46.03 | 176.8338 | 92.3962 | 18.6627 | 77.41 | 102.586
E | 9031 | Analyte | triazine | Hexazinone | Hexazinone TP D | 68613 | 200 | 34 | 95.3534 | 34.4779 | 0 | 176.4227 | 91.7753 | 33.0566 | 72.958 | 117.5513
9030 | 9031 | Analyte | acetanilide & amide | Propanil | Propanil | L903F | 58 | 34 | 98.6364 | 12.2157 | 74.4565 | 122.9824 | 99.0592 | 13.6129 | 90.4885 | 108.8523
9030 | 9031 | Analyte | sulfonylurea & urea | Chlorimuron-ethyl | Chlorimuron-ethyl | L903C | 12 | 34 | 93.0448 | 20.9474 | 61.3015 | 136.753 | 90.002 | 22.3197 | 76.0048 | 106.114
9030 | 9031 | Analyte |  | Chlorsulfuron | Chlorsulfuron | L903R | 96 | 24 | 102.7884 | 21.5996 | 67.4388 | 152.9127 | 100.0155 | 28.9438 | 83.7998 | 122.8449
9030 | 9031 | Analyte |  | Novaluron | Novaluron | L903J | 10 | 34 | 95.7684 | 58.3349 | 14.2411 | 244.7896 | 95.7284 | 59.8328 | 50.176 | 130.8905
9030 | 9031 | Analyte |  | Prosulfuron | Prosulfuron | L903N | 10 | 34 | 100.405 | 20.9947 | 63.2863 | 162.7479 | 98.7277 | 20.6018 | 83.8705 | 111.6624
9030 | 9031 | Surrogate | sulfonylurea & urea | Nicosulfuron | Nicosulfuron-d6 (surrogate) | 90554 | 100 | 34 | 100.1802 | 14.9646 | 71.2448 | 161.9451 | 98.754 | 11.5564 | 90.8301 | 106.4197
9031 | 9030 | Analyte | acetanilide & amide | Dimethenamid | Dimethenamid SAA | L903G | 500 | 32 | 93.0711 | 25.719 | 60.6195 | 151.8476 | 86.9198 | 22.7735 | 75.8966 | 106.618
D | 9030 | Analyte | carbamate & thiocarbamate | Carbofuran | 3-Ketocarbofuran | 68509 | 176 | 32 | 37.1072 | 111.6305 | 0 | 127.9832 | 22.9024 | 52.5573 | 0 | 70.8998
D | 9030 | Analyte |  | Formetanate | Formetanate | 68609 | 2 | 32 | 28.0806 | 60.5444 | 0 | 77.2128 | 27.9986 | 15.9099 | 16.6898 | 38.1523
D | 9030 | Analyte | miscellaneous | Aminopyralid | Aminopyralid | 68534 | 500 | 32 | 113.9144 | 55.2247 | 0 | 315.9641 | 99.5957 | 37.2066 | 83.3691 | 133.5609
D | 9030 | Analyte |  | Bifenazate | Bifenazate | 68540 | 10 | 32 | 162.4123 | 227.4982 | 0 | 1573.5697 | 75.2073 | 41.3479 | 41.1193 | 96.8976
D | 9030 | Analyte |  |  | Bifenazate diazene | 68541 | 4 | 32 | 46.6812 | 87.825 | 0 | 126.4244 | 25.3795 | 54.7282 | 16.4377 | 90.266
D | 9030 | Analyte |  | flumiclorac-pentyl | flumiclorac-pentyl | 68607 | 4 | 32 | 29.4843 | 123.4678 | 0 | 118.1798 | 10.7337 | 39.7063 | 1.5777 | 55.1415
D | 9030 | Analyte | organophosphate | Phosmet | Phosmet oxon | 68674 | 20 | 32 | 16.5797 | 278.0739 | 0 | 184.9273 | 0 | 0 | 0 | 0
D | 9030 | Analyte | fungicide | Iprodione | Deisopropyliprodione | 68565 | 500 | 34 | 280.7606 | 104.1335 | 19.7195 | 1401.9634 | 226.7328 | 159.7811 | 120.3649 | 335.9096
D | 9030 | Analyte |  |  | Iprodione | 66617 | 2000 | 14 | 438.8596 | 257.7787 | 0 | 3551.2395 | 0 | 0 | 0 | 0
D | 9030 | Analyte | triazine | Atrazine | Ammelide | 68535 | 74 | 34 | 81.9556 | 40.1394 | 34.2427 | 157.623 | 77.3223 | 34.0017 | 57.6933 | 103.5616
Note: results for bifenthrin, butralin, cis-permethrin, etoxazole, fenbutatin oxide, propargite, pyridaben, trans-permethrin, tribuphos in lab reagent spike 20123349001 were deleted because of interference