## Tbl 1–1. Acquisition Method
Table 1–1. Typical instrument operating conditions specific to the Agilent Technologies 6460 triple-quadrupole tandem mass spectrometer.
[°C, degrees Celsius; L/min, liters per minute; lb/in2, pounds per square inch; --, none; TIC, total ion current; N/A, not applicable; µL/min, microliters per minute; µL, microliters; mm, millimeters; mL/min, milliliters per minute; M, molar; mL, milliliters; %, percent; ±, plus or minus; #, number]
Electrospray ionization source parameters |  | Chromatogram plotted during analysis |  |  |  | Wellplate sampler conditions |  | Binary pump |  | Pump time table |  |  |  | Signals | Thermostated column compartment
Parameter | Value | Chromatogram type | Label | Offset | Y-Range | Condition | Type or value used | Condition | Type or value used | Time | Flow | Pressure | Solvent ratio B | Description | Condition | Type or value used
Gas tempature (°C) | 300 | TIC | TIC | 0 | 10000000 | Name | h-ALS-SL+ | Name | BinPump | 0 | 0.6 | 600 | 10 | Pressure | Name | Column-SL
Gas flow (L/min) | 10 | -- | -- | -- | -- | Ordinal # | 1 | Ordinal # | 1 | 1 | 0.6 | 600 | 10 | Flow | Ordinal # | 1
Nebulizer (lb/in2) | 45 | -- | -- | -- | -- | Stop time (minutes) | As pump | Stop time (minutes) | 33 | 5 | 0.6 | 600 | 40 | Solvent% B | Stop time (minutes) | As pump
Sheath gas heater (°C) | 325 | -- | -- | -- | -- | Injection type | Needle wash | Flow (mL/min) | 0.6 | 10 | 0.6 | 600 | 60 | -- | Left temperature (°C) | 60
Sheath gas flow (L/min) | 11 | -- | -- | -- | -- | Overlap time (minutes) | Disable overlapped injection | Pressure max (bar) | 600 | 15 | 0.6 | 600 | 100 | -- | Left ready (°C) | When temperature within set point ±0.8
Capillary (volts) | 3500 | -- | -- | -- | -- | Draw position detection | 0 | Solvent A | Aqueous mobile phase solution: 10 mL 1M ammonium formate and 12 mL 1M formic acid (98 percent) | 25 | 0.6 | 600 | 100 | -- | Valve position | 1
Charging voltage (volts) | 500 | -- | -- | -- | -- | Eject speed (µL/min) | 500 | Solvent ratio A | 90 | 25.01 | 0.6 | 600 | 10 | -- | Contact 1 | Off
-- | -- | -- | -- | -- | -- | Automatic delay volume reduction | No | Solvent type A1 | N/A | 33 | 0.6 | 600 | 10 | -- | Contact 2 | Off
-- | -- | -- | -- | -- | -- | Wash vessel | N/A | Solvent type A2 | N/A | -- | -- | -- | -- | -- | Contact 3 | Off
-- | -- | -- | -- | -- | -- | Wash time (seconds) | 20 | Compression A (*10-6/bar) | 100 | -- | -- | -- | -- | -- | Contact 4 | Off
-- | -- | -- | -- | -- | -- | Contact 1 | Off | Stroke A (µL) | Auto | -- | -- | -- | -- | -- | Model | G1316B
-- | -- | -- | -- | -- | -- | Contact 2 | Off | Stroke syncronization | N/A | -- | -- | -- | -- | -- | Options | Column switching valve
-- | -- | -- | -- | -- | -- | Contact 3 | Off | Contact 1 | Off | -- | -- | -- | -- | -- | Post time (minutes) | Off
-- | -- | -- | -- | -- | -- | Contact 4 | Off | Contact 2 | Off | -- | -- | -- | -- | -- | Right temperature (°C) | 60
-- | -- | -- | -- | -- | -- | Model | G1367D | Contact 3 | Off | -- | -- | -- | -- | -- | Right ready (°C) | When temperature within set point ±0.8
-- | -- | -- | -- | -- | -- | Options | Thermostatted temperature control module | Contact 4 | Off | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Post time (minutes) | Off | Model | G1312B | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Injection volume (µL) | 100 | Options | Solvent selection valve | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Draw position (mm) | 0 | Post time (minutes) | Off | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Draw speed (µL/min) | 140 | Pressure minimum (bar) | 0 | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Flush out factor | 5 | Maximum flow gradient (mL/min) | 100 | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Equilibration time (seconds) | 0 | Solvent B | N/A | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Wash location | Flush port | Solvent ratio B | 10 | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | Wash cycles | N/A | Solvent type B1 | N/A | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | -- | -- | Solvent type B2 | N/A | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | -- | -- | Compression B (*10-6/bar) | 115 | -- | -- | -- | -- | -- | -- | --
-- | -- | -- | -- | -- | -- | -- | -- | Stroke B (µL) | Auto | -- | -- | -- | -- | -- | -- | --

## Tbl 1–2. Performance Criteria
Table 1–2. General quality-control guidelines for performance criteria and corrective actions applied to quality-control samples.
[QC, quality control; ±, plus or minus; NWQL, National Water Quality Laboratory; IDS, isotope-dilution standard]
QC sample or compound | Performance criteria | Corrective action
Continuing calibration verification (CCV) samples | Recoveries of individual pharmaceutials in CCV samples bracketing environmental samples fall within ±3 standard deviations of the long-term mean CCV recovery, which is calculated from the recoveries determined for at least 30 previously analyzed CCV samples. | If recovery of a pharmaceutical in a particular CCV is outside the performance criteria, any concentrations of that pharmaceutical in an environmental sample are qualified as an estimate. If more than five pharmaceuticals fail to meet criteria, samples bracketed by the failed CCV samples are reanalyzed.
Calibration curve quality calculated from analysis of calibration samples | The correlation coefficient, R2, should be greater than or equal to 0.98, for each pharmaceutical calibration curve, defined as a quadratic fit of the plot of the response of the calibrant sample concentration against the isotope-dilution standard normalized pharmaceutical response, using a 1/x2 weighting, and not forcing the calibration curve through the plot origin. | When the calibration curve for a pharmaceutical fails to meet the performance criteria, any detected concentrations of that pharmaceutical are reported as an estimate. If more than five pharmaceuticals fail to meet criteria, standards are remade, and the instrument is serviced and recalibrated prior to reanalysis.
Laboratory reagent spike (LRS) samples | Recoveries of individual pharmaceutials in LRS samples fall within ±3 standard deviations of the long-term mean LRS recovery, which is calculated from the recoveries determined for at least 30 previously analyzed LRS samples. | Concentrations of any pharamceuticals detected in environmental samples are qualified as estimates if that pharmaceutical in the LRS fails to meet performance criteria. If more than five pharmaceuticals fail to meet criteria, or other data indicate an LRS failure, reanalysis may be required.
Laboratory reagent blank (LRB) and continuing calibration blank (CCB) samples | Any pharmaceuticals detected in CCB or LRB samples are less than the laboratory reporting level (LRL). Note that the CCB samples, which bracket groups of environmental and QC samples, are used to determine if blank contamination during analysis is present and requires qualifying environmental or QC sample results. | The NWQL's quality management manual (Stevenson and Barnard, 2013) provides general guidance on evaluating blank contamination and applying corrective action. For this method, corrective action is cleaning of pertinent HPLC, autosampler, and mass spectrometer components. The Office of Water Quality Technical Memorandum 2012.01 (http://water.usgs.gov/admin/memo/QW/qw12.01.pdf; last accessed April 2012) provides specific guidance for qualifying environmental sample results when detections occur in blanks, specifically for CCB samples for this method.
Limit of quantitation (LOQ) samples | All pharmaceuticals in the LOQ sample are detected at the lower bound of the method operational range. | For analytes not detected, the reporting level is raised to value of lowest detected standard.
Surrogates | Sample surrogate recoveries fall within long-term control limits (±3 standard deviations from the mean of surrogate recoveries, calculated from reagent spike samples). | If the surrogate recovery in the set spike or bracketing CCV is outside of control limits, the surrogate recovery in samples is reported as an estimate.
Isotope Dilution Standard (IDS) | Response of sample IDS compounds must fall within ±50 percent deviation from the mean response of the calibration standards. | Possible causes are evaluated. Affected samples are re-prepared and re-analyzed for detected analytes compromised by the failed IDS or the concentrations for the affected samples and analytes are reported as estimates.
Third-party check (TPC) | Calculated TPC concentrations must fall within ±30 percent of the expected concentration. | Values are qualified as estimated in all data files for detected analytes that failed criteria.

## Tbl 1–3. User Annotation
Table 1–3. Examples of Laboratory Information System Management data actions applied by analysts using entries to MassHunterTM software user annotation field.
[NWIS, National Water Information System; QC, quality control; <, less than; IRL, interim reporting level; LRL, laboratory reporting level; LIMS, Laboratory Information Management System]
MassHunterTM user annotation entry | Laboratory Information Management System (LIMS) data action | Corresponding NWIS remark code, if applicable | Corresponding NWIS data quality indicator code, if applicable | Explanation
q,Q | Q-deleted |  |  | Failure of QC sample, only applicable to QC samples, data are used for statistics, unknown cause.
r,R | R-deleted |  |  | Sample ruined in preparation or analysis.
u,U | U-deleted |  | i | Unable to determine; matrix interference.
x,X | X-deleted |  |  | Failure of QC sample, only applicable to QC samples, data excluded from statistics, known cause.
o,O | O-deleted |  | q | Insufficient sample received for analysis.
e,E | E | E |  | Estimated concentration. Several potential reasons for applying the E remark code can occur, including contributions from an interference, internal standards not meeting acceptance criteria, and set QC sample results out of range. When the "E" remark is applied by analyst, an explanatory comment is provided.
rrl, RRL, rrL,rRL,Rrl, RRl | < | < |  | < Final concentration, that is, a raised reporting level. An explanatory comment is provided.
z,Z | 0 |  |  | < IRL or < LRL.
“v” | LIMS or analyst applies data qualifier code “v” (or other appropriate qualifier code corresponding to specific situations described in LIMS) |  | v | The "v" data quality indicator code is applied to a result if carryover or other sources of blank contamination are detected and the calculated blank concentration is less than one third of the environmental concentration, and the environmental concentration is above the method reporting limit.

## Tbl 1–4. IDS Perf. Criteria
Table 1–4. General quality-control guidelines for performance criteria and corrective actions applied to potentially unacceptable isotope-dilution standard response.
[IDS, isotope dilution standard, °C, degrees Celsius; MS, mass spectrometer]
Symptoms of potentially unaceptable IDS response | Probable cause | Corrective action
Low response of an IDS compound relative to the mean response | Matrix related or suppression effect in analysis | Re-prepare sample and re-analyze. If symptom duplicates, consider referencing to another IDS and re-calibrating, consider external standard quantitation, or qualify the compromised values as estimated.
Variability of IDS areas outside performance criteria | Characteristic of some labeled compounds, especially fluoxetine-d6, long duration in water matrix at 4 °C, and other circumstances | Consider external standard quantitation or qualify the compromised values as estimated, or consider freezing then allow to thaw just before injection.
Low response of all IDS compounds | Not enough IDS added in sample preparation, mixture of solution components not uniform throughout or solution got warm, poor injection, and other circumstances | Re-prepare sample and re-analyze. Check surrogate recoveries; if unacceptable, report results for sample as "deleted, ruined."
Loss of IDS | Retention time shift in column or MS mass axis adjustment needed. Possibly matrix related | Enter correct retention time into the acquisition method and perform autotune. Re-prepare sample and re-analyze. If symptom duplicates, consider referencing to another IDS and re-calibrating, or consider external standard quantitation.
Unsymmetrical IDS quantitation ion peak and other compound peak shapes OK | Matrix interference | Integrate appropriately; consider reporting compounds quantified with the affected IDS as "deleted unrecoverable."